## Final Report
Final Report
Job 2014-0009
|  | Isotope(s) requested: | δ13C, δ15N
|  | Instrument Used: | Costech 4010 Elemental Analyzer coupled to a Thermo Delta Plus XP IRMS
|  | Analytical Code: | 010 (combustion/reduction organics)
|  | Units: | δ13C and δ15N values are reported w.r.t. VPDB and AIR respectively in parts per thousand (per mil)
|  | Principal Investigator: | Jason Stockwell
|  | Job submission contact: | Peter Euclide
|  | Sample Material(s): | Mysis Tissue
|  | Date Submited: | 2013-01-14
|  | Number of unknown samples analyzed: | 72
|  | Number of reference samples analyzed: | 24
|  | Quality Control Reference Material 1: | 36-UWSIF-Glutamic 1
|  | Quality Control Reference Material 2: | 39-UWSIF-Glutamic 2
|  | Quality Assurance Reference Material 3: | UWSIF01 (Liver)
Quality Control Data |  |  |  |  | Quality Assurance Data
Reference Material 1 | Known |  | Known |  | Reference Material 3 | Known | Normalized | Known | Normalized
36-UWSIF-Glutamic 1 | δ13C PDB |  | δ15N AIR |  | UWSIF01 (Liver) | δ13C PDB | δ13C PDB | δ15N AIR | δ15N AIR
average (N=8) | -28.28 |  | -4.62 |  | average (N=8) | -17.82 | -17.7177 | 6.78 | 6.8116
standard uncertainty | 0.15287 |  | 0.047355 |  | standard uncertainty |  | 0.058844 |  | 0.056114
Reference Material 2 | Known |  | Known |  | Long-Term | δ13C PDB | -18.12 | δ15N AIR | 6.38
39-UWSIF-Glutamic 2 | δ13C PDB |  | δ15N AIR |  | Acceptable Range | 2 σ = 0.3 | -17.52 | 2 σ = 0.4 | 7.18
average (N=8) | 24.36 |  | 27.89 |  | Weight Percent | Known | Measured | Known | Measured
standard uncertainty | 0.14264 |  | 0.037814 |  |  | Wt% C | Wt% C | Wt% N | Wt% N
[5 empty] | average (N=8) | 49.54 | 49.879 | 12.9 | 13.0131
[5 empty] | standard uncertainty |  | 0.18575 |  | 0.039135
[5 empty] | relative error (%) |  | 0.68434 |  | 0.87679
Record Keeping |  |  |  |  | Quality Assurance Approval
Date Reported: | 2014-03-13 |  |  |  | Reviewer: | Craig Cook
Date Invoiced: | 2014-03-13 |  |  |  | Title: | Facility Director
Initial: | cjm |  |  |  | Date Reviewed: | 2014-03-12
Comments: |  |  |  |  | Comments: | Data meet all QAQC criteria in the SOP
Analytical Comments:
*Sample weight percents are calculated using the sample weights reported by the user and are dependent upon the accuracy of these values.
Quality Control Color Legend
Yellow | Peak amplitudes too low for reliable results. Use with extreme caution or rerun the sample.
Green | Weight percent values wrong. Use with caution.
Pink | Potential outlier. Use with caution.
Blue | Sample lost during the analysis. Reload and rerun the sample.
Olive | Possible sample ID problem. Check loading documents.
SIF ID | Sample ID | δ15N | δ13C | Wt% N* | Wt% C* | C:N ratio | Comments
20140009.031 | 111 | 12.7414 | -28.5557 | 9.5511 | 39.1998 | 4.0693
20140009.002 | 139 | 14.1791 | -32.0225 | 9.7715 | 50.4546 | 5.1627
20140009.061 | 141 | 12.7504 | -29.5915 | 10.9164 | 42.062 | 3.7875
20140009.062 | 160 | 12.5967 | -28.3462 | 11.523 | 44.7748 | 3.8244
20140009.001 | 164 | 14.268 | -32.7719 | 9.381 | 53.8562 | 5.7448
20140009.032 | 171 | 13.3195 | -29.8933 | 10.3908 | 53.4252 | 5.1383
20140009.003 | 225 | 13.628 | -29.7126 | 10.3487 | 50.4243 | 4.8693
20140009.064 | 237 | 13.3934 | -28.7384 | 12.0004 | 47.268 | 3.8811
20140009.063 | 238 | 12.999 | -28.6645 | 11.8868 | 46.9399 | 3.8916
20140009.033 | 248 | 12.6945 | -29.3164 | 10.8462 | 48.8916 | 4.487
20140009.004 | 273 | 13.3505 | -29.2578 | 10.9302 | 49.3625 | 4.5099
20140009.034 | 300 | 12.9601 | -29.7681 | 10.294 | 51.4617 | 4.9915
20140009.005 | 310 | 13.3005 | -29.9077 | 10.7923 | 52.2853 | 4.8409
20140009.036 | 324 | 13.8656 | -29.1613 | 11.4125 | 49.104 | 4.276
20140009.035 | 341 | 12.7784 | -28.723 | 9.8682 | 42.1304 | 4.238
20140009.065 | 361 | 12.1564 | -27.72 | 10.0114 | 35.9338 | 3.4998
20140009.066 | 376 | 12.99 | -28.2436 | 11.4333 | 45.1653 | 3.8886
20140009.006 | 398 | 14.3888 | -28.7938 | 11.0263 | 46.6316 | 4.2194
20140009.007 | 410 | 14.9938 | -31.0441 | 9.6724 | 50.8289 | 5.255
20140009.038 | 429 | 13.04 | -28.9283 | 10.4256 | 47.4484 | 4.5287
20140009.037 | 446 | 13.4703 | -30.2044 | 10.0995 | 50.623 | 5.0045
20140009.008 | 479 | 13.8227 | -29.457 | 11.2397 | 50.1741 | 4.4558
20140009.068 | 483 | 14.3718 | -30.5719 | 10.6766 | 51.159 | 4.7675
20140009.067 | 499 | 12.6716 | -29.6664 | 10.2849 | 49.4576 | 4.7848
20140009.070 | 507 | 14.3798 | -29.1737 | 12.9789 | 50.4119 | 3.8247
20140009.069 | 539 | 13.8946 | -28.5341 | 12.6952 | 46.5932 | 3.6006
20140009.010 | 541 | 13.2027 | -28.0526 | 12.1176 | 45.077 | 3.7038
20140009.039 | 542 | 14.9269 | -30.4662 | 12.2445 | 49.732 | 4.0288
20140009.009 | 544 | 14.6714 | -30.3194 | 11.9254 | 46.669 | 3.8997
20140009.040 | 576 | 14.8381 | -30.2547 | 13.4845 | 49.2936 | 3.6151
20140009.017 | 603 | 13.9166 | -29.3574 | 11.0082 | 46.7912 | 4.2351
20140009.047 | 606 | 12.1494 | -28.4571 | 10.5823 | 43.8279 | 4.1009
20140009.048 | 613 | 13.9625 | -29.3225 | 11.2857 | 47.272 | 4.1516
20140009.018 | 638 | 14.3509 | -28.6296 | 13.5308 | 49.1373 | 3.6106
20140009.045 | 705 | 13.7688 | -29.3851 | 11.0269 | 48.1699 | 4.3383
20140009.046 | 711 | 14.274 | -28.9283 | 11.7552 | 47.1416 | 3.9701
20140009.015 | 714 | 14.0663 | -29.0546 | 11.6112 | 48.4991 | 4.1623
20140009.016 | 718 | 13.9725 | -29.3359 | 11.0121 | 47.6357 | 4.312
20140009.013 | 834 | 13.0949 | -29.032 | 12.0237 | 49.3525 | 4.0907
20140009.014 | 846 | 13.3964 | -28.7558 | 10.8012 | 48.0051 | 4.4331
20140009.044 | 898 | 14.6953 | -29.302 | 11.0947 | 50.8978 | 4.5654
20140009.043 | 899 | 13.4004 | -28.4047 | 10.5173 | 40.1987 | 3.7728
20140009.012 | 938 | [5 empty] | Sample lost during the analysis.
20140009.011 | 941 | [5 empty] | Sample lost during the analysis.
20140009.042 | 980 | 12.8333 | -29.2671 | 11.5651 | 49.7339 | 4.271
20140009.041 | 989 | 13.045 | -27.5763 | 12.2425 | 43.1998 | 3.4796
20140009.020 | 1003 | 13.2756 | -28.8267 | 10.0346 | 40.5065 | 4.014
20140009.019 | 1031 | 14.3748 | -29.3328 | 10.7222 | 45.836 | 4.2583
20140009.050 | 1035 | 12.6716 | -29.4211 | 11.2297 | 49.2825 | 4.3545
20140009.049 | 1039 | 14.3119 | -28.2538 | 11.4145 | 47.2623 | 4.0981
20140009.021 | 1115 | 13.9245 | -28.0095 | 11.5053 | 46.0268 | 3.9795
20140009.022 | 1128 | 13.3784 | -29.3348 | 12.0732 | 46.5392 | 3.8317
20140009.052 | 1132 | 14.5276 | -28.1912 | 10.233 | 37.474 | 3.5912
20140009.051 | 1146 | 15.0467 | -28.1933 | 12.2149 | 46.1067 | 3.7206
20140009.023 | 1224 | 14.2171 | -27.9838 | 12.2332 | 44.6371 | 3.622
20140009.053 | 1242 | 14.3918 | -29.2024 | 11.5851 | 47.3962 | 4.045
20140009.024 | 1291 | 13.9575 | -28.9386 | 12.1923 | 45.9058 | 3.7395
20140009.054 | 1299 | 14.8371 | -28.4376 | 12.2642 | 44.7246 | 3.5859
20140009.025 | 1331 | 14.1642 | -28.6994 | 12.2399 | 45.4656 | 3.6828
20140009.056 | 1358 | 13.8616 | -28.8903 | 12.7598 | 45.2931 | 3.486
20140009.055 | 1367 | 14.1612 | -29.1911 | 12.8117 | 45.7774 | 3.5114
20140009.026 | 1369 | 14.5366 | -28.8852 | 12.507 | 45.3327 | 3.591
20140009.030 | 1481 | 13.4613 | -28.4047 | 12.1428 | 43.8654 | 3.5744
20140009.059 | 1485 | 12.2982 | -27.7344 | 11.6219 | 41.9578 | 3.54
20140009.060 | 1494 | 14.4447 | -29.304 | 11.661 | 48.1302 | 4.0788
20140009.029 | 1497 | 12.4938 | -28.1296 | 11.5433 | 43.1459 | 3.7014
20140009.028 | 1515 | 14.5326 | -29.687 | 10.5566 | 49.4485 | 4.6694
20140009.027 | 1522 | 14.6174 | -29.42 | 10.9199 | 49.117 | 4.4796
20140009.057 | 1543 | 14.3469 | -31.6591 | 10.2116 | 49.4522 | 4.823
20140009.058 | 1550 | 12.4479 | -29.6305 | 10.0216 | 47.3394 | 4.6974
20140009.071 | 1601 | 13.3934 | -29.4909 | 10.79 | 50.402 | 4.64
20140009.072 | 1602 | 14.6354 | -29.6028 | 10.7792 | 49.5291 | 4.5592
Stable Isotope Facility

## QAQC, calculations
Line | Identifier 1 | δ15N | δ13C | δ15N | δ13C | Line | Identifier 1 | δ15N | δ13C | δ15N | δ13C | Line | Identifier 1 | Absolute Difference |  | Relative Error |  | δ15N |  | δ13C
|  | Normalized | Normalized | known | known |  |  | Normalized | Normalized | known | known |  |  | % Nitrogen | % Carbon | % Nitrogen | % Carbon | normalized |  | normalized
QC Reference Materials | [5 empty] | QC Reference Materials | [5 empty] | QA Reference Material
7 | 36-UWSIF-UT Glut 120140009.14 | -4.6467 | -27.9808 |  |  | 9 | 39-UWSIF-UW Glut 220140009.11 | 27.8582 | 24.5565 |  |  | 11 | 01-UWSIF-Liver 20140009.11 | 0.16586 | 0.55914 | 13.0659 | 50.0991 | 3.8275 | 6.8149 | -17.6347
8 | 36-UWSIF-UT Glut 120140009.15 | -4.6117 | -28.1666 |  |  | 10 | 39-UWSIF-UW Glut 220140009.12 | 27.8882 | 24.5791 |  |  | 12 | 01-UWSIF-Liver 20140009.12 | 0.12086 | 0.41371 | 13.0209 | 49.9537 | 3.8289 | 6.771 | -17.6942
37 | 36-UWSIF-UT Glut 120140009.16 | -4.6467 | -28.2703 |  |  | 39 | 39-UWSIF-UW Glut 220140009.13 | 27.9111 | 24.3399 |  |  | 41 | 01-UWSIF-Liver 20140009.13 | 0.084143 | 0.25218 | 12.9841 | 49.7922 | 3.8086 | 6.8559 | -17.687
38 | 36-UWSIF-UT Glut 120140009.17 | -4.5309 | -28.3082 |  |  | 40 | 39-UWSIF-UW Glut 220140009.14 | 27.9221 | 24.2722 |  |  | 42 | 01-UWSIF-Liver 20140009.14 | 0.091157 | 0.19122 | 12.9912 | 49.7312 | 3.8011 | 6.796 | -17.7651
67 | 36-UWSIF-UT Glut 120140009.18 | -4.6297 | -28.3709 |  |  | 69 | 39-UWSIF-UW Glut 220140009.15 | 27.9391 | 24.2013 |  |  | 71 | 01-UWSIF-Liver 20140009.15 | 0.072806 | 0.085599 | 12.9728 | 49.6256 | 3.779 | 6.779 | -17.6686
68 | 36-UWSIF-UT Glut 120140009.19 | -4.6327 | -28.3298 |  |  | 70 | 39-UWSIF-UW Glut 220140009.16 | 27.7883 | 24.3584 |  |  | 72 | 01-UWSIF-Liver 20140009.16 | 0.073822 | 0.17697 | 12.9738 | 49.717 | 3.7853 | 6.9218 | -17.763
97 | 36-UWSIF-UT Glut 120140009.2 | -4.6008 | -28.3945 |  |  | 99 | 39-UWSIF-UW Glut 220140009.17 | 27.9111 | 24.1592 |  |  | 101 | 01-UWSIF-Liver 20140009.17 | 0.12642 | 0.46928 | 13.0264 | 50.0093 | 3.7731 | 6.7411 | -17.7158
98 | 36-UWSIF-UT Glut 120140009.21 | -4.6287 | -28.4037 |  |  | 100 | 39-UWSIF-UW Glut 220140009.18 | 27.8822 | 24.4179 |  |  | 102 | 01-UWSIF-Liver 20140009.18 | 0.16978 | 0.56407 | 13.0698 | 50.1041 | 3.7669 | 6.8129 | -17.8133
8 | average | -4.616 | -28.2781 | -4.62 | -28.28 | 8 | average | 27.8875 | 24.3606 | 27.89 | 24.36 | 8 |  |  | Known % | 12.9 | 49.54 | δ15N known | 6.78 | δ13C known | -17.82
| std uncertainty | 0.037814 | 0.14264 |  |  |  | std uncertainty | 0.047355 | 0.15287 | [5 empty] | Average | 13.0131 | 49.879 | average δ15N | 6.8116 | average δ13C | -17.7177
[15 empty] | Stdev | 0.039135 | 0.18575 | stdev δ15N | 0.056114 | stdev δ13C | 0.058844
[15 empty] | Absolute Error | 0.11311 | 0.33902
[15 empty] | Relative Error | 0.87679 | 0.68434 | Long-Term | 7.18 | Long-Term | -17.52
[18 empty] | Acceptable Range | 6.38 | Acceptable Range | -18.12

## Run 1
Date | Line | Identifier 1 | Identifier 2 | Amount | Ampl 28 | d 15N/14N | Ampl 44 | d 13C/12C |  | Nitrogen % | Drift Corr. %N | Carbon % | Drift Corr for C% | C:N | Corr. d15N | Corr. d13C
2014/03/04 17:36:07 | 13 | 20140009.001 | 164 | 0.858 | 3174 | 14.689 | 3425 | -18.53 |  | 9.3066 | 9.381 | 53.4648 | 53.8562 | 5.7448 | 14.268 | -32.7719
2014/03/04 17:44:35 | 14 | 20140009.002 | 139 | 0.879 | 3376 | 14.6 | 3281 | -17.8 |  | 9.6913 | 9.7715 | 50.0331 | 50.4546 | 5.1627 | 14.1791 | -32.0225
2014/03/04 17:53:03 | 15 | 20140009.003 | 225 | 0.877 | 3572 | 14.048 | 3272 | -15.55 |  | 10.2628 | 10.3487 | 49.9727 | 50.4243 | 4.8693 | 13.628 | -29.7126
2014/03/04 18:01:31 | 16 | 20140009.004 | 273 | 0.784 | 3356 | 13.77 | 2876 | -15.107 |  | 10.8386 | 10.9302 | 48.8808 | 49.3625 | 4.5099 | 13.3505 | -29.2578
2014/03/04 18:09:59 | 17 | 20140009.005 | 310 | 0.825 | 3513 | 13.72 | 3203 | -15.74 |  | 10.695 | 10.7923 | 51.7735 | 52.2853 | 4.8409 | 13.3005 | -29.9077
2014/03/04 18:18:28 | 18 | 20140009.006 | 398 | 0.785 | 3381 | 14.81 | 2718 | -14.655 |  | 10.9233 | 11.0263 | 46.0897 | 46.6316 | 4.2194 | 14.3888 | -28.7938
2014/03/04 18:26:55 | 19 | 20140009.007 | 410 | 0.879 | 3359 | 15.416 | 3319 | -16.847 |  | 9.5636 | 9.6724 | 50.2569 | 50.8289 | 5.255 | 14.9938 | -31.0441
2014/03/04 18:35:23 | 20 | 20140009.008 | 479 | 0.86 | 3820 | 14.243 | 3203 | -15.301 |  | 11.1252 | 11.2397 | 49.572 | 50.1741 | 4.4558 | 13.8227 | -29.457
2014/03/04 18:43:51 | 21 | 20140009.009 | 544 | 0.801 | 3752 | 15.093 | 2781 | -16.141 |  | 11.8052 | 11.9254 | 46.0368 | 46.669 | 3.8997 | 14.6714 | -30.3194
2014/03/04 18:52:19 | 22 | 20140009.010 | 541 | 0.82 | 3924 | 13.622 | 2760 | -13.933 |  | 11.9916 | 12.1176 | 44.4147 | 45.077 | 3.7038 | 13.2027 | -28.0526
2014/03/04 19:00:48 | 23 | 20140009.011 | 941 | 0.813
2014/03/04 19:09:16 | 24 | 20140009.012 | 938 | 0.813 | [12 empty] | Samples Dropped together
2014/03/04 19:17:45 | 25 | 20140009.013 | 834 | 0.831 | 3950 | 13.514 | 3048 | -14.887 |  | 11.8806 | 12.0237 | 48.5998 | 49.3525 | 4.0907 | 13.0949 | -29.032
2014/03/04 19:26:13 | 26 | 20140009.014 | 846 | 0.827 | 3530 | 13.816 | 2965 | -14.618 |  | 10.6523 | 10.8012 | 47.2223 | 48.0051 | 4.4331 | 13.3964 | -28.7558
2014/03/04 19:34:41 | 27 | 20140009.015 | 714 | 0.797 | 3665 | 14.487 | 2887 | -14.909 |  | 11.4566 | 11.6112 | 47.6863 | 48.4991 | 4.1623 | 14.0663 | -29.0546
2014/03/04 19:43:10 | 28 | 20140009.016 | 718 | 0.786 | 3424 | 14.393 | 2798 | -15.183 |  | 10.8518 | 11.0121 | 46.7928 | 47.6357 | 4.312 | 13.9725 | -29.3359
2014/03/04 19:51:38 | 29 | 20140009.017 | 603 | 0.82 | 3572 | 14.337 | 2869 | -15.204 |  | 10.8422 | 11.0082 | 45.9182 | 46.7912 | 4.2351 | 13.9166 | -29.3574
2014/03/04 20:00:07 | 30 | 20140009.018 | 638 | 0.847 | 4532 | 14.772 | 3082 | -14.495 |  | 13.359 | 13.5308 | 48.2342 | 49.1373 | 3.6106 | 14.3509 | -28.6296
2014/03/04 20:08:35 | 31 | 20140009.019 | 1031 | 0.872 | 3689 | 14.796 | 2966 | -15.18 |  | 10.5447 | 10.7222 | 44.9028 | 45.836 | 4.2583 | 14.3748 | -29.3328
2014/03/04 20:17:04 | 32 | 20140009.020 | 1003 | 0.82 | 3229 | 13.695 | 2481 | -14.687 |  | 9.8514 | 10.0346 | 39.5431 | 40.5065 | 4.014 | 13.2756 | -28.8267
2014/03/04 20:25:33 | 33 | 20140009.021 | 1115 | 0.852 | 3881 | 14.345 | 2920 | -13.891 |  | 11.3163 | 11.5053 | 45.0333 | 46.0268 | 3.9795 | 13.9245 | -28.0095
2014/03/04 20:34:02 | 34 | 20140009.022 | 1128 | 0.795 | 3809 | 13.798 | 2765 | -15.182 |  | 11.8786 | 12.0732 | 45.5156 | 46.5392 | 3.8317 | 13.3784 | -29.3348
2014/03/04 20:42:30 | 35 | 20140009.023 | 1224 | 0.801 | 3865 | 14.638 | 2657 | -13.866 |  | 12.0329 | 12.2332 | 43.5834 | 44.6371 | 3.622 | 14.2171 | -27.9838
2014/03/04 20:50:59 | 36 | 20140009.024 | 1291 | 0.778 | 3749 | 14.378 | 2660 | -14.796 |  | 11.9862 | 12.1923 | 44.822 | 45.9058 | 3.7395 | 13.9575 | -28.9386
2014/03/04 21:50:21 | 43 | 20140009.025 | 1331 | 0.829 | 3993 | 14.585 | 2785 | -14.563 |  | 11.9937 | 12.2399 | 44.171 | 45.4656 | 3.6828 | 14.1642 | -28.6994
2014/03/04 21:58:50 | 44 | 20140009.026 | 1369 | 0.847 | 4206 | 14.958 | 2855 | -14.744 |  | 12.2551 | 12.507 | 44.0081 | 45.3327 | 3.591 | 14.5366 | -28.8852
2014/03/04 22:07:19 | 45 | 20140009.027 | 1522 | 0.831 | 3588 | 15.039 | 3026 | -15.265 |  | 10.6623 | 10.9199 | 47.7622 | 49.117 | 4.4796 | 14.6174 | -29.42
2014/03/04 22:15:48 | 46 | 20140009.028 | 1515 | 0.83 | 3446 | 14.954 | 3038 | -15.525 |  | 10.2933 | 10.5566 | 48.0637 | 49.4485 | 4.6694 | 14.5326 | -29.687
2014/03/04 22:24:17 | 47 | 20140009.029 | 1497 | 0.777 | 3529 | 12.912 | 2484 | -14.008 |  | 11.2742 | 11.5433 | 41.7309 | 43.1459 | 3.7014 | 12.4938 | -28.1296
2014/03/04 22:32:46 | 48 | 20140009.030 | 1481 | 0.809 | 3865 | 13.881 | 2621 | -14.276 |  | 11.8679 | 12.1428 | 42.4204 | 43.8654 | 3.5744 | 13.4613 | -28.4047
2014/03/04 22:41:15 | 49 | 20140009.031 | 111 | 0.83 | 3111 | 13.16 | 2420 | -14.423 |  | 9.2706 | 9.5511 | 37.7246 | 39.1998 | 4.0693 | 12.7414 | -28.5557
2014/03/04 22:49:44 | 50 | 20140009.032 | 171 | 0.751 | 3065 | 13.739 | 2988 | -15.726 |  | 10.1045 | 10.3908 | 51.9199 | 53.4252 | 5.1383 | 13.3195 | -29.8933
2014/03/04 22:58:13 | 51 | 20140009.033 | 248 | 0.774 | 3293 | 13.113 | 2799 | -15.164 |  | 10.5542 | 10.8462 | 47.3562 | 48.8916 | 4.487 | 12.6945 | -29.3164
2014/03/04 23:06:43 | 52 | 20140009.034 | 300 | 0.782 | 3176 | 13.379 | 2994 | -15.604 |  | 9.9963 | 10.294 | 49.8962 | 51.4617 | 4.9915 | 12.9601 | -29.7681
2014/03/04 23:15:12 | 53 | 20140009.035 | 341 | 0.844 | 3244 | 13.197 | 2617 | -14.586 |  | 9.5647 | 9.8682 | 40.5348 | 42.1304 | 4.238 | 12.7784 | -28.723
2014/03/04 23:23:41 | 54 | 20140009.036 | 324 | 0.839 | 3769 | 14.286 | 3042 | -15.013 |  | 11.1033 | 11.4125 | 47.4783 | 49.104 | 4.276 | 13.8656 | -29.1613
2014/03/04 23:32:11 | 55 | 20140009.037 | 446 | 0.772 | 3055 | 13.89 | 2893 | -16.029 |  | 9.7846 | 10.0995 | 48.9672 | 50.623 | 5.0045 | 13.4703 | -30.2044
2014/03/04 23:40:40 | 56 | 20140009.038 | 429 | 0.751 | 3065 | 13.459 | 2639 | -14.786 |  | 10.105 | 10.4256 | 45.7624 | 47.4484 | 4.5287 | 13.04 | -28.9283
2014/03/04 23:49:10 | 57 | 20140009.039 | 542 | 0.806 | 3857 | 15.349 | 2938 | -16.284 |  | 11.9181 | 12.2445 | 48.016 | 49.732 | 4.0288 | 14.9269 | -30.4662
2014/03/04 23:57:39 | 58 | 20140009.040 | 576 | 0.808 | 4293 | 15.26 | 2936 | -16.078 |  | 13.1525 | 13.4845 | 47.5474 | 49.2936 | 3.6151 | 14.8381 | -30.2547
2014/03/05 00:06:08 | 59 | 20140009.041 | 989 | 0.816 | 3929 | 13.464 | 2597 | -13.469 |  | 11.9047 | 12.2425 | 41.4235 | 43.1998 | 3.4796 | 13.045 | -27.5763
2014/03/05 00:14:38 | 60 | 20140009.042 | 980 | 0.798 | 3626 | 13.252 | 2928 | -15.116 |  | 11.2216 | 11.5651 | 47.9276 | 49.7339 | 4.271 | 12.8333 | -29.2671
2014/03/05 00:23:07 | 61 | 20140009.043 | 899 | 0.849 | 3485 | 13.82 | 2502 | -14.276 |  | 10.1681 | 10.5173 | 38.3622 | 40.1987 | 3.7728 | 13.4004 | -28.4047
2014/03/05 00:31:37 | 62 | 20140009.044 | 898 | 0.788 | 3418 | 15.117 | 2947 | -15.15 |  | 10.7398 | 11.0947 | 49.0312 | 50.8978 | 4.5654 | 14.6953 | -29.302
2014/03/05 00:40:07 | 63 | 20140009.045 | 705 | 0.75 | 3238 | 14.189 | 2666 | -15.231 |  | 10.6663 | 11.0269 | 46.2732 | 48.1699 | 4.3383 | 13.7688 | -29.3851
2014/03/05 00:48:36 | 64 | 20140009.046 | 711 | 0.786 | 3620 | 14.695 | 2724 | -14.786 |  | 11.3888 | 11.7552 | 45.2148 | 47.1416 | 3.9701 | 14.274 | -28.9283
2014/03/05 00:57:06 | 65 | 20140009.047 | 606 | 0.845 | 3477 | 12.567 | 2701 | -14.327 |  | 10.2102 | 10.5823 | 41.8711 | 43.8279 | 4.1009 | 12.1494 | -28.4571
2014/03/05 01:05:35 | 66 | 20140009.048 | 613 | 0.811 | 3595 | 14.383 | 2824 | -15.17 |  | 10.9079 | 11.2857 | 45.2851 | 47.272 | 4.1516 | 13.9625 | -29.3225
2014/03/05 02:05:02 | 73 | 20140009.049 | 1039 | 0.756 | 3351 | 14.733 | 2614 | -14.129 |  | 10.9966 | 11.4145 | 45.0646 | 47.2623 | 4.0981 | 14.3119 | -28.2538
2014/03/05 02:13:32 | 74 | 20140009.050 | 1035 | 0.78 | 3410 | 13.09 | 2804 | -15.266 |  | 10.8061 | 11.2297 | 47.0547 | 49.2825 | 4.3545 | 12.6716 | -29.4211
2014/03/05 02:22:02 | 75 | 20140009.051 | 1146 | 0.801 | 3825 | 15.469 | 2692 | -14.07 |  | 11.7855 | 12.2149 | 43.8488 | 46.1067 | 3.7206 | 15.0467 | -28.1933
2014/03/05 02:30:31 | 76 | 20140009.052 | 1132 | 0.797 | 3137 | 14.949 | 2162 | -14.068 |  | 9.7978 | 10.233 | 35.186 | 37.474 | 3.5912 | 14.5276 | -28.1912
2014/03/05 02:39:01 | 77 | 20140009.053 | 1242 | 0.777 | 3495 | 14.813 | 2676 | -15.053 |  | 11.1442 | 11.5851 | 45.0781 | 47.3962 | 4.045 | 14.3918 | -29.2024
2014/03/05 02:47:31 | 78 | 20140009.054 | 1299 | 0.772 | 3676 | 15.259 | 2506 | -14.308 |  | 11.8176 | 12.2642 | 42.3764 | 44.7246 | 3.5859 | 14.8371 | -28.4376
2014/03/05 02:56:01 | 79 | 20140009.055 | 1367 | 0.785 | 3912 | 14.582 | 2610 | -15.042 |  | 12.3594 | 12.8117 | 43.3991 | 45.7774 | 3.5114 | 14.1612 | -29.1911
2014/03/05 03:04:31 | 80 | 20140009.056 | 1358 | 0.756 | 3755 | 14.282 | 2487 | -14.749 |  | 12.3018 | 12.7598 | 42.8846 | 45.2931 | 3.486 | 13.8616 | -28.8903
2014/03/05 03:13:00 | 81 | 20140009.057 | 1543 | 0.796 | 3153 | 14.768 | 2878 | -17.446 |  | 9.7479 | 10.2116 | 47.0136 | 49.4522 | 4.823 | 14.3469 | -31.6591
2014/03/05 03:21:29 | 82 | 20140009.058 | 1550 | 0.787 | 3049 | 12.866 | 2720 | -15.47 |  | 9.5522 | 10.0216 | 44.8707 | 47.3394 | 4.6974 | 12.4479 | -29.6305
2014/03/05 03:29:59 | 83 | 20140009.059 | 1485 | 0.843 | 3790 | 12.716 | 2548 | -13.623 |  | 11.1467 | 11.6219 | 39.4591 | 41.9578 | 3.54 | 12.2982 | -27.7344
2014/03/05 03:38:30 | 84 | 20140009.060 | 1494 | 0.804 | 3631 | 14.866 | 2797 | -15.152 |  | 11.18 | 11.661 | 45.6014 | 48.1302 | 4.0788 | 14.4447 | -29.304
2014/03/05 03:47:00 | 85 | 20140009.061 | 141 | 0.787 | 3321 | 13.169 | 2399 | -15.432 |  | 10.4298 | 10.9164 | 39.5031 | 42.062 | 3.7875 | 12.7504 | -29.5915
2014/03/05 03:55:30 | 86 | 20140009.062 | 160 | 0.791 | 3525 | 13.015 | 2564 | -14.219 |  | 11.0307 | 11.523 | 42.1857 | 44.7748 | 3.8244 | 12.5967 | -28.3462
2014/03/05 04:04:00 | 87 | 20140009.063 | 238 | 0.834 | 3853 | 13.418 | 2832 | -14.529 |  | 11.3887 | 11.8868 | 44.3207 | 46.9399 | 3.8916 | 12.999 | -28.6645
2014/03/05 04:12:30 | 88 | 20140009.064 | 237 | 0.8 | 3722 | 13.813 | 2730 | -14.601 |  | 11.4965 | 12.0004 | 44.6187 | 47.268 | 3.8811 | 13.3934 | -28.7384
2014/03/05 04:21:00 | 89 | 20140009.065 | 361 | 0.837 | 3205 | 12.574 | 2147 | -13.609 |  | 9.5019 | 10.0114 | 33.2544 | 35.9338 | 3.4998 | 12.1564 | -27.72
2014/03/05 04:29:30 | 90 | 20140009.066 | 376 | 0.847 | 3737 | 13.409 | 2750 | -14.119 |  | 10.918 | 11.4333 | 42.4558 | 45.1653 | 3.8886 | 12.99 | -28.2436
2014/03/05 04:38:00 | 91 | 20140009.067 | 499 | 0.754 | 2985 | 13.09 | 2717 | -15.505 |  | 9.7639 | 10.2849 | 46.718 | 49.4576 | 4.7848 | 12.6716 | -29.6664
2014/03/05 04:46:30 | 92 | 20140009.068 | 483 | 0.8 | 3294 | 14.793 | 2966 | -16.387 |  | 10.1499 | 10.6766 | 48.3893 | 51.159 | 4.7675 | 14.3718 | -30.5719
2014/03/05 04:55:01 | 93 | 20140009.069 | 539 | 0.825 | 4039 | 14.315 | 2757 | -14.402 |  | 12.1627 | 12.6952 | 43.7933 | 46.5932 | 3.6006 | 13.8946 | -28.5341
2014/03/05 05:03:31 | 94 | 20140009.070 | 507 | 0.791 | 3984 | 14.801 | 2886 | -15.025 |  | 12.4407 | 12.9789 | 47.5819 | 50.4119 | 3.8247 | 14.3798 | -29.1737
2014/03/05 05:12:01 | 95 | 20140009.071 | 1601 | 0.842 | 3502 | 13.813 | 3063 | -15.334 |  | 10.2461 | 10.79 | 47.5419 | 50.402 | 4.64 | 13.3934 | -29.4909
2014/03/05 05:20:32 | 96 | 20140009.072 | 1602 | 0.842 | 3505 | 15.057 | 3005 | -15.443 |  | 10.2296 | 10.7792 | 46.6389 | 49.5291 | 4.5592 | 14.6354 | -29.6028
Reference Check
2014/03/04 17:19:11 | 11 | 01-UWSIF-Liver 20140009.11 | 01-UWSIF-Liver | 0.7472 | 3828 | 7.224 | 2795 | -3.785 | -14.035 | 13.0029 | 13.0659 | 49.768 | 50.0991 | 3.8275 | 6.8149 | -17.6347
2014/03/04 17:27:39 | 12 | 01-UWSIF-Liver 20140009.12 | 01-UWSIF-Liver | 0.7172 | 3665 | 7.18 | 2671 | -3.843 | -13.977 | 12.9522 | 13.0209 | 49.5924 | 49.9537 | 3.8289 | 6.771 | -17.6942
2014/03/04 21:33:23 | 41 | 01-UWSIF-Liver 20140009.13 | 01-UWSIF-Liver | 0.7646 | 3957 | 7.265 | 2861 | -3.836 | -13.984 | 12.7494 | 12.9841 | 48.5578 | 49.7922 | 3.8086 | 6.8559 | -17.687
2014/03/04 21:41:52 | 42 | 01-UWSIF-Liver 20140009.14 | 01-UWSIF-Liver | 0.7301 | 3765 | 7.205 | 2727 | -3.912 | -13.908 | 12.7507 | 12.9912 | 48.4668 | 49.7312 | 3.8011 | 6.796 | -17.7651
2014/03/05 01:48:03 | 71 | 01-UWSIF-Liver 20140009.15 | 01-UWSIF-Liver | 0.7261 | 3690 | 7.188 | 2655 | -3.818 | -14.002 | 12.5663 | 12.9728 | 47.4881 | 49.6256 | 3.779 | 6.779 | -17.6686
2014/03/05 01:56:32 | 72 | 01-UWSIF-Liver 20140009.16 | 01-UWSIF-Liver | 0.7155 | 3638 | 7.331 | 2627 | -3.91 | -13.91 | 12.5616 | 12.9738 | 47.5494 | 49.717 | 3.7853 | 6.9218 | -17.763
2014/03/05 06:03:02 | 101 | 01-UWSIF-Liver 20140009.17 | 01-UWSIF-Liver | 0.7536 | 3832 | 7.15 | 2744 | -3.864 | -13.956 | 12.4482 | 13.0264 | 46.9686 | 50.0093 | 3.7731 | 6.7411 | -17.7158
2014/03/05 06:11:33 | 102 | 01-UWSIF-Liver 20140009.18 | 01-UWSIF-Liver | 0.7296 | 3700 | 7.222 | 2657 | -3.959 | -13.861 | 12.4858 | 13.0698 | 47.0333 | 50.1041 | 3.7669 | 6.8129 | -17.8133
[5 empty] | average | 7.2206 |  | -3.8659 |  | 12.6896 | 13.0131 | 48.1781 | 49.879 |  | 6.8116 | -17.7177
[5 empty] | std. dev. | 0.056204 |  | 0.05732 |  | 0.20901 | 0.039135 | 1.0955 | 0.18575 |  | 0.056114 | 0.058844
Reference Material
2014/03/04 16:45:19 | 7 | 36-UWSIF-UT Glut 120140009.14 | 36-UWSIF-UT Glut 1 | 0.4665 | 1714 | -4.256 | 1407 | -13.863 | -14.416 | 9.5031 | 9.5431 | 39.3406 | 39.5513 | 4.1398 | -4.6467 | -27.9808
2014/03/04 16:53:47 | 8 | 36-UWSIF-UT Glut 120140009.15 | 36-UWSIF-UT Glut 1 | 0.7271 | 2686 | -4.221 | 2211 | -14.044 | -14.235 | 9.4896 | 9.5354 | 40.23 | 40.4709 | 4.2394 | -4.6117 | -28.1666
2014/03/04 20:59:28 | 37 | 36-UWSIF-UT Glut 120140009.16 | 36-UWSIF-UT Glut 1 | 0.764 | 2880 | -4.256 | 2340 | -14.145 | -14.134 | 9.34 | 9.5519 | 39.3844 | 40.4983 | 4.2167 | -4.6467 | -28.2703
2014/03/04 21:07:57 | 38 | 36-UWSIF-UT Glut 120140009.17 | 36-UWSIF-UT Glut 1 | 0.7864 | 2956 | -4.14 | 2398 | -14.182 | -14.097 | 9.3369 | 9.5544 | 39.4155 | 40.5596 | 4.2215 | -4.5309 | -28.3082
2014/03/05 01:14:05 | 67 | 36-UWSIF-UT Glut 120140009.18 | 36-UWSIF-UT Glut 1 | 0.7936 | 2948 | -4.239 | 2379 | -14.243 | -14.036 | 9.1701 | 9.5537 | 38.5975 | 40.6146 | 4.2091 | -4.6297 | -28.3709
2014/03/05 01:22:34 | 68 | 36-UWSIF-UT Glut 120140009.19 | 36-UWSIF-UT Glut 1 | 0.7865 | 2920 | -4.242 | 2366 | -14.203 | -14.076 | 9.1827 | 9.572 | 38.6144 | 40.6616 | 4.2051 | -4.6327 | -28.3298
2014/03/05 05:29:02 | 97 | 36-UWSIF-UT Glut 120140009.2 | 36-UWSIF-UT Glut 1 | 0.7704 | 2845 | -4.21 | 2294 | -14.266 | -14.013 | 9.0773 | 9.6327 | 38.0667 | 40.9869 | 4.1936 | -4.6008 | -28.3945
2014/03/05 05:37:32 | 98 | 36-UWSIF-UT Glut 120140009.21 | 36-UWSIF-UT Glut 1 | 0.7738 | 2851 | -4.238 | 2294 | -14.275 | -14.004 | 9.0783 | 9.6394 | 37.9928 | 40.9432 | 4.185 | -4.6287 | -28.4037
[5 empty] | average | -4.2253 |  | -14.1526 |  | 9.2723 | 9.5728 | 38.9552 | 40.5358 |  | -4.616 | -28.2781
[5 empty] | std. dev. | 0.037875 |  | 0.13894 |  | 0.17044 | 0.040428 | 0.76807 | 0.4419 |  | 0.037814 | 0.14264
2014/03/04 17:02:15 | 9 | 39-UWSIF-UW Glut 220140009.11 | 39-UWSIF-UW Glut 2 | 0.7405 | 2975 | 28.301 | 2436 | 37.313 | -12.951 | 10.2882 | 10.3397 | 43.6803 | 43.9512 | 4.2457 | 27.8582 | 24.5565
2014/03/04 17:10:43 | 10 | 39-UWSIF-UW Glut 220140009.12 | 39-UWSIF-UW Glut 2 | 0.7648 | 3078 | 28.331 | 2519 | 37.335 | -12.973 | 10.2817 | 10.3389 | 43.7447 | 44.0458 | 4.2546 | 27.8882 | 24.5791
2014/03/04 21:16:25 | 39 | 39-UWSIF-UW Glut 220140009.13 | 39-UWSIF-UW Glut 2 | 0.712 | 2880 | 28.354 | 2341 | 37.102 | -12.74 | 10.0756 | 10.2989 | 42.4736 | 43.6477 | 4.2155 | 27.9111 | 24.3399
2014/03/04 21:24:54 | 40 | 39-UWSIF-UW Glut 220140009.14 | 39-UWSIF-UW Glut 2 | 0.7464 | 3047 | 28.365 | 2472 | 37.036 | -12.674 | 10.1542 | 10.3832 | 42.8813 | 44.0855 | 4.223 | 27.9221 | 24.2722
2014/03/05 01:31:03 | 69 | 39-UWSIF-UW Glut 220140009.15 | 39-UWSIF-UW Glut 2 | 0.7345 | 2953 | 28.382 | 2383 | 36.967 | -12.605 | 9.979 | 10.374 | 41.8761 | 43.9534 | 4.1964 | 27.9391 | 24.2013
2014/03/05 01:39:33 | 70 | 39-UWSIF-UW Glut 220140009.16 | 39-UWSIF-UW Glut 2 | 0.754 | 3016 | 28.231 | 2441 | 37.12 | -12.758 | 9.8916 | 10.2924 | 41.7211 | 43.8285 | 4.2178 | 27.7883 | 24.3584
2014/03/05 05:46:02 | 99 | 39-UWSIF-UW Glut 220140009.17 | 39-UWSIF-UW Glut 2 | 0.7692 | 3093 | 28.354 | 2477 | 36.926 | -12.564 | 9.9117 | 10.4785 | 41.5063 | 44.4867 | 4.1876 | 27.9111 | 24.1592
2014/03/05 05:54:32 | 100 | 39-UWSIF-UW Glut 220140009.18 | 39-UWSIF-UW Glut 2 | 0.7779 | 3112 | 28.325 | 2494 | 37.178 | -12.816 | 9.8759 | 10.4485 | 41.3991 | 44.4096 | 4.1919 | 27.8822 | 24.4179
[5 empty] | average | 28.3304 |  | 37.1221 |  | 10.0572 | 10.3693 | 42.4103 | 44.0511 |  | 27.8875 | 24.3606
[5 empty] | std. dev. | 0.047431 |  | 0.14891 |  | 0.16947 | 0.06663 | 0.94261 | 0.28065 |  | 0.047355 | 0.15287
15N Normilization
|  |  | Meas. | Actual
|  | UWSIF 36- UT Glutamic | -4.2253 | -4.616
|  | UWSIF 39- UW Glutamic 2 | 28.3304 | 27.888
Lab QC Check
|  | Ref. Check | Corrected | Actual | 0.031572
|  | UWSIF01 (Liver) | 6.8116 | 6.78
Percentage
|  | UWSIF01 (Liver) | 13.0131 | 12.897
13C Normilization
|  |  | Meas. | Actual
|  | UWSIF 36- UT Glutamic | -14.1526 | -28.279
|  | UWSIF 39- UW Glutamic 2 | 37.1221 | 24.362
Lab QC Check
|  | Ref. Check | Corrected | Actual | 0.10229
|  | UWSIF01 (Liver) | -17.7177 | -17.82
Percentage
|  | UWSIF01 (Liver) | 49.879 | 49.537
Date Analyzed:3/5/14
Analyst:DEW.
[6 empty] | Instrument: | Costech EA 4010

## Original 1
Time Code | Row | Identifier 1 | Identifier 2 | Comment | Amount | Peak Nr | Ampl 28 | d 15N/14N | Ampl 44 | d 13C/12C | Amt% | Area All | Area 28 | BGD 28 | BGD 29 | BGD 30 | Area 44 | BGD 44 | BGD 45 | BGD 46 | AT% 13C/12C | AT% 15N/14N | rR 45CO2/44CO2 | rR 29N2/28N2
2014/03/04 16:01:58 | 1 | 36-UWSIF-UT Glut 120140009.11 | 36-UWSIF-UT Glut 1 | Job: 2014-0009 Run 1 Stockwell | 0.3915 | 1 | 4011 | 0.004 |  |  | 22.7422 | 79.099 | 78.498 | 21.4 | 15.9 | 331.0 | [5 empty] | 0.36647 |  | 0.72265
2014/03/04 16:01:58 | 1 | 36-UWSIF-UT Glut 120140009.11 | 36-UWSIF-UT Glut 1 | Job: 2014-0009 Run 1 Stockwell | 0.3915 | 2 | 4009 | 0 |  |  | 22.7297 | 79.055 | 78.454 | 21.5 | 16.0 | 351.1 | [5 empty] | 0.36647 |  | 0.72265
2014/03/04 16:01:58 | 1 | 36-UWSIF-UT Glut 120140009.11 | 36-UWSIF-UT Glut 1 | Job: 2014-0009 Run 1 Stockwell | 0.3915 | 3 | 1429 | -3.921 |  |  | 9.5353 | 33.164 | 32.914 | 21.5 | 16.0 | 337.6 | [5 empty] | 0.36504 |  | 0.71981
2014/03/04 16:01:58 | 1 | 36-UWSIF-UT Glut 120140009.11 | 36-UWSIF-UT Glut 1 | Job: 2014-0009 Run 1 Stockwell | 0.3915 | 4 |  |  | 1144 | -13.897 | 37.0404 | 26.892 |  |  |  |  | 26.474 | 0.2 | 0.3 | 2.2 | 1.0905 |  | 1.153
2014/03/04 16:01:58 | 1 | 36-UWSIF-UT Glut 120140009.11 | 36-UWSIF-UT Glut 1 | Job: 2014-0009 Run 1 Stockwell | 0.3915 | 5 |  |  | 4329 | 0 | 116.9443 | 84.904 |  |  |  |  | 83.575 | 3.0 | 3.6 | 6.2 | 1.1057 |  | 1.1678
2014/03/04 16:10:27 | 2 | 36-UWSIF-UT Glut 120140009.12 | 36-UWSIF-UT Glut 1 | Analyst: DEW | 0.7039 | 1 | 4000 | -0.009 |  |  | 12.6218 | 78.929 | 78.328 | 23.2 | 17.9 | 363.8 | [5 empty] | 0.36647 |  | 0.72303
2014/03/04 16:10:27 | 2 | 36-UWSIF-UT Glut 120140009.12 | 36-UWSIF-UT Glut 1 | Analyst: DEW | 0.7039 | 2 | 4000 | 0 |  |  | 12.6237 | 78.941 | 78.34 | 22.9 | 17.6 | 379.0 | [5 empty] | 0.36647 |  | 0.72304
2014/03/04 16:10:27 | 2 | 36-UWSIF-UT Glut 120140009.12 | 36-UWSIF-UT Glut 1 | Analyst: DEW | 0.7039 | 3 | 2619 | -4.208 |  |  | 9.5328 | 59.613 | 59.159 | 22.8 | 17.4 | 360.3 | [5 empty] | 0.36494 |  | 0.71999
2014/03/04 16:10:27 | 2 | 36-UWSIF-UT Glut 120140009.12 | 36-UWSIF-UT Glut 1 | Analyst: DEW | 0.7039 | 4 |  |  | 2160 | -13.998 | 39.689 | 51.808 |  |  |  |  | 51.003 | 1.1 | 1.4 | 3.6 | 1.0903 |  | 1.1527
2014/03/04 16:10:27 | 2 | 36-UWSIF-UT Glut 120140009.12 | 36-UWSIF-UT Glut 1 | Analyst: DEW | 0.7039 | 5 |  |  | 4326 | 0 | 65.2203 | 85.136 |  |  |  |  | 83.804 | 4.0 | 4.8 | 7.7 | 1.1057 |  | 1.1676
2014/03/04 16:17:00 | 3 |  |  | Sample Count: 0706 |  | 1 | 3996 | 0 |  |  | 0 | 78.824 | 78.224 | 23.8 | 18.5 | 379.0 | [5 empty] | 0.36647 |  | 0.72231
2014/03/04 16:17:00 | 3 |  |  | Sample Count: 0706 |  | 2 | 4001 | 0.041 |  |  | 0 | 78.835 | 78.235 | 23.2 | 17.8 | 336.6 | [5 empty] | 0.36649 |  | 0.72234
2014/03/04 16:17:00 | 3 |  |  | Sample Count: 0706 |  | 3 |  |  | 4317 | 0.282 | 0 | 84.635 |  |  |  |  | 83.311 | 2.4 | 2.9 | 5.5 | 1.106 |  | 1.1678
2014/03/04 16:17:00 | 3 |  |  | Sample Count: 0706 |  | 4 |  |  | 4318 | 0 | 0 | 84.947 |  |  |  |  | 83.618 | 5.0 | 6.0 | 9.1 | 1.1057 |  | 1.1674
2014/03/04 16:22:42 | 4 |  |  |  |  | 1 | 3990 | 0 |  |  | 0 | 78.72 | 78.118 | 24.8 | 19.6 | 312.7 | [5 empty] | 0.36647 |  | 0.72288
2014/03/04 16:22:42 | 4 |  |  |  |  | 2 | 3993 | 0.02 |  |  | 0 | 78.789 | 78.188 | 23.6 | 18.3 | 322.0 | [5 empty] | 0.36648 |  | 0.72289
2014/03/04 16:22:42 | 4 |  |  |  |  | 3 |  |  | 4308 | 0.255 | 0 | 84.551 |  |  |  |  | 83.228 | 2.6 | 3.1 | 5.7 | 1.1059 |  | 1.1677
2014/03/04 16:22:42 | 4 |  |  |  |  | 4 |  |  | 4312 | 0 | 0 | 84.947 |  |  |  |  | 83.618 | 5.0 | 6.0 | 9.1 | 1.1057 |  | 1.1675
2014/03/04 16:28:24 | 5 |  |  |  |  | 1 | 3988 | 0 |  |  | 0 | 78.691 | 78.09 | 24.9 | 19.7 | 308.0 | [5 empty] | 0.36647 |  | 0.72277
2014/03/04 16:28:24 | 5 |  |  |  |  | 2 | 3987 | 0.062 |  |  | 0 | 78.661 | 78.061 | 23.7 | 18.4 | 319.2 | [5 empty] | 0.3665 |  | 0.72281
2014/03/04 16:28:24 | 5 |  |  |  |  | 3 |  |  | 4310 | 0.186 | 0 | 84.563 |  |  |  |  | 83.239 | 2.6 | 3.1 | 5.7 | 1.1059 |  | 1.1677
2014/03/04 16:28:24 | 5 |  |  |  |  | 4 |  |  | 4310 | 0 | 0 | 84.857 |  |  |  |  | 83.529 | 5.0 | 5.9 | 9.1 | 1.1057 |  | 1.1675
2014/03/04 16:36:51 | 6 | 36-UWSIF-UT Glut 120140009.13 | 36-UWSIF-UT Glut 1 | Sample Count: | 1.0124 | 1 | 3982 | 0.017 |  |  |  | 78.576 | 77.976 | 24.2 | 18.9 | 385.6 | [5 empty] | 0.36648 |  | 0.72277
2014/03/04 16:36:51 | 6 | 36-UWSIF-UT Glut 120140009.13 | 36-UWSIF-UT Glut 1 | Sample Count: | 1.0124 | 2 | 3983 | 0 |  |  |  | 78.553 | 77.953 | 23.8 | 18.4 | 399.5 | [5 empty] | 0.36647 |  | 0.72276
2014/03/04 16:36:51 | 6 | 36-UWSIF-UT Glut 120140009.13 | 36-UWSIF-UT Glut 1 | Sample Count: | 1.0124 | 3 | 3794 | -4.088 |  |  | 9.52 | 85.545 | 84.89 | 23.6 | 18.1 | 378.5 | [5 empty] | 0.36498 |  | 0.7198
2014/03/04 16:36:51 | 6 | 36-UWSIF-UT Glut 120140009.13 | 36-UWSIF-UT Glut 1 | Sample Count: | 1.0124 | 4 |  |  | 3101 | -14.131 | 40.81 | 75.542 |  |  |  |  | 74.369 | 1.5 | 1.9 | 4.1 | 1.0902 |  | 1.1522
2014/03/04 16:36:51 | 6 | 36-UWSIF-UT Glut 120140009.13 | 36-UWSIF-UT Glut 1 | Sample Count: | 1.0124 | 5 |  |  | 4301 | 0 |  | 84.823 |  |  |  |  | 83.496 | 5.0 | 5.9 | 9.1 | 1.1057 |  | 1.1673
2014/03/04 16:45:19 | 7 | 36-UWSIF-UT Glut 120140009.14 | 36-UWSIF-UT Glut 1 | Helium Pressure: 1700 | 0.4665 | 1 | 3979 | -0.034 |  |  | 18.9661 | 78.53 | 77.931 | 24.3 | 19.0 | 391.1 | [5 empty] | 0.36646 |  | 0.72258
2014/03/04 16:45:19 | 7 | 36-UWSIF-UT Glut 120140009.14 | 36-UWSIF-UT Glut 1 | Helium Pressure: 1700 | 0.4665 | 2 | 3978 | 0 |  |  | 18.959 | 78.501 | 77.902 | 23.9 | 18.6 | 405.4 | [5 empty] | 0.36647 |  | 0.72261
2014/03/04 16:45:19 | 7 | 36-UWSIF-UT Glut 120140009.14 | 36-UWSIF-UT Glut 1 | Helium Pressure: 1700 | 0.4665 | 3 | 1714 | -4.256 |  |  | 9.5031 | 39.348 | 39.05 | 23.8 | 18.3 | 384.6 | [5 empty] | 0.36492 |  | 0.71953
2014/03/04 16:45:19 | 7 | 36-UWSIF-UT Glut 120140009.14 | 36-UWSIF-UT Glut 1 | Helium Pressure: 1700 | 0.4665 | 4 |  |  | 1407 | -13.863 | 39.3406 | 33.556 |  |  |  |  | 33.034 | 1.6 | 1.9 | 4.2 | 1.0905 |  | 1.1527
2014/03/04 16:45:19 | 7 | 36-UWSIF-UT Glut 120140009.14 | 36-UWSIF-UT Glut 1 | Helium Pressure: 1700 | 0.4665 | 5 |  |  | 4300 | 0 | 99.2524 | 84.657 |  |  |  |  | 83.333 | 3.9 | 4.7 | 7.6 | 1.1057 |  | 1.1675
2014/03/04 16:53:47 | 8 | 36-UWSIF-UT Glut 120140009.15 | 36-UWSIF-UT Glut 1 | Oxygen pressure: 600 | 0.7271 | 1 | 3975 | 0.024 |  |  | 12.1465 | 78.388 | 77.79 | 24.3 | 18.9 | 389.8 | [5 empty] | 0.36648 |  | 0.72259
2014/03/04 16:53:47 | 8 | 36-UWSIF-UT Glut 120140009.15 | 36-UWSIF-UT Glut 1 | Oxygen pressure: 600 | 0.7271 | 2 | 3980 | 0 |  |  | 12.1646 | 78.505 | 77.906 | 23.9 | 18.5 | 404.7 | [5 empty] | 0.36647 |  | 0.72257
2014/03/04 16:53:47 | 8 | 36-UWSIF-UT Glut 120140009.15 | 36-UWSIF-UT Glut 1 | Oxygen pressure: 600 | 0.7271 | 3 | 2686 | -4.221 |  |  | 9.4896 | 61.242 | 60.775 | 23.8 | 18.3 | 385.3 | [5 empty] | 0.36493 |  | 0.71952
2014/03/04 16:53:47 | 8 | 36-UWSIF-UT Glut 120140009.15 | 36-UWSIF-UT Glut 1 | Oxygen pressure: 600 | 0.7271 | 4 |  |  | 2211 | -14.044 | 40.23 | 53.483 |  |  |  |  | 52.652 | 1.5 | 1.8 | 4.2 | 1.0903 |  | 1.1524
2014/03/04 16:53:47 | 8 | 36-UWSIF-UT Glut 120140009.15 | 36-UWSIF-UT Glut 1 | Oxygen pressure: 600 | 0.7271 | 5 |  |  | 4292 | 0 | 63.6009 | 84.553 |  |  |  |  | 83.23 | 4.4 | 5.2 | 8.3 | 1.1057 |  | 1.1674
2014/03/04 17:02:15 | 9 | 39-UWSIF-UW Glut 220140009.11 | 39-UWSIF-UW Glut 2 | Mass 28:21 | 0.7405 | 1 | 3975 | -0.01 |  |  | 11.9352 | 78.444 | 77.846 | 24.4 | 19.0 | 395.6 | [5 empty] | 0.36647 |  | 0.72253
2014/03/04 17:02:15 | 9 | 39-UWSIF-UW Glut 220140009.11 | 39-UWSIF-UW Glut 2 | Mass 28:21 | 0.7405 | 2 | 3975 | 0 |  |  | 11.9316 | 78.421 | 77.822 | 24.0 | 18.6 | 410.1 | [5 empty] | 0.36647 |  | 0.72254
2014/03/04 17:02:15 | 9 | 39-UWSIF-UW Glut 220140009.11 | 39-UWSIF-UW Glut 2 | Mass 28:21 | 0.7405 | 3 | 2975 | 28.301 |  |  | 10.2882 | 67.619 | 67.087 | 23.9 | 18.4 | 390.0 | [5 empty] | 0.3768 |  | 0.74299
2014/03/04 17:02:15 | 9 | 39-UWSIF-UW Glut 220140009.11 | 39-UWSIF-UW Glut 2 | Mass 28:21 | 0.7405 | 4 |  |  | 2436 | 37.313 | 43.6803 | 59.14 |  |  |  |  | 58.189 | 1.6 | 1.9 | 4.2 | 1.1464 |  | 1.2087
2014/03/04 17:02:15 | 9 | 39-UWSIF-UW Glut 220140009.11 | 39-UWSIF-UW Glut 2 | Mass 28:21 | 0.7405 | 5 |  |  | 4286 | 0 | 62.3381 | 84.402 |  |  |  |  | 83.081 | 4.6 | 5.6 | 8.6 | 1.1057 |  | 1.1677
2014/03/04 17:10:43 | 10 | 39-UWSIF-UW Glut 220140009.12 | 39-UWSIF-UW Glut 2 | Mass 29:16 | 0.7648 | 1 | 3970 | -0.025 |  |  | 11.5406 | 78.34 | 77.742 | 24.5 | 19.2 | 399.6 | [5 empty] | 0.36646 |  | 0.72255
2014/03/04 17:10:43 | 10 | 39-UWSIF-UW Glut 220140009.12 | 39-UWSIF-UW Glut 2 | Mass 29:16 | 0.7648 | 2 | 3970 | 0 |  |  | 11.5413 | 78.344 | 77.747 | 24.1 | 18.7 | 414.3 | [5 empty] | 0.36647 |  | 0.72257
2014/03/04 17:10:43 | 10 | 39-UWSIF-UW Glut 220140009.12 | 39-UWSIF-UW Glut 2 | Mass 29:16 | 0.7648 | 3 | 3078 | 28.331 |  |  | 10.2817 | 69.794 | 69.244 | 23.9 | 18.5 | 393.8 | [5 empty] | 0.37682 |  | 0.74304
2014/03/04 17:10:43 | 10 | 39-UWSIF-UW Glut 220140009.12 | 39-UWSIF-UW Glut 2 | Mass 29:16 | 0.7648 | 4 |  |  | 2519 | 37.335 | 43.7447 | 61.171 |  |  |  |  | 60.187 | 1.6 | 2.0 | 4.3 | 1.1465 |  | 1.2088
2014/03/04 17:10:43 | 10 | 39-UWSIF-UW Glut 220140009.12 | 39-UWSIF-UW Glut 2 | Mass 29:16 | 0.7648 | 5 |  |  | 4286 | 0 | 60.3561 | 84.4 |  |  |  |  | 83.079 | 4.7 | 5.8 | 8.7 | 1.1057 |  | 1.1677
2014/03/04 17:19:11 | 11 | 01-UWSIF-Liver 20140009.11 | 01-UWSIF-Liver | Mass 30: 246 | 0.7472 | 1 | 3971 | -0.027 |  |  | 11.813 | 78.344 | 77.746 | 24.6 | 19.3 | 404.7 | [5 empty] | 0.36646 |  | 0.72239
2014/03/04 17:19:11 | 11 | 01-UWSIF-Liver 20140009.11 | 01-UWSIF-Liver | Mass 30: 246 | 0.7472 | 2 | 3970 | 0 |  |  | 11.8226 | 78.407 | 77.809 | 24.2 | 18.8 | 418.8 | [5 empty] | 0.36647 |  | 0.72241
2014/03/04 17:19:11 | 11 | 01-UWSIF-Liver 20140009.11 | 01-UWSIF-Liver | Mass 30: 246 | 0.7472 | 3 | 3828 | 7.224 |  |  | 13.0029 | 86.235 | 85.567 | 24.1 | 18.6 | 398.9 | [5 empty] | 0.36911 |  | 0.72763
2014/03/04 17:19:11 | 11 | 01-UWSIF-Liver 20140009.11 | 01-UWSIF-Liver | Mass 30: 246 | 0.7472 | 4 |  |  | 2795 | -3.785 | 49.768 | 67.992 |  |  |  |  | 66.928 | 1.6 | 2.0 | 4.3 | 1.1015 |  | 1.1636
2014/03/04 17:19:11 | 11 | 01-UWSIF-Liver 20140009.11 | 01-UWSIF-Liver | Mass 30: 246 | 0.7472 | 5 |  |  | 4284 | 0 | 61.7667 | 84.385 |  |  |  |  | 83.064 | 4.8 | 5.7 | 8.9 | 1.1057 |  | 1.1673
2014/03/04 17:27:39 | 12 | 01-UWSIF-Liver 20140009.12 | 01-UWSIF-Liver | Peak Center:3.065 | 0.7172 | 1 | 3969 | -0.035 |  |  | 12.2981 | 78.286 | 77.689 | 24.6 | 19.3 | 410.1 | [5 empty] | 0.36646 |  | 0.72246
2014/03/04 17:27:39 | 12 | 01-UWSIF-Liver 20140009.12 | 01-UWSIF-Liver | Peak Center:3.065 | 0.7172 | 2 | 3970 | 0 |  |  | 12.3109 | 78.367 | 77.769 | 24.2 | 18.8 | 424.8 | [5 empty] | 0.36647 |  | 0.72249
2014/03/04 17:27:39 | 12 | 01-UWSIF-Liver 20140009.12 | 01-UWSIF-Liver | Peak Center:3.065 | 0.7172 | 3 | 3665 | 7.18 |  |  | 12.9522 | 82.45 | 81.811 | 24.1 | 18.6 | 403.7 | [5 empty] | 0.36909 |  | 0.72768
2014/03/04 17:27:39 | 12 | 01-UWSIF-Liver 20140009.12 | 01-UWSIF-Liver | Peak Center:3.065 | 0.7172 | 4 |  |  | 2671 | -3.843 | 49.5924 | 65.032 |  |  |  |  | 64.015 | 1.6 | 2.0 | 4.3 | 1.1015 |  | 1.1633
2014/03/04 17:27:39 | 12 | 01-UWSIF-Liver 20140009.12 | 01-UWSIF-Liver | Peak Center:3.065 | 0.7172 | 5 |  |  | 4285 | 0 | 64.3375 | 84.368 |  |  |  |  | 83.048 | 4.8 | 5.7 | 8.9 | 1.1057 |  | 1.1671
2014/03/04 17:36:07 | 13 | 20140009.001 | 164 | Instrument: ANNIE | 0.858 | 1 | 3974 | -0.02 |  |  | 10.2912 | 78.372 | 77.774 | 24.6 | 19.3 | 415.3 | [5 empty] | 0.36646 |  | 0.7217
2014/03/04 17:36:07 | 13 | 20140009.001 | 164 | Instrument: ANNIE | 0.858 | 2 | 3975 | 0 |  |  | 10.2949 | 78.4 | 77.802 | 24.2 | 18.9 | 429.3 | [5 empty] | 0.36647 |  | 0.72171
2014/03/04 17:36:07 | 13 | 20140009.001 | 164 | Instrument: ANNIE | 0.858 | 3 | 3174 | 14.689 |  |  | 9.3066 | 70.874 | 70.323 | 24.1 | 18.6 | 409.0 | [5 empty] | 0.37184 |  | 0.73231
2014/03/04 17:36:07 | 13 | 20140009.001 | 164 | Instrument: ANNIE | 0.858 | 4 |  |  | 3425 | -18.53 | 53.4648 | 83.874 |  |  |  |  | 82.575 | 1.6 | 2.0 | 4.4 | 1.0854 |  | 1.1471
2014/03/04 17:36:07 | 13 | 20140009.001 | 164 | Instrument: ANNIE | 0.858 | 5 |  |  | 4282 | 0 | 53.8209 | 84.433 |  |  |  |  | 83.112 | 5.2 | 6.1 | 9.4 | 1.1057 |  | 1.1669
2014/03/04 17:44:35 | 14 | 20140009.002 | 139 | Data analyst: | 0.879 | 1 | 3969 | -0.016 |  |  | 10.0429 | 78.353 | 77.755 | 24.7 | 19.3 | 419.2 | [5 empty] | 0.36647 |  | 0.72235
2014/03/04 17:44:35 | 14 | 20140009.002 | 139 | Data analyst: | 0.879 | 2 | 3971 | 0 |  |  | 10.0394 | 78.325 | 77.728 | 24.3 | 18.9 | 433.4 | [5 empty] | 0.36647 |  | 0.72236
2014/03/04 17:44:35 | 14 | 20140009.002 | 139 | Data analyst: | 0.879 | 3 | 3376 | 14.6 |  |  | 9.6913 | 75.61 | 75.021 | 24.1 | 18.7 | 413.1 | [5 empty] | 0.3718 |  | 0.7329
2014/03/04 17:44:35 | 14 | 20140009.002 | 139 | Data analyst: | 0.879 | 4 |  |  | 3281 | -17.8 | 50.0331 | 80.412 |  |  |  |  | 79.165 | 1.7 | 2.0 | 4.4 | 1.0862 |  | 1.1479
2014/03/04 17:44:35 | 14 | 20140009.002 | 139 | Data analyst: | 0.879 | 5 |  |  | 4279 | 0 | 52.458 | 84.309 |  |  |  |  | 82.99 | 5.1 | 6.0 | 9.2 | 1.1057 |  | 1.1668
2014/03/04 17:53:03 | 15 | 20140009.003 | 225 | Comments: | 0.877 | 1 | 3968 | -0.023 |  |  | 10.0585 | 78.296 | 77.698 | 24.6 | 19.3 | 422.6 | [5 empty] | 0.36646 |  | 0.72219
2014/03/04 17:53:03 | 15 | 20140009.003 | 225 | Comments: | 0.877 | 2 | 3967 | 0 |  |  | 10.0562 | 78.278 | 77.681 | 24.2 | 18.8 | 436.6 | [5 empty] | 0.36647 |  | 0.72221
2014/03/04 17:53:03 | 15 | 20140009.003 | 225 | Comments: | 0.877 | 3 | 3572 | 14.048 |  |  | 10.2628 | 79.886 | 79.264 | 24.0 | 18.6 | 415.5 | [5 empty] | 0.3716 |  | 0.73235
2014/03/04 17:53:03 | 15 | 20140009.003 | 225 | Comments: | 0.877 | 4 |  |  | 3272 | -15.55 | 49.9727 | 80.132 |  |  |  |  | 78.888 | 1.7 | 2.1 | 4.4 | 1.0887 |  | 1.1502
2014/03/04 17:53:03 | 15 | 20140009.003 | 225 | Comments: | 0.877 | 5 |  |  | 4276 | 0 | 52.5501 | 84.265 |  |  |  |  | 82.947 | 5.1 | 6.0 | 9.4 | 1.1057 |  | 1.1668
2014/03/04 18:01:31 | 16 | 20140009.004 | 273 |  | 0.784 | 1 | 3966 | 0.015 |  |  | 11.2477 | 78.268 | 77.67 | 24.6 | 19.3 | 426.1 | [5 empty] | 0.36648 |  | 0.7223
2014/03/04 18:01:31 | 16 | 20140009.004 | 273 |  | 0.784 | 2 | 3969 | 0 |  |  | 11.2417 | 78.227 | 77.63 | 24.1 | 18.8 | 440.3 | [5 empty] | 0.36647 |  | 0.72229
2014/03/04 18:01:31 | 16 | 20140009.004 | 273 |  | 0.784 | 3 | 3356 | 13.77 |  |  | 10.8386 | 75.421 | 74.834 | 24.0 | 18.6 | 419.5 | [5 empty] | 0.3715 |  | 0.73223
2014/03/04 18:01:31 | 16 | 20140009.004 | 273 |  | 0.784 | 4 |  |  | 2876 | -15.107 | 48.8808 | 70.069 |  |  |  |  | 68.981 | 1.7 | 2.1 | 4.4 | 1.0891 |  | 1.1508
2014/03/04 18:01:31 | 16 | 20140009.004 | 273 |  | 0.784 | 5 |  |  | 4274 | 0 | 58.7415 | 84.204 |  |  |  |  | 82.887 | 4.9 | 5.7 | 9.0 | 1.1057 |  | 1.1668
2014/03/04 18:09:59 | 17 | 20140009.005 | 310 |  | 0.825 | 1 | 3969 | -0.001 |  |  | 10.7066 | 78.399 | 77.801 | 24.5 | 19.2 | 428.5 | [5 empty] | 0.36647 |  | 0.72151
2014/03/04 18:09:59 | 17 | 20140009.005 | 310 |  | 0.825 | 2 | 3969 | 0 |  |  | 10.6826 | 78.223 | 77.627 | 24.1 | 18.8 | 442.3 | [5 empty] | 0.36647 |  | 0.72151
2014/03/04 18:09:59 | 17 | 20140009.005 | 310 |  | 0.825 | 3 | 3513 | 13.72 |  |  | 10.695 | 78.314 | 77.705 | 23.9 | 18.5 | 420.7 | [5 empty] | 0.37148 |  | 0.73141
2014/03/04 18:09:59 | 17 | 20140009.005 | 310 |  | 0.825 | 4 |  |  | 3203 | -15.74 | 51.7735 | 78.097 |  |  |  |  | 76.885 | 1.7 | 2.1 | 4.4 | 1.0884 |  | 1.15
2014/03/04 18:09:59 | 17 | 20140009.005 | 310 |  | 0.825 | 5 |  |  | 4273 | 0 | 55.7975 | 84.167 |  |  |  |  | 82.85 | 5.1 | 6.0 | 9.3 | 1.1057 |  | 1.1667
2014/03/04 18:18:28 | 18 | 20140009.006 | 398 |  | 0.785 | 1 | 3967 | -0.037 |  |  | 11.2319 | 78.258 | 77.661 | 24.5 | 19.3 | 431.1 | [5 empty] | 0.36646 |  | 0.72158
2014/03/04 18:18:28 | 18 | 20140009.006 | 398 |  | 0.785 | 2 | 3967 | 0 |  |  | 11.2289 | 78.237 | 77.64 | 24.1 | 18.8 | 445.3 | [5 empty] | 0.36647 |  | 0.72161
2014/03/04 18:18:28 | 18 | 20140009.006 | 398 |  | 0.785 | 3 | 3381 | 14.81 |  |  | 10.9233 | 76.108 | 75.515 | 23.9 | 18.5 | 424.4 | [5 empty] | 0.37188 |  | 0.7323
2014/03/04 18:18:28 | 18 | 20140009.006 | 398 |  | 0.785 | 4 |  |  | 2718 | -14.655 | 46.0897 | 66.152 |  |  |  |  | 65.125 | 1.7 | 2.1 | 4.4 | 1.0896 |  | 1.1511
2014/03/04 18:18:28 | 18 | 20140009.006 | 398 |  | 0.785 | 5 |  |  | 4270 | 0 | 58.5844 | 84.086 |  |  |  |  | 82.771 | 4.8 | 5.6 | 8.9 | 1.1057 |  | 1.1666
2014/03/04 18:26:55 | 19 | 20140009.007 | 410 |  | 0.879 | 1 | 3967 | -0.029 |  |  | 10.0289 | 78.243 | 77.647 | 24.4 | 19.1 | 432.4 | [5 empty] | 0.36646 |  | 0.72129
2014/03/04 18:26:55 | 19 | 20140009.007 | 410 |  | 0.879 | 2 | 3969 | 0 |  |  | 10.0306 | 78.256 | 77.66 | 24.0 | 18.7 | 447.0 | [5 empty] | 0.36647 |  | 0.72131
2014/03/04 18:26:55 | 19 | 20140009.007 | 410 |  | 0.879 | 3 | 3359 | 15.416 |  |  | 9.5636 | 74.614 | 74.033 | 23.8 | 18.4 | 425.1 | [5 empty] | 0.3721 |  | 0.73243
2014/03/04 18:26:55 | 19 | 20140009.007 | 410 |  | 0.879 | 4 |  |  | 3319 | -16.847 | 50.2569 | 80.771 |  |  |  |  | 79.518 | 1.7 | 2.0 | 4.3 | 1.0872 |  | 1.1487
2014/03/04 18:26:55 | 19 | 20140009.007 | 410 |  | 0.879 | 5 |  |  | 4265 | 0 | 52.3012 | 84.057 |  |  |  |  | 82.742 | 5.1 | 6.0 | 9.4 | 1.1057 |  | 1.1666
2014/03/04 18:35:23 | 20 | 20140009.008 | 479 |  | 0.86 | 1 | 3961 | -0.051 |  |  | 10.2355 | 78.129 | 77.533 | 24.3 | 19.2 | 435.9 | [5 empty] | 0.36645 |  | 0.72205
2014/03/04 18:35:23 | 20 | 20140009.008 | 479 |  | 0.86 | 2 | 3962 | 0 |  |  | 10.2365 | 78.137 | 77.541 | 23.9 | 18.7 | 449.8 | [5 empty] | 0.36647 |  | 0.72208
2014/03/04 18:35:23 | 20 | 20140009.008 | 479 |  | 0.86 | 3 | 3820 | 14.243 |  |  | 11.1252 | 84.92 | 84.258 | 23.8 | 18.4 | 428.4 | [5 empty] | 0.37167 |  | 0.73237
2014/03/04 18:35:23 | 20 | 20140009.008 | 479 |  | 0.86 | 4 |  |  | 3203 | -15.301 | 49.572 | 77.948 |  |  |  |  | 76.738 | 1.7 | 2.1 | 4.4 | 1.0889 |  | 1.1503
2014/03/04 18:35:23 | 20 | 20140009.008 | 479 |  | 0.86 | 5 |  |  | 4263 | 0 | 53.4351 | 84.023 |  |  |  |  | 82.709 | 5.1 | 5.9 | 9.2 | 1.1057 |  | 1.1665
2014/03/04 18:43:51 | 21 | 20140009.009 | 544 |  | 0.801 | 1 | 3962 | 0.006 |  |  | 10.9986 | 78.194 | 77.598 | 24.3 | 19.1 | 439.7 | [5 empty] | 0.36647 |  | 0.72206
2014/03/04 18:43:51 | 21 | 20140009.009 | 544 |  | 0.801 | 2 | 3962 | 0 |  |  | 10.9906 | 78.137 | 77.541 | 23.9 | 18.7 | 453.8 | [5 empty] | 0.36647 |  | 0.72205
2014/03/04 18:43:51 | 21 | 20140009.009 | 544 |  | 0.801 | 3 | 3752 | 15.093 |  |  | 11.8052 | 83.929 | 83.274 | 23.7 | 18.4 | 432.5 | [5 empty] | 0.37198 |  | 0.73295
2014/03/04 18:43:51 | 21 | 20140009.009 | 544 |  | 0.801 | 4 |  |  | 2781 | -16.141 | 46.0368 | 67.423 |  |  |  |  | 66.377 | 1.7 | 2.1 | 4.4 | 1.088 |  | 1.1493
2014/03/04 18:43:51 | 21 | 20140009.009 | 544 |  | 0.801 | 5 |  |  | 4263 | 0 | 57.2823 | 83.893 |  |  |  |  | 82.581 | 4.8 | 5.7 | 8.9 | 1.1057 |  | 1.1664
2014/03/04 18:52:19 | 22 | 20140009.010 | 541 |  | 0.82 | 1 | 3964 | -0.033 |  |  | 10.7408 | 78.173 | 77.576 | 24.2 | 19.1 | 440.5 | [5 empty] | 0.36646 |  | 0.72138
2014/03/04 18:52:19 | 22 | 20140009.010 | 541 |  | 0.82 | 2 | 3963 | 0 |  |  | 10.7482 | 78.227 | 77.63 | 23.8 | 18.6 | 454.8 | [5 empty] | 0.36647 |  | 0.72141
2014/03/04 18:52:19 | 22 | 20140009.010 | 541 |  | 0.82 | 3 | 3924 | 13.622 |  |  | 11.9916 | 87.277 | 86.597 | 23.6 | 18.3 | 433.5 | [5 empty] | 0.37145 |  | 0.73124
2014/03/04 18:52:19 | 22 | 20140009.010 | 541 |  | 0.82 | 4 |  |  | 2760 | -13.933 | 44.4147 | 66.591 |  |  |  |  | 65.556 | 1.7 | 2.0 | 4.4 | 1.0904 |  | 1.1517
2014/03/04 18:52:19 | 22 | 20140009.010 | 541 |  | 0.82 | 5 |  |  | 4259 | 0 | 55.9399 | 83.87 |  |  |  |  | 82.559 | 4.8 | 5.6 | 8.8 | 1.1057 |  | 1.1664
2014/03/04 19:00:48 | 23 | 20140009.011 | 941 |  | 0.813 | 1 | 3962 | 0.028 |  |  | 10.8421 | 78.237 | 77.64 | 24.1 | 19.0 | 443.9 | [5 empty] | 0.36648 |  | 0.72131
2014/03/04 19:00:48 | 23 | 20140009.011 | 941 |  | 0.813 | 2 | 3962 | 0 |  |  | 10.8353 | 78.187 | 77.591 | 23.7 | 18.5 | 457.7 | [5 empty] | 0.36647 |  | 0.72129
2014/03/04 19:00:48 | 23 | 20140009.011 | 941 |  | 0.813 | 3 |  |  | 4260 | 0 | 55.9149 | 83.117 |  |  |  |  | 81.817 | 1.1 | 1.4 | 3.5 | 1.1057 |  | 1.167
2014/03/04 19:09:16 | 24 | 20140009.012 | 938 |  | 0.813 | 1 | 3965 | 0.033 |  |  | 10.8382 | 78.209 | 77.613 | 23.4 | 18.3 | 427.8 | [5 empty] | 0.36648 |  | 0.72129
2014/03/04 19:09:16 | 24 | 20140009.012 | 938 |  | 0.813 | 2 | 3966 | 0 |  |  | 10.8239 | 78.105 | 77.51 | 23.1 | 17.9 | 443.6 | [5 empty] | 0.36647 |  | 0.72126
2014/03/04 19:09:16 | 24 | 20140009.012 | 938 |  | 0.813 | 3 | 7170 | 14.154 |  |  | 21.4218 | 154.579 | 153.37 | 23.1 | 17.8 | 424.7 | [5 empty] | 0.37164 |  | 0.73147
2014/03/04 19:09:16 | 24 | 20140009.012 | 938 |  | 0.813 | 4 |  |  | 5972 | -14.739 | 99.7799 | 148.322 |  |  |  |  | 146.019 | 1.3 | 1.7 | 3.9 | 1.0895 |  | 1.1505
2014/03/04 19:09:16 | 24 | 20140009.012 | 938 |  | 0.813 | 5 |  |  | 4255 | 0 | 56.4125 | 83.857 |  |  |  |  | 82.546 | 6.2 | 7.2 | 10.8 | 1.1057 |  | 1.1662
2014/03/04 19:17:45 | 25 | 20140009.013 | 834 |  | 0.831 | 1 | 3961 | -0.05 |  |  | 10.5922 | 78.126 | 77.53 | 24.3 | 19.2 | 451.4 | [5 empty] | 0.36645 |  | 0.72119
2014/03/04 19:17:45 | 25 | 20140009.013 | 834 |  | 0.831 | 2 | 3961 | 0 |  |  | 10.5795 | 78.032 | 77.437 | 23.8 | 18.7 | 464.5 | [5 empty] | 0.36647 |  | 0.72123
2014/03/04 19:17:45 | 25 | 20140009.013 | 834 |  | 0.831 | 3 | 3950 | 13.514 |  |  | 11.8806 | 87.628 | 86.946 | 23.6 | 18.3 | 442.5 | [5 empty] | 0.37141 |  | 0.73097
2014/03/04 19:17:45 | 25 | 20140009.013 | 834 |  | 0.831 | 4 |  |  | 3048 | -14.887 | 48.5998 | 73.843 |  |  |  |  | 72.696 | 1.8 | 2.1 | 4.5 | 1.0894 |  | 1.1505
2014/03/04 19:17:45 | 25 | 20140009.013 | 834 |  | 0.831 | 5 |  |  | 4250 | 0 | 55.1418 | 83.783 |  |  |  |  | 82.473 | 4.9 | 5.8 | 9.1 | 1.1057 |  | 1.1662
2014/03/04 19:26:13 | 26 | 20140009.014 | 846 |  | 0.827 | 1 | 3953 | -0.033 |  |  | 10.6258 | 77.996 | 77.401 | 24.1 | 19.0 | 450.7 | [5 empty] | 0.36646 |  | 0.72183
2014/03/04 19:26:13 | 26 | 20140009.014 | 846 |  | 0.827 | 2 | 3952 | 0 |  |  | 10.6263 | 78 | 77.405 | 23.6 | 18.5 | 464.4 | [5 empty] | 0.36647 |  | 0.72185
2014/03/04 19:26:13 | 26 | 20140009.014 | 846 |  | 0.827 | 3 | 3530 | 13.816 |  |  | 10.6523 | 78.191 | 77.582 | 23.5 | 18.2 | 442.5 | [5 empty] | 0.37152 |  | 0.73183
2014/03/04 19:26:13 | 26 | 20140009.014 | 846 |  | 0.827 | 4 |  |  | 2965 | -14.618 | 47.2223 | 71.404 |  |  |  |  | 70.295 | 1.7 | 2.0 | 4.4 | 1.0897 |  | 1.1507
2014/03/04 19:26:13 | 26 | 20140009.014 | 846 |  | 0.827 | 5 |  |  | 4246 | 0 | 55.1712 | 83.424 |  |  |  |  | 82.12 | 4.9 | 5.7 | 9.0 | 1.1057 |  | 1.1662
2014/03/04 19:34:41 | 27 | 20140009.015 | 714 |  | 0.797 | 1 | 3958 | -0.014 |  |  | 11.038 | 78.083 | 77.488 | 24.0 | 18.9 | 450.2 | [5 empty] | 0.36647 |  | 0.72002
2014/03/04 19:34:41 | 27 | 20140009.015 | 714 |  | 0.797 | 2 | 3958 | 0 |  |  | 11.0362 | 78.07 | 77.476 | 23.5 | 18.4 | 464.4 | [5 empty] | 0.36647 |  | 0.72003
2014/03/04 19:34:41 | 27 | 20140009.015 | 714 |  | 0.797 | 3 | 3665 | 14.487 |  |  | 11.4566 | 81.044 | 80.414 | 23.4 | 18.1 | 442.8 | [5 empty] | 0.37176 |  | 0.73046
2014/03/04 19:34:41 | 27 | 20140009.015 | 714 |  | 0.797 | 4 |  |  | 2887 | -14.909 | 47.6863 | 69.49 |  |  |  |  | 68.411 | 1.7 | 2.0 | 4.3 | 1.0894 |  | 1.1504
2014/03/04 19:34:41 | 27 | 20140009.015 | 714 |  | 0.797 | 5 |  |  | 4245 | 0 | 57.3664 | 83.596 |  |  |  |  | 82.29 | 4.8 | 5.7 | 9.0 | 1.1057 |  | 1.1661
2014/03/04 19:43:10 | 28 | 20140009.016 | 718 |  | 0.786 | 1 | 3954 | -0.04 |  |  | 11.1811 | 78.003 | 77.409 | 23.9 | 18.9 | 452.4 | [5 empty] | 0.36646 |  | 0.7211
2014/03/04 19:43:10 | 28 | 20140009.016 | 718 |  | 0.786 | 2 | 3955 | 0 |  |  | 11.1788 | 77.987 | 77.392 | 23.5 | 18.3 | 466.2 | [5 empty] | 0.36647 |  | 0.72113
2014/03/04 19:43:10 | 28 | 20140009.016 | 718 |  | 0.786 | 3 | 3424 | 14.393 |  |  | 10.8518 | 75.706 | 75.117 | 23.3 | 18.1 | 444.4 | [5 empty] | 0.37173 |  | 0.73151
2014/03/04 19:43:10 | 28 | 20140009.016 | 718 |  | 0.786 | 4 |  |  | 2798 | -15.183 | 46.7928 | 67.247 |  |  |  |  | 66.203 | 1.7 | 2.0 | 4.3 | 1.0891 |  | 1.15
2014/03/04 19:43:10 | 28 | 20140009.016 | 718 |  | 0.786 | 5 |  |  | 4240 | 0 | 58.0933 | 83.487 |  |  |  |  | 82.182 | 4.8 | 5.6 | 8.8 | 1.1057 |  | 1.1661
2014/03/04 19:51:38 | 29 | 20140009.017 | 603 |  | 0.82 | 1 | 3955 | -0.011 |  |  | 10.7216 | 78.033 | 77.439 | 23.9 | 18.8 | 453.6 | [5 empty] | 0.36647 |  | 0.72002
2014/03/04 19:51:38 | 29 | 20140009.017 | 603 |  | 0.82 | 2 | 3956 | 0 |  |  | 10.7177 | 78.005 | 77.411 | 23.4 | 18.3 | 467.5 | [5 empty] | 0.36647 |  | 0.72003
2014/03/04 19:51:38 | 29 | 20140009.017 | 603 |  | 0.82 | 3 | 3572 | 14.337 |  |  | 10.8422 | 78.911 | 78.298 | 23.3 | 18.0 | 446.2 | [5 empty] | 0.37171 |  | 0.73035
2014/03/04 19:51:38 | 29 | 20140009.017 | 603 |  | 0.82 | 4 |  |  | 2869 | -15.204 | 45.9182 | 68.845 |  |  |  |  | 67.776 | 1.7 | 2.0 | 4.3 | 1.089 |  | 1.15
2014/03/04 19:51:38 | 29 | 20140009.017 | 603 |  | 0.82 | 5 |  |  | 4236 | 0 | 55.6648 | 83.458 |  |  |  |  | 82.153 | 4.8 | 5.6 | 8.8 | 1.1057 |  | 1.1661
2014/03/04 20:00:07 | 30 | 20140009.018 | 638 |  | 0.847 | 1 | 3953 | 0.02 |  |  | 10.372 | 77.974 | 77.381 | 23.9 | 18.8 | 455.7 | [5 empty] | 0.36648 |  | 0.71989
2014/03/04 20:00:07 | 30 | 20140009.018 | 638 |  | 0.847 | 2 | 3953 | 0 |  |  | 10.3723 | 77.977 | 77.383 | 23.4 | 18.3 | 469.8 | [5 empty] | 0.36647 |  | 0.71987
2014/03/04 20:00:07 | 30 | 20140009.018 | 638 |  | 0.847 | 3 | 4532 | 14.772 |  |  | 13.359 | 100.43 | 99.647 | 23.3 | 18.1 | 448.7 | [5 empty] | 0.37187 |  | 0.7305
2014/03/04 20:00:07 | 30 | 20140009.018 | 638 |  | 0.847 | 4 |  |  | 3082 | -14.495 | 48.2342 | 74.698 |  |  |  |  | 73.538 | 1.7 | 2.1 | 4.3 | 1.0898 |  | 1.1507
2014/03/04 20:00:07 | 30 | 20140009.018 | 638 |  | 0.847 | 5 |  |  | 4236 | 0 | 53.8503 | 83.396 |  |  |  |  | 82.092 | 4.9 | 5.8 | 9.0 | 1.1057 |  | 1.166
2014/03/04 20:08:35 | 31 | 20140009.019 | 1031 |  | 0.872 | 1 | 3948 | -0.041 |  |  | 10.0643 | 77.894 | 77.3 | 23.8 | 18.7 | 460.4 | [5 empty] | 0.36646 |  | 0.72101
2014/03/04 20:08:35 | 31 | 20140009.019 | 1031 |  | 0.872 | 2 | 3949 | 0 |  |  | 10.0605 | 77.865 | 77.272 | 23.3 | 18.2 | 474.2 | [5 empty] | 0.36647 |  | 0.72104
2014/03/04 20:08:35 | 31 | 20140009.019 | 1031 |  | 0.872 | 3 | 3689 | 14.796 |  |  | 10.5447 | 81.613 | 80.977 | 23.2 | 18.0 | 452.3 | [5 empty] | 0.37187 |  | 0.73171
2014/03/04 20:08:35 | 31 | 20140009.019 | 1031 |  | 0.872 | 4 |  |  | 2966 | -15.18 | 44.9028 | 71.592 |  |  |  |  | 70.48 | 1.7 | 2.0 | 4.4 | 1.0891 |  | 1.1499
2014/03/04 20:08:35 | 31 | 20140009.019 | 1031 |  | 0.872 | 5 |  |  | 4229 | 0 | 52.2525 | 83.31 |  |  |  |  | 82.008 | 4.9 | 5.7 | 8.9 | 1.1057 |  | 1.1659
2014/03/04 20:17:04 | 32 | 20140009.020 | 1003 |  | 0.82 | 1 | 3951 | -0.026 |  |  | 10.6992 | 77.87 | 77.277 | 23.7 | 18.7 | 460.9 | [5 empty] | 0.36646 |  | 0.71999
2014/03/04 20:17:04 | 32 | 20140009.020 | 1003 |  | 0.82 | 2 | 3949 | 0 |  |  | 10.7032 | 77.899 | 77.306 | 23.3 | 18.2 | 474.8 | [5 empty] | 0.36647 |  | 0.72001
2014/03/04 20:17:04 | 32 | 20140009.020 | 1003 |  | 0.82 | 3 | 3229 | 13.695 |  |  | 9.8514 | 71.7 | 71.144 | 23.1 | 17.9 | 453.3 | [5 empty] | 0.37147 |  | 0.72987
2014/03/04 20:17:04 | 32 | 20140009.020 | 1003 |  | 0.82 | 4 |  |  | 2481 | -14.687 | 39.5431 | 59.287 |  |  |  |  | 58.366 | 1.7 | 2.0 | 4.4 | 1.0896 |  | 1.1504
2014/03/04 20:17:04 | 32 | 20140009.020 | 1003 |  | 0.82 | 5 |  |  | 4226 | 0 | 55.5147 | 83.233 |  |  |  |  | 81.932 | 4.6 | 5.4 | 8.5 | 1.1057 |  | 1.1659
2014/03/04 20:25:33 | 33 | 20140009.021 | 1115 |  | 0.852 | 1 | 3948 | -0.046 |  |  | 10.3013 | 77.9 | 77.307 | 23.6 | 18.6 | 460.7 | [5 empty] | 0.36645 |  | 0.71986
2014/03/04 20:25:33 | 33 | 20140009.021 | 1115 |  | 0.852 | 2 | 3948 | 0 |  |  | 10.2949 | 77.852 | 77.259 | 23.2 | 18.1 | 474.8 | [5 empty] | 0.36647 |  | 0.71989
2014/03/04 20:25:33 | 33 | 20140009.021 | 1115 |  | 0.852 | 3 | 3881 | 14.345 |  |  | 11.3163 | 85.576 | 84.91 | 23.0 | 17.8 | 452.9 | [5 empty] | 0.37171 |  | 0.73022
2014/03/04 20:25:33 | 33 | 20140009.021 | 1115 |  | 0.852 | 4 |  |  | 2920 | -13.891 | 45.0333 | 70.153 |  |  |  |  | 69.063 | 1.7 | 2.0 | 4.3 | 1.0905 |  | 1.1512
2014/03/04 20:25:33 | 33 | 20140009.021 | 1115 |  | 0.852 | 5 |  |  | 4223 | 0 | 53.4024 | 83.19 |  |  |  |  | 81.89 | 4.8 | 5.7 | 8.8 | 1.1057 |  | 1.1659
2014/03/04 20:34:02 | 34 | 20140009.022 | 1128 |  | 0.795 | 1 | 3945 | -0.001 |  |  | 11.024 | 77.788 | 77.195 | 23.6 | 18.6 | 463.3 | [5 empty] | 0.36647 |  | 0.71981
2014/03/04 20:34:02 | 34 | 20140009.022 | 1128 |  | 0.795 | 2 | 3945 | 0 |  |  | 11.0241 | 77.789 | 77.197 | 23.1 | 18.1 | 477.3 | [5 empty] | 0.36647 |  | 0.71981
2014/03/04 20:34:02 | 34 | 20140009.022 | 1128 |  | 0.795 | 3 | 3809 | 13.798 |  |  | 11.8786 | 83.818 | 83.166 | 23.0 | 17.8 | 455.4 | [5 empty] | 0.37151 |  | 0.72974
2014/03/04 20:34:02 | 34 | 20140009.022 | 1128 |  | 0.795 | 4 |  |  | 2765 | -15.182 | 45.5156 | 66.161 |  |  |  |  | 65.134 | 1.7 | 2.0 | 4.4 | 1.0891 |  | 1.1498
2014/03/04 20:34:02 | 34 | 20140009.022 | 1128 |  | 0.795 | 5 |  |  | 4221 | 0 | 57.2298 | 83.188 |  |  |  |  | 81.888 | 4.7 | 5.6 | 8.7 | 1.1057 |  | 1.1658
2014/03/04 20:42:30 | 35 | 20140009.023 | 1224 |  | 0.801 | 1 | 3945 | -0.023 |  |  | 10.9504 | 77.852 | 77.259 | 23.5 | 18.6 | 465.8 | [5 empty] | 0.36646 |  | 0.71973
2014/03/04 20:42:30 | 35 | 20140009.023 | 1224 |  | 0.801 | 2 | 3944 | 0 |  |  | 10.9442 | 77.808 | 77.215 | 23.1 | 18.0 | 479.4 | [5 empty] | 0.36647 |  | 0.71975
2014/03/04 20:42:30 | 35 | 20140009.023 | 1224 |  | 0.801 | 3 | 3865 | 14.638 |  |  | 12.0329 | 85.547 | 84.882 | 22.9 | 17.8 | 458.2 | [5 empty] | 0.37182 |  | 0.73029
2014/03/04 20:42:30 | 35 | 20140009.023 | 1224 |  | 0.801 | 4 |  |  | 2657 | -13.866 | 43.5834 | 63.83 |  |  |  |  | 62.839 | 1.7 | 2.0 | 4.3 | 1.0905 |  | 1.1512
2014/03/04 20:42:30 | 35 | 20140009.023 | 1224 |  | 0.801 | 5 |  |  | 4217 | 0 | 56.7042 | 83.046 |  |  |  |  | 81.748 | 4.7 | 5.5 | 8.6 | 1.1057 |  | 1.1658
2014/03/04 20:50:59 | 36 | 20140009.024 | 1291 |  | 0.778 | 1 | 3940 | 0.043 |  |  | 11.2546 | 77.717 | 77.125 | 23.4 | 18.5 | 466.8 | [5 empty] | 0.36649 |  | 0.71979
2014/03/04 20:50:59 | 36 | 20140009.024 | 1291 |  | 0.778 | 2 | 3940 | 0 |  |  | 11.2582 | 77.742 | 77.15 | 23.0 | 18.0 | 480.8 | [5 empty] | 0.36647 |  | 0.71976
2014/03/04 20:50:59 | 36 | 20140009.024 | 1291 |  | 0.778 | 3 | 3749 | 14.378 |  |  | 11.9862 | 82.769 | 82.125 | 22.8 | 17.7 | 458.4 | [5 empty] | 0.37172 |  | 0.73011
2014/03/04 20:50:59 | 36 | 20140009.024 | 1291 |  | 0.778 | 4 |  |  | 2660 | -14.796 | 44.822 | 63.759 |  |  |  |  | 62.769 | 1.7 | 2.0 | 4.3 | 1.0895 |  | 1.1501
2014/03/04 20:50:59 | 36 | 20140009.024 | 1291 |  | 0.778 | 5 |  |  | 4215 | 0 | 58.3652 | 83.024 |  |  |  |  | 81.727 | 4.7 | 5.5 | 8.7 | 1.1057 |  | 1.1658
2014/03/04 20:59:28 | 37 | 36-UWSIF-UT Glut 120140009.16 | 36-UWSIF-UT Glut 1 |  | 0.764 | 1 | 3942 | -0.039 |  |  | 11.4663 | 77.754 | 77.162 | 23.3 | 18.4 | 468.6 | [5 empty] | 0.36646 |  | 0.71966
2014/03/04 20:59:28 | 37 | 36-UWSIF-UT Glut 120140009.16 | 36-UWSIF-UT Glut 1 |  | 0.764 | 2 | 3940 | 0 |  |  | 11.4626 | 77.729 | 77.138 | 22.9 | 17.9 | 482.4 | [5 empty] | 0.36647 |  | 0.71969
2014/03/04 20:59:28 | 37 | 36-UWSIF-UT Glut 120140009.16 | 36-UWSIF-UT Glut 1 |  | 0.764 | 3 | 2880 | -4.256 |  |  | 9.34 | 63.336 | 62.853 | 22.7 | 17.7 | 460.8 | [5 empty] | 0.36492 |  | 0.71663
2014/03/04 20:59:28 | 37 | 36-UWSIF-UT Glut 120140009.16 | 36-UWSIF-UT Glut 1 |  | 0.764 | 4 |  |  | 2340 | -14.145 | 39.3844 | 55.016 |  |  |  |  | 54.162 | 1.7 | 2.0 | 4.2 | 1.0902 |  | 1.1509
2014/03/04 20:59:28 | 37 | 36-UWSIF-UT Glut 120140009.16 | 36-UWSIF-UT Glut 1 |  | 0.764 | 5 |  |  | 4215 | 0 | 59.4396 | 83.031 |  |  |  |  | 81.734 | 4.4 | 5.2 | 8.2 | 1.1057 |  | 1.1658
2014/03/04 21:07:57 | 38 | 36-UWSIF-UT Glut 120140009.17 | 36-UWSIF-UT Glut 1 |  | 0.7864 | 1 | 3941 | -0.014 |  |  | 11.1429 | 77.776 | 77.184 | 23.2 | 18.3 | 467.8 | [5 empty] | 0.36647 |  | 0.71989
2014/03/04 21:07:57 | 38 | 36-UWSIF-UT Glut 120140009.17 | 36-UWSIF-UT Glut 1 |  | 0.7864 | 2 | 3941 | 0 |  |  | 11.134 | 77.714 | 77.123 | 22.8 | 17.8 | 481.5 | [5 empty] | 0.36647 |  | 0.7199
2014/03/04 21:07:57 | 38 | 36-UWSIF-UT Glut 120140009.17 | 36-UWSIF-UT Glut 1 |  | 0.7864 | 3 | 2956 | -4.14 |  |  | 9.3369 | 65.171 | 64.674 | 22.6 | 17.6 | 459.4 | [5 empty] | 0.36496 |  | 0.71692
2014/03/04 21:07:57 | 38 | 36-UWSIF-UT Glut 120140009.17 | 36-UWSIF-UT Glut 1 |  | 0.7864 | 4 |  |  | 2398 | -14.182 | 39.4155 | 56.674 |  |  |  |  | 55.794 | 1.6 | 2.0 | 4.2 | 1.0901 |  | 1.1507
2014/03/04 21:07:57 | 38 | 36-UWSIF-UT Glut 120140009.17 | 36-UWSIF-UT Glut 1 |  | 0.7864 | 5 |  |  | 4214 | 0 | 57.7243 | 82.999 |  |  |  |  | 81.702 | 4.4 | 5.2 | 8.3 | 1.1057 |  | 1.1657
2014/03/04 21:16:25 | 39 | 39-UWSIF-UW Glut 220140009.13 | 39-UWSIF-UW Glut 2 |  | 0.712 | 1 | 3939 | 0.002 |  |  | 12.2875 | 77.651 | 77.06 | 23.2 | 18.3 | 467.6 | [5 empty] | 0.36647 |  | 0.71981
2014/03/04 21:16:25 | 39 | 39-UWSIF-UW Glut 220140009.13 | 39-UWSIF-UW Glut 2 |  | 0.712 | 2 | 3939 | 0 |  |  | 12.2974 | 77.714 | 77.122 | 22.7 | 17.8 | 481.5 | [5 empty] | 0.36647 |  | 0.71981
2014/03/04 21:16:25 | 39 | 39-UWSIF-UW Glut 220140009.13 | 39-UWSIF-UW Glut 2 |  | 0.712 | 3 | 2880 | 28.354 |  |  | 10.0756 | 63.673 | 63.174 | 22.5 | 17.5 | 460.3 | [5 empty] | 0.37682 |  | 0.74022
2014/03/04 21:16:25 | 39 | 39-UWSIF-UW Glut 220140009.13 | 39-UWSIF-UW Glut 2 |  | 0.712 | 4 |  |  | 2341 | 37.102 | 42.4736 | 55.293 |  |  |  |  | 54.404 | 1.6 | 2.0 | 4.3 | 1.1462 |  | 1.207
2014/03/04 21:16:25 | 39 | 39-UWSIF-UW Glut 220140009.13 | 39-UWSIF-UW Glut 2 |  | 0.712 | 5 |  |  | 4211 | 0 | 63.6484 | 82.859 |  |  |  |  | 81.564 | 4.4 | 5.3 | 8.1 | 1.1057 |  | 1.1662
2014/03/04 21:24:54 | 40 | 39-UWSIF-UW Glut 220140009.14 | 39-UWSIF-UW Glut 2 |  | 0.7464 | 1 | 3938 | -0.034 |  |  | 11.7265 | 77.686 | 77.095 | 23.1 | 18.3 | 467.5 | [5 empty] | 0.36646 |  | 0.71979
2014/03/04 21:24:54 | 40 | 39-UWSIF-UW Glut 220140009.14 | 39-UWSIF-UW Glut 2 |  | 0.7464 | 2 | 3938 | 0 |  |  | 11.7247 | 77.675 | 77.083 | 22.7 | 17.8 | 481.8 | [5 empty] | 0.36647 |  | 0.71981
2014/03/04 21:24:54 | 40 | 39-UWSIF-UW Glut 220140009.14 | 39-UWSIF-UW Glut 2 |  | 0.7464 | 3 | 3047 | 28.365 |  |  | 10.1542 | 67.27 | 66.741 | 22.5 | 17.5 | 460.3 | [5 empty] | 0.37683 |  | 0.74023
2014/03/04 21:24:54 | 40 | 39-UWSIF-UW Glut 220140009.14 | 39-UWSIF-UW Glut 2 |  | 0.7464 | 4 |  |  | 2472 | 37.036 | 42.8813 | 58.521 |  |  |  |  | 57.58 | 1.6 | 2.0 | 4.2 | 1.1461 |  | 1.207
2014/03/04 21:24:54 | 40 | 39-UWSIF-UW Glut 220140009.14 | 39-UWSIF-UW Glut 2 |  | 0.7464 | 5 |  |  | 4209 | 0 | 60.762 | 82.923 |  |  |  |  | 81.627 | 4.4 | 5.4 | 8.3 | 1.1057 |  | 1.1662
2014/03/04 21:33:23 | 41 | 01-UWSIF-Liver 20140009.13 | 01-UWSIF-Liver |  | 0.7646 | 1 | 3937 | 0.035 |  |  | 11.4531 | 77.726 | 77.134 | 23.1 | 18.3 | 469.6 | [5 empty] | 0.36649 |  | 0.71964
2014/03/04 21:33:23 | 41 | 01-UWSIF-Liver 20140009.13 | 01-UWSIF-Liver |  | 0.7646 | 2 | 3939 | 0 |  |  | 11.4462 | 77.679 | 77.088 | 22.6 | 17.8 | 483.4 | [5 empty] | 0.36647 |  | 0.71962
2014/03/04 21:33:23 | 41 | 01-UWSIF-Liver 20140009.13 | 01-UWSIF-Liver |  | 0.7646 | 3 | 3957 | 7.265 |  |  | 12.7494 | 86.523 | 85.854 | 22.4 | 17.5 | 461.7 | [5 empty] | 0.36912 |  | 0.72485
2014/03/04 21:33:23 | 41 | 01-UWSIF-Liver 20140009.13 | 01-UWSIF-Liver |  | 0.7646 | 4 |  |  | 2861 | -3.836 | 48.5578 | 67.884 |  |  |  |  | 66.822 | 1.6 | 2.0 | 4.2 | 1.1015 |  | 1.1621
2014/03/04 21:33:23 | 41 | 01-UWSIF-Liver 20140009.13 | 01-UWSIF-Liver |  | 0.7646 | 5 |  |  | 4205 | 0 | 59.1725 | 82.723 |  |  |  |  | 81.43 | 4.7 | 5.5 | 8.6 | 1.1057 |  | 1.1658
2014/03/04 21:41:52 | 42 | 01-UWSIF-Liver 20140009.14 | 01-UWSIF-Liver |  | 0.7301 | 1 | 3934 | -0.084 |  |  | 11.9797 | 77.631 | 77.04 | 23.1 | 18.3 | 473.1 | [5 empty] | 0.36644 |  | 0.71981
2014/03/04 21:41:52 | 42 | 01-UWSIF-Liver 20140009.14 | 01-UWSIF-Liver |  | 0.7301 | 2 | 3935 | 0 |  |  | 11.9789 | 77.625 | 77.034 | 22.7 | 17.8 | 486.5 | [5 empty] | 0.36647 |  | 0.71987
2014/03/04 21:41:52 | 42 | 01-UWSIF-Liver 20140009.14 | 01-UWSIF-Liver |  | 0.7301 | 3 | 3765 | 7.205 |  |  | 12.7507 | 82.627 | 81.989 | 22.4 | 17.5 | 464.4 | [5 empty] | 0.3691 |  | 0.72505
2014/03/04 21:41:52 | 42 | 01-UWSIF-Liver 20140009.14 | 01-UWSIF-Liver |  | 0.7301 | 4 |  |  | 2727 | -3.912 | 48.4668 | 64.699 |  |  |  |  | 63.687 | 1.7 | 2.0 | 4.3 | 1.1014 |  | 1.1619
2014/03/04 21:41:52 | 42 | 01-UWSIF-Liver 20140009.14 | 01-UWSIF-Liver |  | 0.7301 | 5 |  |  | 4201 | 0 | 62.013 | 82.782 |  |  |  |  | 81.489 | 4.6 | 5.5 | 8.4 | 1.1057 |  | 1.1657
2014/03/04 21:50:21 | 43 | 20140009.025 | 1331 |  | 0.829 | 1 | 3933 | 0.027 |  |  | 10.5496 | 77.624 | 77.033 | 23.0 | 18.2 | 473.5 | [5 empty] | 0.36648 |  | 0.71961
2014/03/04 21:50:21 | 43 | 20140009.025 | 1331 |  | 0.829 | 2 | 3934 | 0 |  |  | 10.5514 | 77.637 | 77.046 | 22.5 | 17.7 | 487.5 | [5 empty] | 0.36647 |  | 0.71959
2014/03/04 21:50:21 | 43 | 20140009.025 | 1331 |  | 0.829 | 3 | 3993 | 14.585 |  |  | 11.9937 | 88.25 | 87.563 | 22.4 | 17.4 | 466.1 | [5 empty] | 0.3718 |  | 0.73009
2014/03/04 21:50:21 | 43 | 20140009.025 | 1331 |  | 0.829 | 4 |  |  | 2785 | -14.563 | 44.171 | 66.952 |  |  |  |  | 65.913 | 1.7 | 2.0 | 4.3 | 1.0897 |  | 1.1501
2014/03/04 21:50:21 | 43 | 20140009.025 | 1331 |  | 0.829 | 5 |  |  | 4200 | 0 | 54.5865 | 82.739 |  |  |  |  | 81.447 | 4.6 | 5.5 | 8.5 | 1.1057 |  | 1.1655
2014/03/04 21:58:50 | 44 | 20140009.026 | 1369 |  | 0.847 | 1 | 3931 | -0.062 |  |  | 10.314 | 77.538 | 76.948 | 23.0 | 18.2 | 475.7 | [5 empty] | 0.36645 |  | 0.71975
2014/03/04 21:58:50 | 44 | 20140009.026 | 1369 |  | 0.847 | 2 | 3932 | 0 |  |  | 10.3126 | 77.528 | 76.938 | 22.5 | 17.7 | 489.7 | [5 empty] | 0.36647 |  | 0.71979
2014/03/04 21:58:50 | 44 | 20140009.026 | 1369 |  | 0.847 | 3 | 4206 | 14.958 |  |  | 12.2551 | 92.131 | 91.413 | 22.3 | 17.4 | 467.3 | [5 empty] | 0.37193 |  | 0.73056
2014/03/04 21:58:50 | 44 | 20140009.026 | 1369 |  | 0.847 | 4 |  |  | 2855 | -14.744 | 44.0081 | 68.153 |  |  |  |  | 67.096 | 1.7 | 2.0 | 4.3 | 1.0895 |  | 1.1499
2014/03/04 21:58:50 | 44 | 20140009.026 | 1369 |  | 0.847 | 5 |  |  | 4199 | 0 | 53.3942 | 82.689 |  |  |  |  | 81.397 | 4.7 | 5.5 | 8.6 | 1.1057 |  | 1.1655
2014/03/04 22:07:19 | 45 | 20140009.027 | 1522 |  | 0.831 | 1 | 3931 | -0.04 |  |  | 10.5186 | 77.583 | 76.992 | 23.0 | 18.2 | 478.3 | [5 empty] | 0.36646 |  | 0.71963
2014/03/04 22:07:19 | 45 | 20140009.027 | 1522 |  | 0.831 | 2 | 3931 | 0 |  |  | 10.5123 | 77.536 | 76.946 | 22.5 | 17.7 | 492.2 | [5 empty] | 0.36647 |  | 0.71966
2014/03/04 22:07:19 | 45 | 20140009.027 | 1522 |  | 0.831 | 3 | 3588 | 15.039 |  |  | 10.6623 | 78.642 | 78.031 | 22.3 | 17.4 | 469.6 | [5 empty] | 0.37196 |  | 0.73048
2014/03/04 22:07:19 | 45 | 20140009.027 | 1522 |  | 0.831 | 4 |  |  | 3026 | -15.265 | 47.7622 | 72.57 |  |  |  |  | 71.444 | 1.7 | 2.0 | 4.3 | 1.089 |  | 1.1493
2014/03/04 22:07:19 | 45 | 20140009.027 | 1522 |  | 0.831 | 5 |  |  | 4194 | 0 | 54.394 | 82.646 |  |  |  |  | 81.355 | 4.8 | 5.6 | 8.8 | 1.1057 |  | 1.1654
2014/03/04 22:15:48 | 46 | 20140009.028 | 1515 |  | 0.83 | 1 | 3928 | -0.006 |  |  | 10.518 | 77.485 | 76.895 | 22.9 | 18.2 | 478.5 | [5 empty] | 0.36647 |  | 0.7199
2014/03/04 22:15:48 | 46 | 20140009.028 | 1515 |  | 0.83 | 2 | 3929 | 0 |  |  | 10.5164 | 77.473 | 76.883 | 22.5 | 17.7 | 492.5 | [5 empty] | 0.36647 |  | 0.71991
2014/03/04 22:15:48 | 46 | 20140009.028 | 1515 |  | 0.83 | 3 | 3446 | 14.954 |  |  | 10.2933 | 75.829 | 75.239 | 22.3 | 17.4 | 470.3 | [5 empty] | 0.37193 |  | 0.73067
2014/03/04 22:15:48 | 46 | 20140009.028 | 1515 |  | 0.83 | 4 |  |  | 3038 | -15.525 | 48.0637 | 72.94 |  |  |  |  | 71.809 | 1.7 | 2.0 | 4.4 | 1.0887 |  | 1.1488
2014/03/04 22:15:48 | 46 | 20140009.028 | 1515 |  | 0.83 | 5 |  |  | 4194 | 0 | 54.4256 | 82.595 |  |  |  |  | 81.305 | 4.8 | 5.6 | 8.8 | 1.1057 |  | 1.1653
2014/03/04 22:24:17 | 47 | 20140009.029 | 1497 |  | 0.777 | 1 | 3927 | 0.006 |  |  | 11.2382 | 77.504 | 76.914 | 22.9 | 18.2 | 479.1 | [5 empty] | 0.36647 |  | 0.71949
2014/03/04 22:24:17 | 47 | 20140009.029 | 1497 |  | 0.777 | 2 | 3927 | 0 |  |  | 11.2325 | 77.465 | 76.875 | 22.4 | 17.6 | 493.2 | [5 empty] | 0.36647 |  | 0.71949
2014/03/04 22:24:17 | 47 | 20140009.029 | 1497 |  | 0.777 | 3 | 3529 | 12.912 |  |  | 11.2742 | 77.752 | 77.149 | 22.2 | 17.3 | 470.6 | [5 empty] | 0.37119 |  | 0.72878
2014/03/04 22:24:17 | 47 | 20140009.029 | 1497 |  | 0.777 | 4 |  |  | 2484 | -14.008 | 41.7309 | 59.286 |  |  |  |  | 58.365 | 1.7 | 2.1 | 4.3 | 1.0903 |  | 1.1506
2014/03/04 22:24:17 | 47 | 20140009.029 | 1497 |  | 0.777 | 5 |  |  | 4189 | 0 | 58.1094 | 82.554 |  |  |  |  | 81.264 | 4.5 | 5.3 | 8.3 | 1.1057 |  | 1.1654
2014/03/04 22:32:46 | 48 | 20140009.030 | 1481 |  | 0.809 | 1 | 3925 | -0.019 |  |  | 10.7876 | 77.46 | 76.871 | 22.8 | 18.1 | 479.4 | [5 empty] | 0.36646 |  | 0.7196
2014/03/04 22:32:46 | 48 | 20140009.030 | 1481 |  | 0.809 | 2 | 3924 | 0 |  |  | 10.7837 | 77.432 | 76.843 | 22.3 | 17.5 | 493.4 | [5 empty] | 0.36647 |  | 0.71961
2014/03/04 22:32:46 | 48 | 20140009.030 | 1481 |  | 0.809 | 3 | 3865 | 13.881 |  |  | 11.8679 | 85.218 | 84.555 | 22.1 | 17.2 | 471.5 | [5 empty] | 0.37154 |  | 0.7296
2014/03/04 22:32:46 | 48 | 20140009.030 | 1481 |  | 0.809 | 4 |  |  | 2621 | -14.276 | 42.4204 | 62.747 |  |  |  |  | 61.773 | 1.6 | 2.0 | 4.3 | 1.09 |  | 1.1502
2014/03/04 22:32:46 | 48 | 20140009.030 | 1481 |  | 0.809 | 5 |  |  | 4189 | 0 | 55.7674 | 82.49 |  |  |  |  | 81.201 | 4.5 | 5.4 | 8.4 | 1.1057 |  | 1.1653
2014/03/04 22:41:15 | 49 | 20140009.031 | 111 |  | 0.83 | 1 | 3921 | 0.011 |  |  | 10.5019 | 77.366 | 76.777 | 22.7 | 18.0 | 481.2 | [5 empty] | 0.36648 |  | 0.71982
2014/03/04 22:41:15 | 49 | 20140009.031 | 111 |  | 0.83 | 2 | 3923 | 0 |  |  | 10.4992 | 77.346 | 76.757 | 22.3 | 17.5 | 495.2 | [5 empty] | 0.36647 |  | 0.71981
2014/03/04 22:41:15 | 49 | 20140009.031 | 111 |  | 0.83 | 3 | 3111 | 13.16 |  |  | 9.2706 | 68.295 | 67.766 | 22.1 | 17.2 | 472.2 | [5 empty] | 0.37128 |  | 0.72929
2014/03/04 22:41:15 | 49 | 20140009.031 | 111 |  | 0.83 | 4 |  |  | 2420 | -14.423 | 37.7246 | 57.25 |  |  |  |  | 56.361 | 1.7 | 2.0 | 4.3 | 1.0899 |  | 1.15
2014/03/04 22:41:15 | 49 | 20140009.031 | 111 |  | 0.83 | 5 |  |  | 4184 | 0 | 54.2887 | 82.387 |  |  |  |  | 81.1 | 4.4 | 5.2 | 8.3 | 1.1057 |  | 1.1652
2014/03/04 22:49:44 | 50 | 20140009.032 | 171 |  | 0.751 | 1 | 3921 | 0.001 |  |  | 11.6146 | 77.42 | 76.83 | 22.6 | 17.9 | 480.8 | [5 empty] | 0.36647 |  | 0.71969
2014/03/04 22:49:44 | 50 | 20140009.032 | 171 |  | 0.751 | 2 | 3922 | 0 |  |  | 11.61 | 77.389 | 76.8 | 22.2 | 17.4 | 494.6 | [5 empty] | 0.36647 |  | 0.71969
2014/03/04 22:49:44 | 50 | 20140009.032 | 171 |  | 0.751 | 3 | 3065 | 13.739 |  |  | 10.1045 | 67.354 | 66.831 | 22.0 | 17.1 | 473.1 | [5 empty] | 0.37149 |  | 0.72957
2014/03/04 22:49:44 | 50 | 20140009.032 | 171 |  | 0.751 | 4 |  |  | 2988 | -15.726 | 51.9199 | 71.293 |  |  |  |  | 70.187 | 1.6 | 2.0 | 4.3 | 1.0885 |  | 1.1486
2014/03/04 22:49:44 | 50 | 20140009.032 | 171 |  | 0.751 | 5 |  |  | 4182 | 0 | 59.9436 | 82.311 |  |  |  |  | 81.025 | 4.7 | 5.5 | 8.6 | 1.1057 |  | 1.1652
2014/03/04 22:58:13 | 51 | 20140009.033 | 248 |  | 0.774 | 1 | 3921 | -0.087 |  |  | 11.2609 | 77.36 | 76.771 | 22.6 | 18.0 | 481.2 | [5 empty] | 0.36644 |  | 0.71966
2014/03/04 22:58:13 | 51 | 20140009.033 | 248 |  | 0.774 | 2 | 3921 | 0 |  |  | 11.2544 | 77.316 | 76.727 | 22.2 | 17.4 | 495.5 | [5 empty] | 0.36647 |  | 0.71973
2014/03/04 22:58:13 | 51 | 20140009.033 | 248 |  | 0.774 | 3 | 3293 | 13.113 |  |  | 10.5542 | 72.506 | 71.943 | 21.9 | 17.1 | 472.7 | [5 empty] | 0.37126 |  | 0.72916
2014/03/04 22:58:13 | 51 | 20140009.033 | 248 |  | 0.774 | 4 |  |  | 2799 | -15.164 | 47.3562 | 67.018 |  |  |  |  | 65.978 | 1.7 | 2.0 | 4.3 | 1.0891 |  | 1.1491
2014/03/04 22:58:13 | 51 | 20140009.033 | 248 |  | 0.774 | 5 |  |  | 4180 | 0 | 58.1596 | 82.307 |  |  |  |  | 81.021 | 4.6 | 5.4 | 8.5 | 1.1057 |  | 1.1651
2014/03/04 23:06:43 | 52 | 20140009.034 | 300 |  | 0.782 | 1 | 3917 | 0.014 |  |  | 11.1319 | 77.265 | 76.676 | 22.5 | 17.9 | 482.4 | [5 empty] | 0.36648 |  | 0.7196
2014/03/04 23:06:43 | 52 | 20140009.034 | 300 |  | 0.782 | 2 | 3919 | 0 |  |  | 11.1341 | 77.28 | 76.691 | 22.1 | 17.4 | 496.4 | [5 empty] | 0.36647 |  | 0.71959
2014/03/04 23:06:43 | 52 | 20140009.034 | 300 |  | 0.782 | 3 | 3176 | 13.379 |  |  | 9.9963 | 69.383 | 68.844 | 21.9 | 17.1 | 473.6 | [5 empty] | 0.37136 |  | 0.72922
2014/03/04 23:06:43 | 52 | 20140009.034 | 300 |  | 0.782 | 4 |  |  | 2994 | -15.604 | 49.8962 | 71.342 |  |  |  |  | 70.236 | 1.7 | 2.0 | 4.3 | 1.0886 |  | 1.1486
2014/03/04 23:06:43 | 52 | 20140009.034 | 300 |  | 0.782 | 5 |  |  | 4178 | 0 | 57.5856 | 82.337 |  |  |  |  | 81.05 | 4.7 | 5.5 | 8.6 | 1.1057 |  | 1.1651
2014/03/04 23:15:12 | 53 | 20140009.035 | 341 |  | 0.844 | 1 | 3919 | -0.001 |  |  | 10.3257 | 77.352 | 76.762 | 22.5 | 17.9 | 483.8 | [5 empty] | 0.36647 |  | 0.71972
2014/03/04 23:15:12 | 53 | 20140009.035 | 341 |  | 0.844 | 2 | 3917 | 0 |  |  | 10.3089 | 77.226 | 76.638 | 22.1 | 17.4 | 497.3 | [5 empty] | 0.36647 |  | 0.71972
2014/03/04 23:15:12 | 53 | 20140009.035 | 341 |  | 0.844 | 3 | 3244 | 13.197 |  |  | 9.5647 | 71.651 | 71.095 | 21.9 | 17.1 | 475.8 | [5 empty] | 0.37129 |  | 0.72922
2014/03/04 23:15:12 | 53 | 20140009.035 | 341 |  | 0.844 | 4 |  |  | 2617 | -14.586 | 40.5348 | 62.552 |  |  |  |  | 61.582 | 1.7 | 2.0 | 4.3 | 1.0897 |  | 1.1497
2014/03/04 23:15:12 | 53 | 20140009.035 | 341 |  | 0.844 | 5 |  |  | 4175 | 0 | 53.2759 | 82.214 |  |  |  |  | 80.93 | 4.5 | 5.3 | 8.4 | 1.1057 |  | 1.1651
2014/03/04 23:23:41 | 54 | 20140009.036 | 324 |  | 0.839 | 1 | 3916 | -0.081 |  |  | 10.3777 | 77.281 | 76.692 | 22.5 | 17.9 | 483.9 | [5 empty] | 0.36644 |  | 0.71948
2014/03/04 23:23:41 | 54 | 20140009.036 | 324 |  | 0.839 | 2 | 3917 | 0 |  |  | 10.3774 | 77.278 | 76.69 | 22.0 | 17.3 | 498.0 | [5 empty] | 0.36647 |  | 0.71954
2014/03/04 23:23:41 | 54 | 20140009.036 | 324 |  | 0.839 | 3 | 3769 | 14.286 |  |  | 11.1033 | 82.684 | 82.041 | 21.8 | 17.0 | 475.8 | [5 empty] | 0.37169 |  | 0.72982
2014/03/04 23:23:41 | 54 | 20140009.036 | 324 |  | 0.839 | 4 |  |  | 3042 | -15.013 | 47.4783 | 72.833 |  |  |  |  | 71.703 | 1.7 | 2.0 | 4.3 | 1.0892 |  | 1.1492
2014/03/04 23:23:41 | 54 | 20140009.036 | 324 |  | 0.839 | 5 |  |  | 4172 | 0 | 53.577 | 82.189 |  |  |  |  | 80.905 | 4.7 | 5.5 | 8.6 | 1.1057 |  | 1.1651
2014/03/04 23:32:11 | 55 | 20140009.037 | 446 |  | 0.772 | 1 | 3914 | -0.012 |  |  | 11.2723 | 77.239 | 76.65 | 22.5 | 17.8 | 486.4 | [5 empty] | 0.36647 |  | 0.71959
2014/03/04 23:32:11 | 55 | 20140009.037 | 446 |  | 0.772 | 2 | 3912 | 0 |  |  | 11.272 | 77.237 | 76.649 | 22.0 | 17.3 | 500.2 | [5 empty] | 0.36647 |  | 0.7196
2014/03/04 23:32:11 | 55 | 20140009.037 | 446 |  | 0.772 | 3 | 3055 | 13.89 |  |  | 9.7846 | 67.045 | 66.525 | 21.8 | 17.0 | 477.5 | [5 empty] | 0.37154 |  | 0.72959
2014/03/04 23:32:11 | 55 | 20140009.037 | 446 |  | 0.772 | 4 |  |  | 2893 | -16.029 | 48.9672 | 69.119 |  |  |  |  | 68.047 | 1.7 | 2.0 | 4.4 | 1.0881 |  | 1.1481
2014/03/04 23:32:11 | 55 | 20140009.037 | 446 |  | 0.772 | 5 |  |  | 4172 | 0 | 58.1994 | 82.15 |  |  |  |  | 80.867 | 4.6 | 5.4 | 8.5 | 1.1057 |  | 1.165
2014/03/04 23:40:40 | 56 | 20140009.038 | 429 |  | 0.751 | 1 | 3911 | 0.002 |  |  | 11.5873 | 77.237 | 76.649 | 22.3 | 17.7 | 486.2 | [5 empty] | 0.36647 |  | 0.71968
2014/03/04 23:40:40 | 56 | 20140009.038 | 429 |  | 0.751 | 2 | 3911 | 0 |  |  | 11.5761 | 77.163 | 76.575 | 21.9 | 17.2 | 500.2 | [5 empty] | 0.36647 |  | 0.71968
2014/03/04 23:40:40 | 56 | 20140009.038 | 429 |  | 0.751 | 3 | 3065 | 13.459 |  |  | 10.105 | 67.357 | 66.834 | 21.6 | 16.9 | 477.1 | [5 empty] | 0.37139 |  | 0.72936
2014/03/04 23:40:40 | 56 | 20140009.038 | 429 |  | 0.751 | 4 |  |  | 2639 | -14.786 | 45.7624 | 62.838 |  |  |  |  | 61.863 | 1.7 | 2.0 | 4.3 | 1.0895 |  | 1.1494
2014/03/04 23:40:40 | 56 | 20140009.038 | 429 |  | 0.751 | 5 |  |  | 4168 | 0 | 59.8664 | 82.204 |  |  |  |  | 80.92 | 4.5 | 5.3 | 8.4 | 1.1057 |  | 1.165
2014/03/04 23:49:10 | 57 | 20140009.039 | 542 |  | 0.806 | 1 | 3912 | 0.007 |  |  | 10.7901 | 77.191 | 76.603 | 22.0 | 17.5 | 486.1 | [5 empty] | 0.36647 |  | 0.71961
2014/03/04 23:49:10 | 57 | 20140009.039 | 542 |  | 0.806 | 2 | 3913 | 0 |  |  | 10.7912 | 77.199 | 76.611 | 21.6 | 17.0 | 500.0 | [5 empty] | 0.36647 |  | 0.71961
2014/03/04 23:49:10 | 57 | 20140009.039 | 542 |  | 0.806 | 3 | 3857 | 15.349 |  |  | 11.9181 | 85.261 | 84.597 | 21.4 | 16.7 | 478.6 | [5 empty] | 0.37208 |  | 0.73065
2014/03/04 23:49:10 | 57 | 20140009.039 | 542 |  | 0.806 | 4 |  |  | 2938 | -16.284 | 48.016 | 70.761 |  |  |  |  | 69.664 | 1.7 | 2.0 | 4.3 | 1.0878 |  | 1.1477
2014/03/04 23:49:10 | 57 | 20140009.039 | 542 |  | 0.806 | 5 |  |  | 4168 | 0 | 55.7476 | 82.155 |  |  |  |  | 80.872 | 4.6 | 5.4 | 8.5 | 1.1057 |  | 1.165
2014/03/04 23:57:39 | 58 | 20140009.040 | 576 |  | 0.808 | 1 | 3909 | 0.013 |  |  | 10.7498 | 77.094 | 76.506 | 22.0 | 17.5 | 488.8 | [5 empty] | 0.36648 |  | 0.71961
2014/03/04 23:57:39 | 58 | 20140009.040 | 576 |  | 0.808 | 2 | 3911 | 0 |  |  | 10.7507 | 77.1 | 76.513 | 21.5 | 17.0 | 502.6 | [5 empty] | 0.36647 |  | 0.7196
2014/03/04 23:57:39 | 58 | 20140009.040 | 576 |  | 0.808 | 3 | 4293 | 15.26 |  |  | 13.1525 | 94.324 | 93.589 | 21.3 | 16.7 | 480.9 | [5 empty] | 0.37204 |  | 0.73059
2014/03/04 23:57:39 | 58 | 20140009.040 | 576 |  | 0.808 | 4 |  |  | 2936 | -16.078 | 47.5474 | 70.244 |  |  |  |  | 69.156 | 1.7 | 2.0 | 4.3 | 1.0881 |  | 1.1479
2014/03/04 23:57:39 | 58 | 20140009.040 | 576 |  | 0.808 | 5 |  |  | 4166 | 0 | 55.5307 | 82.038 |  |  |  |  | 80.757 | 4.6 | 5.4 | 8.6 | 1.1057 |  | 1.1649
2014/03/05 00:06:08 | 59 | 20140009.041 | 989 |  | 0.816 | 1 | 3909 | -0.018 |  |  | 10.6508 | 77.14 | 76.552 | 21.8 | 17.4 | 492.0 | [5 empty] | 0.36646 |  | 0.71946
2014/03/05 00:06:08 | 59 | 20140009.041 | 989 |  | 0.816 | 2 | 3907 | 0 |  |  | 10.6439 | 77.089 | 76.502 | 21.3 | 16.8 | 505.5 | [5 empty] | 0.36647 |  | 0.71948
2014/03/05 00:06:08 | 59 | 20140009.041 | 989 |  | 0.816 | 3 | 3929 | 13.464 |  |  | 11.9047 | 86.222 | 85.551 | 21.1 | 16.5 | 482.4 | [5 empty] | 0.37139 |  | 0.72916
2014/03/05 00:06:08 | 59 | 20140009.041 | 989 |  | 0.816 | 4 |  |  | 2597 | -13.469 | 41.4235 | 61.803 |  |  |  |  | 60.843 | 1.7 | 2.1 | 4.3 | 1.0909 |  | 1.1508
2014/03/05 00:06:08 | 59 | 20140009.041 | 989 |  | 0.816 | 5 |  |  | 4162 | 0 | 54.8867 | 81.89 |  |  |  |  | 80.611 | 4.5 | 5.3 | 8.3 | 1.1057 |  | 1.165
2014/03/05 00:14:38 | 60 | 20140009.042 | 980 |  | 0.798 | 1 | 3907 | -0.078 |  |  | 10.8821 | 77.076 | 76.489 | 21.6 | 17.2 | 491.8 | [5 empty] | 0.36644 |  | 0.71971
2014/03/05 00:14:38 | 60 | 20140009.042 | 980 |  | 0.798 | 2 | 3907 | 0 |  |  | 10.8746 | 77.023 | 76.437 | 21.2 | 16.7 | 505.7 | [5 empty] | 0.36647 |  | 0.71976
2014/03/05 00:14:38 | 60 | 20140009.042 | 980 |  | 0.798 | 3 | 3626 | 13.252 |  |  | 11.2216 | 79.481 | 78.863 | 20.9 | 16.4 | 483.2 | [5 empty] | 0.37131 |  | 0.7293
2014/03/05 00:14:38 | 60 | 20140009.042 | 980 |  | 0.798 | 4 |  |  | 2928 | -15.116 | 47.9276 | 69.93 |  |  |  |  | 68.845 | 1.7 | 2.0 | 4.3 | 1.0891 |  | 1.1488
2014/03/05 00:14:38 | 60 | 20140009.042 | 980 |  | 0.798 | 5 |  |  | 4160 | 0 | 56.1316 | 81.9 |  |  |  |  | 80.621 | 4.6 | 5.4 | 8.5 | 1.1057 |  | 1.1648
2014/03/05 00:23:07 | 61 | 20140009.043 | 899 |  | 0.849 | 1 | 3903 | -0.077 |  |  | 10.2141 | 76.969 | 76.383 | 21.5 | 17.2 | 492.1 | [5 empty] | 0.36644 |  | 0.71953
2014/03/05 00:23:07 | 61 | 20140009.043 | 899 |  | 0.849 | 2 | 3904 | 0 |  |  | 10.2116 | 76.95 | 76.364 | 21.0 | 16.6 | 506.1 | [5 empty] | 0.36647 |  | 0.71958
2014/03/05 00:23:07 | 61 | 20140009.043 | 899 |  | 0.849 | 3 | 3485 | 13.82 |  |  | 10.1681 | 76.622 | 76.026 | 20.8 | 16.3 | 483.5 | [5 empty] | 0.37152 |  | 0.72953
2014/03/05 00:23:07 | 61 | 20140009.043 | 899 |  | 0.849 | 4 |  |  | 2502 | -14.276 | 38.3622 | 59.55 |  |  |  |  | 58.626 | 1.7 | 2.0 | 4.3 | 1.09 |  | 1.1499
2014/03/05 00:23:07 | 61 | 20140009.043 | 899 |  | 0.849 | 5 |  |  | 4156 | 0 | 52.7408 | 81.87 |  |  |  |  | 80.592 | 4.4 | 5.2 | 8.2 | 1.1057 |  | 1.1649
2014/03/05 00:31:37 | 62 | 20140009.044 | 898 |  | 0.788 | 1 | 3903 | -0.021 |  |  | 11.0065 | 76.981 | 76.394 | 21.3 | 17.0 | 491.8 | [5 empty] | 0.36646 |  | 0.71952
2014/03/05 00:31:37 | 62 | 20140009.044 | 898 |  | 0.788 | 2 | 3903 | 0 |  |  | 10.9998 | 76.934 | 76.348 | 20.9 | 16.5 | 505.6 | [5 empty] | 0.36647 |  | 0.71953
2014/03/05 00:31:37 | 62 | 20140009.044 | 898 |  | 0.788 | 3 | 3418 | 15.117 |  |  | 10.7398 | 75.115 | 74.531 | 20.6 | 16.2 | 483.2 | [5 empty] | 0.37199 |  | 0.73041
2014/03/05 00:31:37 | 62 | 20140009.044 | 898 |  | 0.788 | 4 |  |  | 2947 | -15.15 | 49.0312 | 70.643 |  |  |  |  | 69.548 | 1.7 | 2.0 | 4.3 | 1.0891 |  | 1.1488
2014/03/05 00:31:37 | 62 | 20140009.044 | 898 |  | 0.788 | 5 |  |  | 4157 | 0 | 56.8492 | 81.907 |  |  |  |  | 80.628 | 4.6 | 5.4 | 8.5 | 1.1057 |  | 1.1648
2014/03/05 00:40:07 | 63 | 20140009.045 | 705 |  | 0.75 | 1 | 3902 | -0.067 |  |  | 11.5761 | 77.06 | 76.473 | 21.2 | 17.0 | 493.2 | [5 empty] | 0.36645 |  | 0.71946
2014/03/05 00:40:07 | 63 | 20140009.045 | 705 |  | 0.75 | 2 | 3903 | 0 |  |  | 11.556 | 76.926 | 76.34 | 20.7 | 16.5 | 506.7 | [5 empty] | 0.36647 |  | 0.71951
2014/03/05 00:40:07 | 63 | 20140009.045 | 705 |  | 0.75 | 3 | 3238 | 14.189 |  |  | 10.6663 | 71.003 | 70.452 | 20.5 | 16.1 | 483.7 | [5 empty] | 0.37165 |  | 0.72972
2014/03/05 00:40:07 | 63 | 20140009.045 | 705 |  | 0.75 | 4 |  |  | 2666 | -15.231 | 46.2732 | 63.455 |  |  |  |  | 62.471 | 1.7 | 2.1 | 4.3 | 1.089 |  | 1.1489
2014/03/05 00:40:07 | 63 | 20140009.045 | 705 |  | 0.75 | 5 |  |  | 4151 | 0 | 59.6832 | 81.844 |  |  |  |  | 80.566 | 4.5 | 5.2 | 8.3 | 1.1057 |  | 1.1649
2014/03/05 00:48:36 | 64 | 20140009.046 | 711 |  | 0.786 | 1 | 3900 | -0.053 |  |  | 11.0368 | 76.996 | 76.41 | 21.1 | 16.9 | 492.2 | [5 empty] | 0.36645 |  | 0.71957
2014/03/05 00:48:36 | 64 | 20140009.046 | 711 |  | 0.786 | 2 | 3901 | 0 |  |  | 11.0264 | 76.924 | 76.338 | 20.6 | 16.4 | 506.3 | [5 empty] | 0.36647 |  | 0.71961
2014/03/05 00:48:36 | 64 | 20140009.046 | 711 |  | 0.786 | 3 | 3620 | 14.695 |  |  | 11.3888 | 79.452 | 78.834 | 20.4 | 16.0 | 483.9 | [5 empty] | 0.37184 |  | 0.73018
2014/03/05 00:48:36 | 64 | 20140009.046 | 711 |  | 0.786 | 4 |  |  | 2724 | -14.786 | 45.2148 | 64.979 |  |  |  |  | 63.971 | 1.7 | 2.0 | 4.3 | 1.0895 |  | 1.1492
2014/03/05 00:48:36 | 64 | 20140009.046 | 711 |  | 0.786 | 5 |  |  | 4148 | 0 | 56.8671 | 81.725 |  |  |  |  | 80.449 | 4.5 | 5.3 | 8.4 | 1.1057 |  | 1.1648
2014/03/05 00:57:06 | 65 | 20140009.047 | 606 |  | 0.845 | 1 | 3896 | -0.018 |  |  | 10.251 | 76.883 | 76.297 | 20.9 | 16.7 | 492.9 | [5 empty] | 0.36647 |  | 0.71944
2014/03/05 00:57:06 | 65 | 20140009.047 | 606 |  | 0.845 | 2 | 3896 | 0 |  |  | 10.2488 | 76.867 | 76.281 | 20.5 | 16.2 | 507.1 | [5 empty] | 0.36647 |  | 0.71946
2014/03/05 00:57:06 | 65 | 20140009.047 | 606 |  | 0.845 | 3 | 3477 | 12.567 |  |  | 10.2102 | 76.577 | 75.982 | 20.2 | 15.9 | 484.1 | [5 empty] | 0.37106 |  | 0.7285
2014/03/05 00:57:06 | 65 | 20140009.047 | 606 |  | 0.845 | 4 |  |  | 2701 | -14.327 | 41.8711 | 64.691 |  |  |  |  | 63.687 | 1.7 | 2.0 | 4.3 | 1.09 |  | 1.1497
2014/03/05 00:57:06 | 65 | 20140009.047 | 606 |  | 0.845 | 5 |  |  | 4146 | 0 | 52.8271 | 81.618 |  |  |  |  | 80.343 | 4.5 | 5.3 | 8.4 | 1.1057 |  | 1.1648
2014/03/05 01:05:35 | 66 | 20140009.048 | 613 |  | 0.811 | 1 | 3894 | 0.016 |  |  | 10.68 | 76.877 | 76.292 | 20.9 | 16.8 | 493.8 | [5 empty] | 0.36648 |  | 0.71962
2014/03/05 01:05:35 | 66 | 20140009.048 | 613 |  | 0.811 | 2 | 3896 | 0 |  |  | 10.6801 | 76.878 | 76.292 | 20.5 | 16.3 | 507.9 | [5 empty] | 0.36647 |  | 0.71961
2014/03/05 01:05:35 | 66 | 20140009.048 | 613 |  | 0.811 | 3 | 3595 | 14.383 |  |  | 10.9079 | 78.518 | 77.907 | 20.2 | 15.9 | 485.2 | [5 empty] | 0.37172 |  | 0.72996
2014/03/05 01:05:35 | 66 | 20140009.048 | 613 |  | 0.811 | 4 |  |  | 2824 | -15.17 | 45.2851 | 67.15 |  |  |  |  | 66.109 | 1.7 | 2.0 | 4.3 | 1.0891 |  | 1.1487
2014/03/05 01:05:35 | 66 | 20140009.048 | 613 |  | 0.811 | 5 |  |  | 4143 | 0 | 55.0328 | 81.605 |  |  |  |  | 80.33 | 4.6 | 5.4 | 8.4 | 1.1057 |  | 1.1647
2014/03/05 01:14:05 | 67 | 36-UWSIF-UT Glut 120140009.18 | 36-UWSIF-UT Glut 1 |  | 0.7936 | 1 | 3893 | -0.062 |  |  | 10.9016 | 76.789 | 76.204 | 20.8 | 16.7 | 495.5 | [5 empty] | 0.36645 |  | 0.71963
2014/03/05 01:14:05 | 67 | 36-UWSIF-UT Glut 120140009.18 | 36-UWSIF-UT Glut 1 |  | 0.7936 | 2 | 3894 | 0 |  |  | 10.9061 | 76.82 | 76.235 | 20.3 | 16.1 | 509.3 | [5 empty] | 0.36647 |  | 0.71968
2014/03/05 01:14:05 | 67 | 36-UWSIF-UT Glut 120140009.18 | 36-UWSIF-UT Glut 1 |  | 0.7936 | 3 | 2948 | -4.239 |  |  | 9.1701 | 64.592 | 64.1 | 20.1 | 15.8 | 486.2 | [5 empty] | 0.36492 |  | 0.71663
2014/03/05 01:14:05 | 67 | 36-UWSIF-UT Glut 120140009.18 | 36-UWSIF-UT Glut 1 |  | 0.7936 | 4 |  |  | 2379 | -14.243 | 38.5975 | 56.006 |  |  |  |  | 55.137 | 1.7 | 2.0 | 4.3 | 1.0901 |  | 1.1497
2014/03/05 01:14:05 | 67 | 36-UWSIF-UT Glut 120140009.18 | 36-UWSIF-UT Glut 1 |  | 0.7936 | 5 |  |  | 4142 | 0 | 56.1907 | 81.534 |  |  |  |  | 80.261 | 4.3 | 5.1 | 8.1 | 1.1057 |  | 1.1648
2014/03/05 01:22:34 | 68 | 36-UWSIF-UT Glut 120140009.19 | 36-UWSIF-UT Glut 1 |  | 0.7865 | 1 | 3893 | 0.008 |  |  | 10.9987 | 76.78 | 76.195 | 20.7 | 16.6 | 493.9 | [5 empty] | 0.36647 |  | 0.71959
2014/03/05 01:22:34 | 68 | 36-UWSIF-UT Glut 120140009.19 | 36-UWSIF-UT Glut 1 |  | 0.7865 | 2 | 3891 | 0 |  |  | 11.0005 | 76.792 | 76.207 | 20.3 | 16.1 | 508.2 | [5 empty] | 0.36647 |  | 0.71958
2014/03/05 01:22:34 | 68 | 36-UWSIF-UT Glut 120140009.19 | 36-UWSIF-UT Glut 1 |  | 0.7865 | 3 | 2920 | -4.242 |  |  | 9.1827 | 64.102 | 63.614 | 20.0 | 15.7 | 486.3 | [5 empty] | 0.36492 |  | 0.71653
2014/03/05 01:22:34 | 68 | 36-UWSIF-UT Glut 120140009.19 | 36-UWSIF-UT Glut 1 |  | 0.7865 | 4 |  |  | 2366 | -14.203 | 38.6144 | 55.529 |  |  |  |  | 54.667 | 1.6 | 2.0 | 4.3 | 1.0901 |  | 1.1497
2014/03/05 01:22:34 | 68 | 36-UWSIF-UT Glut 120140009.19 | 36-UWSIF-UT Glut 1 |  | 0.7865 | 5 |  |  | 4139 | 0 | 56.7069 | 81.547 |  |  |  |  | 80.274 | 4.3 | 5.0 | 7.9 | 1.1057 |  | 1.1647
2014/03/05 01:31:03 | 69 | 39-UWSIF-UW Glut 220140009.15 | 39-UWSIF-UW Glut 2 |  | 0.7345 | 1 | 3890 | -0.04 |  |  | 11.7678 | 76.717 | 76.133 | 20.6 | 16.5 | 493.2 | [5 empty] | 0.36646 |  | 0.71935
2014/03/05 01:31:03 | 69 | 39-UWSIF-UW Glut 220140009.15 | 39-UWSIF-UW Glut 2 |  | 0.7345 | 2 | 3890 | 0 |  |  | 11.77 | 76.732 | 76.147 | 20.1 | 16.0 | 507.6 | [5 empty] | 0.36647 |  | 0.71938
2014/03/05 01:31:03 | 69 | 39-UWSIF-UW Glut 220140009.15 | 39-UWSIF-UW Glut 2 |  | 0.7345 | 3 | 2953 | 28.382 |  |  | 9.979 | 65.056 | 64.544 | 19.9 | 15.6 | 485.7 | [5 empty] | 0.37683 |  | 0.7398
2014/03/05 01:31:03 | 69 | 39-UWSIF-UW Glut 220140009.15 | 39-UWSIF-UW Glut 2 |  | 0.7345 | 4 |  |  | 2383 | 36.967 | 41.8761 | 56.238 |  |  |  |  | 55.335 | 1.6 | 2.0 | 4.3 | 1.1461 |  | 1.2059
2014/03/05 01:31:03 | 69 | 39-UWSIF-UW Glut 220140009.15 | 39-UWSIF-UW Glut 2 |  | 0.7345 | 5 |  |  | 4135 | 0 | 60.6304 | 81.424 |  |  |  |  | 80.152 | 4.3 | 5.2 | 8.0 | 1.1057 |  | 1.1653
2014/03/05 01:39:33 | 70 | 39-UWSIF-UW Glut 220140009.16 | 39-UWSIF-UW Glut 2 |  | 0.754 | 1 | 3889 | -0.033 |  |  | 11.4608 | 76.699 | 76.115 | 20.6 | 16.5 | 493.4 | [5 empty] | 0.36646 |  | 0.71945
2014/03/05 01:39:33 | 70 | 39-UWSIF-UW Glut 220140009.16 | 39-UWSIF-UW Glut 2 |  | 0.754 | 2 | 3889 | 0 |  |  | 11.4588 | 76.686 | 76.102 | 20.1 | 16.0 | 507.8 | [5 empty] | 0.36647 |  | 0.71947
2014/03/05 01:39:33 | 70 | 39-UWSIF-UW Glut 220140009.16 | 39-UWSIF-UW Glut 2 |  | 0.754 | 3 | 3016 | 28.231 |  |  | 9.8916 | 66.198 | 65.678 | 19.9 | 15.7 | 485.5 | [5 empty] | 0.37678 |  | 0.73979
2014/03/05 01:39:33 | 70 | 39-UWSIF-UW Glut 220140009.16 | 39-UWSIF-UW Glut 2 |  | 0.754 | 4 |  |  | 2441 | 37.12 | 41.7211 | 57.517 |  |  |  |  | 56.594 | 1.6 | 2.0 | 4.2 | 1.1462 |  | 1.2061
2014/03/05 01:39:33 | 70 | 39-UWSIF-UW Glut 220140009.16 | 39-UWSIF-UW Glut 2 |  | 0.754 | 5 |  |  | 4134 | 0 | 59.0794 | 81.448 |  |  |  |  | 80.176 | 4.3 | 5.2 | 8.0 | 1.1057 |  | 1.1653
2014/03/05 01:48:03 | 71 | 01-UWSIF-Liver 20140009.15 | 01-UWSIF-Liver |  | 0.7261 | 1 | 3887 | -0.048 |  |  | 11.9081 | 76.744 | 76.159 | 20.5 | 16.5 | 494.2 | [5 empty] | 0.36645 |  | 0.71948
2014/03/05 01:48:03 | 71 | 01-UWSIF-Liver 20140009.15 | 01-UWSIF-Liver |  | 0.7261 | 2 | 3887 | 0 |  |  | 11.8941 | 76.654 | 76.07 | 20.0 | 16.0 | 508.5 | [5 empty] | 0.36647 |  | 0.71951
2014/03/05 01:48:03 | 71 | 01-UWSIF-Liver 20140009.15 | 01-UWSIF-Liver |  | 0.7261 | 3 | 3690 | 7.188 |  |  | 12.5663 | 80.986 | 80.36 | 19.8 | 15.6 | 486.3 | [5 empty] | 0.3691 |  | 0.72469
2014/03/05 01:48:03 | 71 | 01-UWSIF-Liver 20140009.15 | 01-UWSIF-Liver |  | 0.7261 | 4 |  |  | 2655 | -3.818 | 47.4881 | 63.045 |  |  |  |  | 62.06 | 1.6 | 2.0 | 4.3 | 1.1015 |  | 1.1612
2014/03/05 01:48:03 | 71 | 01-UWSIF-Liver 20140009.15 | 01-UWSIF-Liver |  | 0.7261 | 5 |  |  | 4132 | 0 | 61.2888 | 81.367 |  |  |  |  | 80.097 | 4.5 | 5.3 | 8.3 | 1.1057 |  | 1.1649
2014/03/05 01:56:32 | 72 | 01-UWSIF-Liver 20140009.16 | 01-UWSIF-Liver |  | 0.7155 | 1 | 3888 | 0.021 |  |  | 12.0795 | 76.712 | 76.129 | 20.6 | 16.5 | 496.5 | [5 empty] | 0.36648 |  | 0.71805
2014/03/05 01:56:32 | 72 | 01-UWSIF-Liver 20140009.16 | 01-UWSIF-Liver |  | 0.7155 | 2 | 3886 | 0 |  |  | 12.0826 | 76.732 | 76.148 | 20.1 | 16.0 | 510.4 | [5 empty] | 0.36647 |  | 0.71803
2014/03/05 01:56:32 | 72 | 01-UWSIF-Liver 20140009.16 | 01-UWSIF-Liver |  | 0.7155 | 3 | 3638 | 7.331 |  |  | 12.5616 | 79.774 | 79.158 | 19.9 | 15.6 | 488.3 | [5 empty] | 0.36915 |  | 0.7233
2014/03/05 01:56:32 | 72 | 01-UWSIF-Liver 20140009.16 | 01-UWSIF-Liver |  | 0.7155 | 4 |  |  | 2627 | -3.91 | 47.5494 | 62.205 |  |  |  |  | 61.233 | 1.7 | 2.0 | 4.3 | 1.1014 |  | 1.161
2014/03/05 01:56:32 | 72 | 01-UWSIF-Liver 20140009.16 | 01-UWSIF-Liver |  | 0.7155 | 5 |  |  | 4131 | 0 | 62.1871 | 81.355 |  |  |  |  | 80.084 | 4.4 | 5.2 | 8.2 | 1.1057 |  | 1.1648
2014/03/05 02:05:02 | 73 | 20140009.049 | 1039 |  | 0.756 | 1 | 3883 | -0.021 |  |  | 11.4275 | 76.679 | 76.095 | 20.4 | 16.4 | 497.9 | [5 empty] | 0.36646 |  | 0.71953
2014/03/05 02:05:02 | 73 | 20140009.049 | 1039 |  | 0.756 | 2 | 3884 | 0 |  |  | 11.4259 | 76.669 | 76.085 | 20.0 | 15.9 | 512.3 | [5 empty] | 0.36647 |  | 0.71955
2014/03/05 02:05:02 | 73 | 20140009.049 | 1039 |  | 0.756 | 3 | 3351 | 14.733 |  |  | 10.9966 | 73.788 | 73.214 | 19.7 | 15.5 | 490.0 | [5 empty] | 0.37185 |  | 0.73015
2014/03/05 02:05:02 | 73 | 20140009.049 | 1039 |  | 0.756 | 4 |  |  | 2614 | -14.129 | 45.0646 | 62.292 |  |  |  |  | 61.325 | 1.7 | 2.0 | 4.3 | 1.0902 |  | 1.1498
2014/03/05 02:05:02 | 73 | 20140009.049 | 1039 |  | 0.756 | 5 |  |  | 4130 | 0 | 58.7789 | 81.249 |  |  |  |  | 79.98 | 4.4 | 5.2 | 8.2 | 1.1057 |  | 1.1647
2014/03/05 02:13:32 | 74 | 20140009.050 | 1035 |  | 0.78 | 1 | 3887 | -0.046 |  |  | 11.0943 | 76.807 | 76.223 | 20.5 | 16.4 | 498.0 | [5 empty] | 0.36645 |  | 0.71793
2014/03/05 02:13:32 | 74 | 20140009.050 | 1035 |  | 0.78 | 2 | 3888 | 0 |  |  | 11.0769 | 76.687 | 76.104 | 20.0 | 15.9 | 512.5 | [5 empty] | 0.36647 |  | 0.71796
2014/03/05 02:13:32 | 74 | 20140009.050 | 1035 |  | 0.78 | 3 | 3410 | 13.09 |  |  | 10.8061 | 74.812 | 74.232 | 19.8 | 15.6 | 489.3 | [5 empty] | 0.37125 |  | 0.72736
2014/03/05 02:13:32 | 74 | 20140009.050 | 1035 |  | 0.78 | 4 |  |  | 2804 | -15.266 | 47.0547 | 67.107 |  |  |  |  | 66.067 | 1.7 | 2.0 | 4.2 | 1.089 |  | 1.1484
2014/03/05 02:13:32 | 74 | 20140009.050 | 1035 |  | 0.78 | 5 |  |  | 4129 | 0 | 57.0107 | 81.306 |  |  |  |  | 80.037 | 4.6 | 5.4 | 8.4 | 1.1057 |  | 1.1645
2014/03/05 02:22:02 | 75 | 20140009.051 | 1146 |  | 0.801 | 1 | 3883 | 0.017 |  |  | 10.7822 | 76.656 | 76.072 | 20.4 | 16.4 | 499.6 | [5 empty] | 0.36648 |  | 0.71941
2014/03/05 02:22:02 | 75 | 20140009.051 | 1146 |  | 0.801 | 2 | 3884 | 0 |  |  | 10.7773 | 76.621 | 76.038 | 20.0 | 15.9 | 513.3 | [5 empty] | 0.36647 |  | 0.7194
2014/03/05 02:22:02 | 75 | 20140009.051 | 1146 |  | 0.801 | 3 | 3825 | 15.469 |  |  | 11.7855 | 83.789 | 83.136 | 19.7 | 15.5 | 490.7 | [5 empty] | 0.37212 |  | 0.73053
2014/03/05 02:22:02 | 75 | 20140009.051 | 1146 |  | 0.801 | 4 |  |  | 2692 | -14.07 | 43.8488 | 64.219 |  |  |  |  | 63.222 | 1.7 | 2.0 | 4.3 | 1.0903 |  | 1.1498
2014/03/05 02:22:02 | 75 | 20140009.051 | 1146 |  | 0.801 | 5 |  |  | 4126 | 0 | 55.5122 | 81.301 |  |  |  |  | 80.031 | 4.5 | 5.3 | 8.3 | 1.1057 |  | 1.1646
2014/03/05 02:30:31 | 76 | 20140009.052 | 1132 |  | 0.797 | 1 | 3881 | -0.007 |  |  | 10.8368 | 76.659 | 76.075 | 20.5 | 16.4 | 501.1 | [5 empty] | 0.36647 |  | 0.7194
2014/03/05 02:30:31 | 76 | 20140009.052 | 1132 |  | 0.797 | 2 | 3883 | 0 |  |  | 10.8317 | 76.623 | 76.04 | 20.0 | 15.9 | 514.8 | [5 empty] | 0.36647 |  | 0.7194
2014/03/05 02:30:31 | 76 | 20140009.052 | 1132 |  | 0.797 | 3 | 3137 | 14.949 |  |  | 9.7978 | 69.31 | 68.771 | 19.7 | 15.5 | 492.1 | [5 empty] | 0.37193 |  | 0.73016
2014/03/05 02:30:31 | 76 | 20140009.052 | 1132 |  | 0.797 | 4 |  |  | 2162 | -14.068 | 35.186 | 51.274 |  |  |  |  | 50.479 | 1.7 | 2.0 | 4.3 | 1.0903 |  | 1.1498
2014/03/05 02:30:31 | 76 | 20140009.052 | 1132 |  | 0.797 | 5 |  |  | 4124 | 0 | 55.7559 | 81.25 |  |  |  |  | 79.981 | 4.2 | 4.9 | 7.9 | 1.1057 |  | 1.1646
2014/03/05 02:39:01 | 77 | 20140009.053 | 1242 |  | 0.777 | 1 | 3883 | -0.006 |  |  | 11.1072 | 76.6 | 76.017 | 20.3 | 16.3 | 499.8 | [5 empty] | 0.36647 |  | 0.71939
2014/03/05 02:39:01 | 77 | 20140009.053 | 1242 |  | 0.777 | 2 | 3882 | 0 |  |  | 11.1018 | 76.563 | 75.98 | 19.9 | 15.8 | 513.5 | [5 empty] | 0.36647 |  | 0.71939
2014/03/05 02:39:01 | 77 | 20140009.053 | 1242 |  | 0.777 | 3 | 3495 | 14.813 |  |  | 11.1442 | 76.856 | 76.258 | 19.6 | 15.4 | 491.1 | [5 empty] | 0.37188 |  | 0.73005
2014/03/05 02:39:01 | 77 | 20140009.053 | 1242 |  | 0.777 | 4 |  |  | 2676 | -15.053 | 45.0781 | 64.041 |  |  |  |  | 63.048 | 1.6 | 2.0 | 4.2 | 1.0892 |  | 1.1486
2014/03/05 02:39:01 | 77 | 20140009.053 | 1242 |  | 0.777 | 5 |  |  | 4122 | 0 | 57.1671 | 81.216 |  |  |  |  | 79.948 | 4.5 | 5.2 | 8.2 | 1.1057 |  | 1.1645
2014/03/05 02:47:31 | 78 | 20140009.054 | 1299 |  | 0.772 | 1 | 3880 | -0.037 |  |  | 11.1784 | 76.595 | 76.012 | 20.4 | 16.4 | 500.8 | [5 empty] | 0.36646 |  | 0.71951
2014/03/05 02:47:31 | 78 | 20140009.054 | 1299 |  | 0.772 | 2 | 3881 | 0 |  |  | 11.1728 | 76.557 | 75.974 | 19.9 | 15.8 | 515.0 | [5 empty] | 0.36647 |  | 0.71953
2014/03/05 02:47:31 | 78 | 20140009.054 | 1299 |  | 0.772 | 3 | 3676 | 15.259 |  |  | 11.8176 | 80.975 | 80.344 | 19.7 | 15.5 | 492.5 | [5 empty] | 0.37204 |  | 0.73051
2014/03/05 02:47:31 | 78 | 20140009.054 | 1299 |  | 0.772 | 4 |  |  | 2506 | -14.308 | 42.3764 | 59.815 |  |  |  |  | 58.888 | 1.7 | 2.0 | 4.3 | 1.09 |  | 1.1494
2014/03/05 02:47:31 | 78 | 20140009.054 | 1299 |  | 0.772 | 5 |  |  | 4120 | 0 | 57.4295 | 81.063 |  |  |  |  | 79.798 | 4.4 | 5.1 | 8.1 | 1.1057 |  | 1.1645
2014/03/05 02:56:01 | 79 | 20140009.055 | 1367 |  | 0.785 | 1 | 3876 | -0.032 |  |  | 10.9903 | 76.574 | 75.991 | 20.3 | 16.3 | 501.3 | [5 empty] | 0.36646 |  | 0.71929
2014/03/05 02:56:01 | 79 | 20140009.055 | 1367 |  | 0.785 | 2 | 3878 | 0 |  |  | 10.9837 | 76.529 | 75.946 | 19.9 | 15.8 | 515.5 | [5 empty] | 0.36647 |  | 0.71931
2014/03/05 02:56:01 | 79 | 20140009.055 | 1367 |  | 0.785 | 3 | 3912 | 14.582 |  |  | 12.3594 | 86.114 | 85.443 | 19.6 | 15.4 | 493.4 | [5 empty] | 0.3718 |  | 0.7298
2014/03/05 02:56:01 | 79 | 20140009.055 | 1367 |  | 0.785 | 4 |  |  | 2610 | -15.042 | 43.3991 | 62.291 |  |  |  |  | 61.325 | 1.7 | 2.0 | 4.3 | 1.0892 |  | 1.1486
2014/03/05 02:56:01 | 79 | 20140009.055 | 1367 |  | 0.785 | 5 |  |  | 4115 | 0 | 56.4606 | 81.038 |  |  |  |  | 79.773 | 4.4 | 5.2 | 8.2 | 1.1057 |  | 1.1645
2014/03/05 03:04:31 | 80 | 20140009.056 | 1358 |  | 0.756 | 1 | 3875 | 0.039 |  |  | 11.3869 | 76.407 | 75.825 | 20.3 | 16.3 | 503.6 | [5 empty] | 0.36649 |  | 0.71944
2014/03/05 03:04:31 | 80 | 20140009.056 | 1358 |  | 0.756 | 2 | 3876 | 0 |  |  | 11.3937 | 76.452 | 75.87 | 19.9 | 15.8 | 517.2 | [5 empty] | 0.36647 |  | 0.71941
2014/03/05 03:04:31 | 80 | 20140009.056 | 1358 |  | 0.756 | 3 | 3755 | 14.282 |  |  | 12.3018 | 82.546 | 81.904 | 19.6 | 15.4 | 494.3 | [5 empty] | 0.37169 |  | 0.72969
2014/03/05 03:04:31 | 80 | 20140009.056 | 1358 |  | 0.756 | 4 |  |  | 2487 | -14.749 | 42.8846 | 59.278 |  |  |  |  | 58.359 | 1.7 | 2.0 | 4.3 | 1.0895 |  | 1.1489
2014/03/05 03:04:31 | 80 | 20140009.056 | 1358 |  | 0.756 | 5 |  |  | 4113 | 0 | 58.5808 | 80.975 |  |  |  |  | 79.71 | 4.4 | 5.1 | 8.1 | 1.1057 |  | 1.1645
2014/03/05 03:13:00 | 81 | 20140009.057 | 1543 |  | 0.796 | 1 | 3874 | -0.079 |  |  | 10.8156 | 76.413 | 75.831 | 20.3 | 16.3 | 504.8 | [5 empty] | 0.36644 |  | 0.71932
2014/03/05 03:13:00 | 81 | 20140009.057 | 1543 |  | 0.796 | 2 | 3873 | 0 |  |  | 10.8068 | 76.351 | 75.77 | 19.8 | 15.7 | 518.2 | [5 empty] | 0.36647 |  | 0.71938
2014/03/05 03:13:00 | 81 | 20140009.057 | 1543 |  | 0.796 | 3 | 3153 | 14.768 |  |  | 9.7479 | 68.87 | 68.335 | 19.5 | 15.4 | 495.1 | [5 empty] | 0.37186 |  | 0.73
2014/03/05 03:13:00 | 81 | 20140009.057 | 1543 |  | 0.796 | 4 |  |  | 2878 | -17.446 | 47.0136 | 68.424 |  |  |  |  | 67.365 | 1.7 | 2.0 | 4.2 | 1.0866 |  | 1.1459
2014/03/05 03:13:00 | 81 | 20140009.057 | 1543 |  | 0.796 | 5 |  |  | 4110 | 0 | 55.6123 | 80.939 |  |  |  |  | 79.675 | 4.5 | 5.3 | 8.3 | 1.1057 |  | 1.1644
2014/03/05 03:21:29 | 82 | 20140009.058 | 1550 |  | 0.787 | 1 | 3873 | 0.032 |  |  | 10.9404 | 76.421 | 75.84 | 20.3 | 16.3 | 503.4 | [5 empty] | 0.36648 |  | 0.71776
2014/03/05 03:21:29 | 82 | 20140009.058 | 1550 |  | 0.787 | 2 | 3874 | 0 |  |  | 10.9353 | 76.386 | 75.805 | 19.9 | 15.8 | 517.3 | [5 empty] | 0.36647 |  | 0.71774
2014/03/05 03:21:29 | 82 | 20140009.058 | 1550 |  | 0.787 | 3 | 3049 | 12.866 |  |  | 9.5522 | 66.724 | 66.208 | 19.6 | 15.4 | 494.2 | [5 empty] | 0.37117 |  | 0.72698
2014/03/05 03:21:29 | 82 | 20140009.058 | 1550 |  | 0.787 | 4 |  |  | 2720 | -15.47 | 44.8707 | 64.567 |  |  |  |  | 63.566 | 1.7 | 2.0 | 4.3 | 1.0887 |  | 1.148
2014/03/05 03:21:29 | 82 | 20140009.058 | 1550 |  | 0.787 | 5 |  |  | 4108 | 0 | 56.2357 | 80.921 |  |  |  |  | 79.657 | 4.5 | 5.3 | 8.3 | 1.1057 |  | 1.1643
2014/03/05 03:29:59 | 83 | 20140009.059 | 1485 |  | 0.843 | 1 | 3869 | -0.018 |  |  | 10.2066 | 76.368 | 75.788 | 20.3 | 16.2 | 503.8 | [5 empty] | 0.36647 |  | 0.71797
2014/03/05 03:29:59 | 83 | 20140009.059 | 1485 |  | 0.843 | 2 | 3871 | 0 |  |  | 10.2119 | 76.408 | 75.827 | 19.8 | 15.7 | 517.4 | [5 empty] | 0.36647 |  | 0.71799
2014/03/05 03:29:59 | 83 | 20140009.059 | 1485 |  | 0.843 | 3 | 3790 | 12.716 |  |  | 11.1467 | 83.402 | 82.755 | 19.5 | 15.4 | 495.7 | [5 empty] | 0.37111 |  | 0.72712
2014/03/05 03:29:59 | 83 | 20140009.059 | 1485 |  | 0.843 | 4 |  |  | 2548 | -13.623 | 39.4591 | 60.82 |  |  |  |  | 59.876 | 1.7 | 2.0 | 4.3 | 1.0908 |  | 1.1499
2014/03/05 03:29:59 | 83 | 20140009.059 | 1485 |  | 0.843 | 5 |  |  | 4106 | 0 | 52.4552 | 80.852 |  |  |  |  | 79.59 | 4.4 | 5.1 | 8.1 | 1.1057 |  | 1.1643
2014/03/05 03:38:30 | 84 | 20140009.060 | 1494 |  | 0.804 | 1 | 3868 | -0.076 |  |  | 10.6972 | 76.336 | 75.756 | 20.2 | 16.2 | 505.1 | [5 empty] | 0.36644 |  | 0.71788
2014/03/05 03:38:30 | 84 | 20140009.060 | 1494 |  | 0.804 | 2 | 3870 | 0 |  |  | 10.6943 | 76.316 | 75.735 | 19.8 | 15.7 | 519.0 | [5 empty] | 0.36647 |  | 0.71793
2014/03/05 03:38:30 | 84 | 20140009.060 | 1494 |  | 0.804 | 3 | 3631 | 14.866 |  |  | 11.18 | 79.782 | 79.162 | 19.5 | 15.3 | 496.5 | [5 empty] | 0.3719 |  | 0.7286
2014/03/05 03:38:30 | 84 | 20140009.060 | 1494 |  | 0.804 | 4 |  |  | 2797 | -15.152 | 45.6014 | 67.036 |  |  |  |  | 65.997 | 1.7 | 2.0 | 4.3 | 1.0891 |  | 1.1483
2014/03/05 03:38:30 | 84 | 20140009.060 | 1494 |  | 0.804 | 5 |  |  | 4103 | 0 | 54.9833 | 80.828 |  |  |  |  | 79.566 | 4.5 | 5.3 | 8.4 | 1.1057 |  | 1.1643
2014/03/05 03:47:00 | 85 | 20140009.061 | 141 |  | 0.787 | 1 | 3868 | -0.06 |  |  | 10.9317 | 76.36 | 75.78 | 20.3 | 16.2 | 506.4 | [5 empty] | 0.36645 |  | 0.71782
2014/03/05 03:47:00 | 85 | 20140009.061 | 141 |  | 0.787 | 2 | 3871 | 0 |  |  | 10.9259 | 76.32 | 75.74 | 19.8 | 15.8 | 520.4 | [5 empty] | 0.36647 |  | 0.71786
2014/03/05 03:47:00 | 85 | 20140009.061 | 141 |  | 0.787 | 3 | 3321 | 13.169 |  |  | 10.4298 | 72.855 | 72.29 | 19.5 | 15.4 | 497.3 | [5 empty] | 0.37128 |  | 0.72731
2014/03/05 03:47:00 | 85 | 20140009.061 | 141 |  | 0.787 | 4 |  |  | 2399 | -15.432 | 39.5031 | 56.843 |  |  |  |  | 55.962 | 1.7 | 2.0 | 4.3 | 1.0888 |  | 1.148
2014/03/05 03:47:00 | 85 | 20140009.061 | 141 |  | 0.787 | 5 |  |  | 4100 | 0 | 56.1515 | 80.8 |  |  |  |  | 79.538 | 4.3 | 5.1 | 8.1 | 1.1057 |  | 1.1643
2014/03/05 03:55:30 | 86 | 20140009.062 | 160 |  | 0.791 | 1 | 3868 | -0.097 |  |  | 10.8745 | 76.347 | 75.768 | 20.2 | 16.1 | 505.4 | [5 empty] | 0.36644 |  | 0.71567
2014/03/05 03:55:30 | 86 | 20140009.062 | 160 |  | 0.791 | 2 | 3870 | 0 |  |  | 10.8649 | 76.28 | 75.701 | 19.7 | 15.6 | 519.5 | [5 empty] | 0.36647 |  | 0.71574
2014/03/05 03:55:30 | 86 | 20140009.062 | 160 |  | 0.791 | 3 | 3525 | 13.015 |  |  | 11.0307 | 77.443 | 76.845 | 19.5 | 15.3 | 497.3 | [5 empty] | 0.37122 |  | 0.72505
2014/03/05 03:55:30 | 86 | 20140009.062 | 160 |  | 0.791 | 4 |  |  | 2564 | -14.219 | 42.1857 | 61.012 |  |  |  |  | 60.066 | 1.7 | 2.0 | 4.3 | 1.0901 |  | 1.1493
2014/03/05 03:55:30 | 86 | 20140009.062 | 160 |  | 0.791 | 5 |  |  | 4099 | 0 | 55.8326 | 80.749 |  |  |  |  | 79.488 | 4.4 | 5.2 | 8.2 | 1.1057 |  | 1.1643
2014/03/05 04:04:00 | 87 | 20140009.063 | 238 |  | 0.834 | 1 | 3863 | 0.01 |  |  | 10.2997 | 76.242 | 75.661 | 20.1 | 16.2 | 506.9 | [5 empty] | 0.36648 |  | 0.71941
2014/03/05 04:04:00 | 87 | 20140009.063 | 238 |  | 0.834 | 2 | 3865 | 0 |  |  | 10.3006 | 76.249 | 75.668 | 19.7 | 15.7 | 521.2 | [5 empty] | 0.36647 |  | 0.7194
2014/03/05 04:04:00 | 87 | 20140009.063 | 238 |  | 0.834 | 3 | 3853 | 13.418 |  |  | 11.3887 | 84.303 | 83.647 | 19.4 | 15.3 | 498.4 | [5 empty] | 0.37137 |  | 0.72905
2014/03/05 04:04:00 | 87 | 20140009.063 | 238 |  | 0.834 | 4 |  |  | 2832 | -14.529 | 44.3207 | 67.584 |  |  |  |  | 66.536 | 1.7 | 2.0 | 4.3 | 1.0898 |  | 1.1489
2014/03/05 04:04:00 | 87 | 20140009.063 | 238 |  | 0.834 | 5 |  |  | 4097 | 0 | 52.8976 | 80.663 |  |  |  |  | 79.404 | 4.5 | 5.3 | 8.3 | 1.1057 |  | 1.1642
2014/03/05 04:12:30 | 88 | 20140009.064 | 237 |  | 0.8 | 1 | 3862 | -0.047 |  |  | 10.7314 | 76.199 | 75.619 | 20.2 | 16.3 | 509.2 | [5 empty] | 0.36645 |  | 0.71924
2014/03/05 04:12:30 | 88 | 20140009.064 | 237 |  | 0.8 | 2 | 3862 | 0 |  |  | 10.7318 | 76.203 | 75.622 | 19.7 | 15.7 | 522.9 | [5 empty] | 0.36647 |  | 0.71928
2014/03/05 04:12:30 | 88 | 20140009.064 | 237 |  | 0.8 | 3 | 3722 | 13.813 |  |  | 11.4965 | 81.633 | 80.997 | 19.5 | 15.4 | 500.4 | [5 empty] | 0.37151 |  | 0.72921
2014/03/05 04:12:30 | 88 | 20140009.064 | 237 |  | 0.8 | 4 |  |  | 2730 | -14.601 | 44.6187 | 65.265 |  |  |  |  | 64.253 | 1.7 | 2.1 | 4.3 | 1.0897 |  | 1.1489
2014/03/05 04:12:30 | 88 | 20140009.064 | 237 |  | 0.8 | 5 |  |  | 4093 | 0 | 55.1102 | 80.611 |  |  |  |  | 79.353 | 4.5 | 5.3 | 8.3 | 1.1057 |  | 1.1643
2014/03/05 04:21:00 | 89 | 20140009.065 | 361 |  | 0.837 | 1 | 3865 | -0.015 |  |  | 10.2631 | 76.245 | 75.666 | 20.2 | 16.1 | 508.8 | [5 empty] | 0.36647 |  | 0.71564
2014/03/05 04:21:00 | 89 | 20140009.065 | 361 |  | 0.837 | 2 | 3866 | 0 |  |  | 10.2683 | 76.283 | 75.705 | 19.7 | 15.6 | 522.8 | [5 empty] | 0.36647 |  | 0.71565
2014/03/05 04:21:00 | 89 | 20140009.065 | 361 |  | 0.837 | 3 | 3205 | 12.574 |  |  | 9.5019 | 70.59 | 70.045 | 19.4 | 15.3 | 499.7 | [5 empty] | 0.37106 |  | 0.72465
2014/03/05 04:21:00 | 89 | 20140009.065 | 361 |  | 0.837 | 4 |  |  | 2147 | -13.609 | 33.2544 | 50.892 |  |  |  |  | 50.102 | 1.7 | 2.0 | 4.3 | 1.0908 |  | 1.1498
2014/03/05 04:21:00 | 89 | 20140009.065 | 361 |  | 0.837 | 5 |  |  | 4095 | 0 | 52.6889 | 80.634 |  |  |  |  | 79.375 | 4.2 | 5.0 | 7.9 | 1.1057 |  | 1.1642
2014/03/05 04:29:30 | 90 | 20140009.066 | 376 |  | 0.847 | 1 | 3861 | -0.017 |  |  | 10.1309 | 76.162 | 75.581 | 20.0 | 16.1 | 508.2 | [5 empty] | 0.36647 |  | 0.71933
2014/03/05 04:29:30 | 90 | 20140009.066 | 376 |  | 0.847 | 2 | 3862 | 0 |  |  | 10.1336 | 76.182 | 75.602 | 19.5 | 15.5 | 522.3 | [5 empty] | 0.36647 |  | 0.71934
2014/03/05 04:29:30 | 90 | 20140009.066 | 376 |  | 0.847 | 3 | 3737 | 13.409 |  |  | 10.918 | 82.079 | 81.441 | 19.3 | 15.2 | 500.5 | [5 empty] | 0.37137 |  | 0.72898
2014/03/05 04:29:30 | 90 | 20140009.066 | 376 |  | 0.847 | 4 |  |  | 2750 | -14.119 | 42.4558 | 65.75 |  |  |  |  | 64.73 | 1.6 | 2.0 | 4.2 | 1.0902 |  | 1.1493
2014/03/05 04:29:30 | 90 | 20140009.066 | 376 |  | 0.847 | 5 |  |  | 4094 | 0 | 52.0664 | 80.633 |  |  |  |  | 79.374 | 4.5 | 5.3 | 8.3 | 1.1057 |  | 1.1642
2014/03/05 04:38:00 | 91 | 20140009.067 | 499 |  | 0.754 | 1 | 3863 | -0.037 |  |  | 11.3812 | 76.167 | 75.589 | 20.1 | 16.1 | 509.5 | [5 empty] | 0.36646 |  | 0.71567
2014/03/05 04:38:00 | 91 | 20140009.067 | 499 |  | 0.754 | 2 | 3863 | 0 |  |  | 11.3919 | 76.238 | 75.66 | 19.6 | 15.6 | 523.2 | [5 empty] | 0.36647 |  | 0.7157
2014/03/05 04:38:00 | 91 | 20140009.067 | 499 |  | 0.754 | 3 | 2985 | 13.09 |  |  | 9.7639 | 65.343 | 64.839 | 19.3 | 15.2 | 500.2 | [5 empty] | 0.37125 |  | 0.72507
2014/03/05 04:38:00 | 91 | 20140009.067 | 499 |  | 0.754 | 4 |  |  | 2717 | -15.505 | 46.718 | 64.406 |  |  |  |  | 63.408 | 1.7 | 2.0 | 4.3 | 1.0887 |  | 1.1478
2014/03/05 04:38:00 | 91 | 20140009.067 | 499 |  | 0.754 | 5 |  |  | 4092 | 0 | 58.3906 | 80.498 |  |  |  |  | 79.242 | 4.5 | 5.3 | 8.3 | 1.1057 |  | 1.1641
2014/03/05 04:46:30 | 92 | 20140009.068 | 483 |  | 0.8 | 1 | 3861 | -0.028 |  |  | 10.7282 | 76.177 | 75.597 | 20.0 | 16.1 | 508.8 | [5 empty] | 0.36646 |  | 0.71768
2014/03/05 04:46:30 | 92 | 20140009.068 | 483 |  | 0.8 | 2 | 3862 | 0 |  |  | 10.7336 | 76.215 | 75.636 | 19.6 | 15.6 | 523.0 | [5 empty] | 0.36647 |  | 0.7177
2014/03/05 04:46:30 | 92 | 20140009.068 | 483 |  | 0.8 | 3 | 3294 | 14.793 |  |  | 10.1499 | 72.07 | 71.511 | 19.3 | 15.2 | 500.4 | [5 empty] | 0.37187 |  | 0.72832
2014/03/05 04:46:30 | 92 | 20140009.068 | 483 |  | 0.8 | 4 |  |  | 2966 | -16.387 | 48.3893 | 70.78 |  |  |  |  | 69.684 | 1.7 | 2.0 | 4.3 | 1.0877 |  | 1.1468
2014/03/05 04:46:30 | 92 | 20140009.068 | 483 |  | 0.8 | 5 |  |  | 4095 | 0 | 55.1247 | 80.632 |  |  |  |  | 79.374 | 4.6 | 5.4 | 8.5 | 1.1057 |  | 1.1642
2014/03/05 04:55:01 | 93 | 20140009.069 | 539 |  | 0.825 | 1 | 3862 | -0.052 |  |  | 10.397 | 76.132 | 75.555 | 20.1 | 16.1 | 509.3 | [5 empty] | 0.36645 |  | 0.71586
2014/03/05 04:55:01 | 93 | 20140009.069 | 539 |  | 0.825 | 2 | 3863 | 0 |  |  | 10.4036 | 76.181 | 75.603 | 19.6 | 15.6 | 523.0 | [5 empty] | 0.36647 |  | 0.7159
2014/03/05 04:55:01 | 93 | 20140009.069 | 539 |  | 0.825 | 3 | 4039 | 14.315 |  |  | 12.1627 | 89.062 | 88.371 | 19.3 | 15.2 | 501.7 | [5 empty] | 0.3717 |  | 0.72614
2014/03/05 04:55:01 | 93 | 20140009.069 | 539 |  | 0.825 | 4 |  |  | 2757 | -14.402 | 43.7933 | 66.059 |  |  |  |  | 65.035 | 1.7 | 2.1 | 4.3 | 1.0899 |  | 1.1488
2014/03/05 04:55:01 | 93 | 20140009.069 | 539 |  | 0.825 | 5 |  |  | 4091 | 0 | 53.4018 | 80.553 |  |  |  |  | 79.296 | 4.5 | 5.3 | 8.3 | 1.1057 |  | 1.164
2014/03/05 05:03:31 | 94 | 20140009.070 | 507 |  | 0.791 | 1 | 3857 | -0.052 |  |  | 10.8446 | 76.137 | 75.557 | 20.0 | 16.1 | 511.4 | [5 empty] | 0.36645 |  | 0.71924
2014/03/05 05:03:31 | 94 | 20140009.070 | 507 |  | 0.791 | 2 | 3857 | 0 |  |  | 10.8308 | 76.04 | 75.461 | 19.6 | 15.6 | 525.1 | [5 empty] | 0.36647 |  | 0.71927
2014/03/05 05:03:31 | 94 | 20140009.070 | 507 |  | 0.791 | 3 | 3984 | 14.801 |  |  | 12.4407 | 87.343 | 86.662 | 19.3 | 15.2 | 503.1 | [5 empty] | 0.37188 |  | 0.72992
2014/03/05 05:03:31 | 94 | 20140009.070 | 507 |  | 0.791 | 4 |  |  | 2886 | -15.025 | 47.5819 | 68.816 |  |  |  |  | 67.75 | 1.7 | 2.1 | 4.3 | 1.0892 |  | 1.1482
2014/03/05 05:03:31 | 94 | 20140009.070 | 507 |  | 0.791 | 5 |  |  | 4090 | 0 | 55.7247 | 80.593 |  |  |  |  | 79.335 | 4.6 | 5.4 | 8.4 | 1.1057 |  | 1.1641
2014/03/05 05:12:01 | 95 | 20140009.071 | 1601 |  | 0.842 | 1 | 3858 | 0.006 |  |  | 10.1918 | 76.167 | 75.587 | 20.0 | 16.1 | 512.4 | [5 empty] | 0.36647 |  | 0.71924
2014/03/05 05:12:01 | 95 | 20140009.071 | 1601 |  | 0.842 | 2 | 3856 | 0 |  |  | 10.1862 | 76.125 | 75.546 | 19.5 | 15.6 | 526.0 | [5 empty] | 0.36647 |  | 0.71923
2014/03/05 05:12:01 | 95 | 20140009.071 | 1601 |  | 0.842 | 3 | 3502 | 13.813 |  |  | 10.2461 | 76.573 | 75.978 | 19.3 | 15.2 | 503.5 | [5 empty] | 0.37151 |  | 0.72917
2014/03/05 05:12:01 | 95 | 20140009.071 | 1601 |  | 0.842 | 4 |  |  | 3063 | -15.334 | 47.5419 | 73.192 |  |  |  |  | 72.058 | 1.7 | 2.0 | 4.3 | 1.0889 |  | 1.1478
2014/03/05 05:12:01 | 95 | 20140009.071 | 1601 |  | 0.842 | 5 |  |  | 4091 | 0 | 52.3119 | 80.535 |  |  |  |  | 79.278 | 4.7 | 5.4 | 8.5 | 1.1057 |  | 1.164
2014/03/05 05:20:32 | 96 | 20140009.072 | 1602 |  | 0.842 | 1 | 3856 | -0.001 |  |  | 10.1832 | 76.103 | 75.523 | 20.0 | 16.1 | 512.1 | [5 empty] | 0.36647 |  | 0.71917
2014/03/05 05:20:32 | 96 | 20140009.072 | 1602 |  | 0.842 | 2 | 3857 | 0 |  |  | 10.1876 | 76.136 | 75.556 | 19.5 | 15.6 | 526.2 | [5 empty] | 0.36647 |  | 0.71917
2014/03/05 05:20:32 | 96 | 20140009.072 | 1602 |  | 0.842 | 3 | 3505 | 15.057 |  |  | 10.2296 | 76.45 | 75.855 | 19.2 | 15.2 | 503.2 | [5 empty] | 0.37197 |  | 0.73
2014/03/05 05:20:32 | 96 | 20140009.072 | 1602 |  | 0.842 | 4 |  |  | 3005 | -15.443 | 46.6389 | 71.801 |  |  |  |  | 70.689 | 1.7 | 2.1 | 4.3 | 1.0888 |  | 1.1477
2014/03/05 05:20:32 | 96 | 20140009.072 | 1602 |  | 0.842 | 5 |  |  | 4089 | 0 | 52.2939 | 80.507 |  |  |  |  | 79.251 | 4.6 | 5.4 | 8.5 | 1.1057 |  | 1.164
2014/03/05 05:29:02 | 97 | 36-UWSIF-UT Glut 120140009.2 | 36-UWSIF-UT Glut 1 |  | 0.7704 | 1 | 3858 | -0.023 |  |  | 11.1347 | 76.138 | 75.559 | 19.9 | 16.0 | 511.5 | [5 empty] | 0.36646 |  | 0.71783
2014/03/05 05:29:02 | 97 | 36-UWSIF-UT Glut 120140009.2 | 36-UWSIF-UT Glut 1 |  | 0.7704 | 2 | 3858 | 0 |  |  | 11.1374 | 76.157 | 75.578 | 19.5 | 15.5 | 525.2 | [5 empty] | 0.36647 |  | 0.71785
2014/03/05 05:29:02 | 97 | 36-UWSIF-UT Glut 120140009.2 | 36-UWSIF-UT Glut 1 |  | 0.7704 | 3 | 2845 | -4.21 |  |  | 9.0773 | 62.07 | 61.597 | 19.2 | 15.1 | 502.1 | [5 empty] | 0.36494 |  | 0.71483
2014/03/05 05:29:02 | 97 | 36-UWSIF-UT Glut 120140009.2 | 36-UWSIF-UT Glut 1 |  | 0.7704 | 4 |  |  | 2294 | -14.266 | 38.0667 | 53.621 |  |  |  |  | 52.789 | 1.7 | 2.0 | 4.3 | 1.0901 |  | 1.1491
2014/03/05 05:29:02 | 97 | 36-UWSIF-UT Glut 120140009.2 | 36-UWSIF-UT Glut 1 |  | 0.7704 | 5 |  |  | 4089 | 0 | 57.1273 | 80.47 |  |  |  |  | 79.214 | 4.2 | 4.9 | 7.9 | 1.1057 |  | 1.1641
2014/03/05 05:37:32 | 98 | 36-UWSIF-UT Glut 120140009.21 | 36-UWSIF-UT Glut 1 |  | 0.7738 | 1 | 3857 | -0.052 |  |  | 11.0893 | 76.162 | 75.582 | 19.7 | 15.9 | 508.8 | [5 empty] | 0.36645 |  | 0.71925
2014/03/05 05:37:32 | 98 | 36-UWSIF-UT Glut 120140009.21 | 36-UWSIF-UT Glut 1 |  | 0.7738 | 2 | 3860 | 0 |  |  | 11.0887 | 76.158 | 75.578 | 19.3 | 15.4 | 523.1 | [5 empty] | 0.36647 |  | 0.71929
2014/03/05 05:37:32 | 98 | 36-UWSIF-UT Glut 120140009.21 | 36-UWSIF-UT Glut 1 |  | 0.7738 | 3 | 2851 | -4.238 |  |  | 9.0783 | 62.351 | 61.875 | 19.1 | 15.0 | 500.4 | [5 empty] | 0.36492 |  | 0.71624
2014/03/05 05:37:32 | 98 | 36-UWSIF-UT Glut 120140009.21 | 36-UWSIF-UT Glut 1 |  | 0.7738 | 4 |  |  | 2294 | -14.275 | 37.9928 | 53.753 |  |  |  |  | 52.919 | 1.6 | 2.0 | 4.2 | 1.09 |  | 1.1491
2014/03/05 05:37:32 | 98 | 36-UWSIF-UT Glut 120140009.21 | 36-UWSIF-UT Glut 1 |  | 0.7738 | 5 |  |  | 4093 | 0 | 56.9379 | 80.557 |  |  |  |  | 79.3 | 4.2 | 4.9 | 7.9 | 1.1057 |  | 1.1641
2014/03/05 05:46:02 | 99 | 39-UWSIF-UW Glut 220140009.17 | 39-UWSIF-UW Glut 2 |  | 0.7692 | 1 | 3862 | -0.072 |  |  | 11.1602 | 76.193 | 75.614 | 19.7 | 15.8 | 507.0 | [5 empty] | 0.36645 |  | 0.71758
2014/03/05 05:46:02 | 99 | 39-UWSIF-UW Glut 220140009.17 | 39-UWSIF-UW Glut 2 |  | 0.7692 | 2 | 3864 | 0 |  |  | 11.1583 | 76.18 | 75.601 | 19.3 | 15.3 | 521.3 | [5 empty] | 0.36647 |  | 0.71764
2014/03/05 05:46:02 | 99 | 39-UWSIF-UW Glut 220140009.17 | 39-UWSIF-UW Glut 2 |  | 0.7692 | 3 | 3093 | 28.354 |  |  | 9.9117 | 67.67 | 67.139 | 19.0 | 15.0 | 498.9 | [5 empty] | 0.37682 |  | 0.73798
2014/03/05 05:46:02 | 99 | 39-UWSIF-UW Glut 220140009.17 | 39-UWSIF-UW Glut 2 |  | 0.7692 | 4 |  |  | 2477 | 36.926 | 41.5063 | 58.375 |  |  |  |  | 57.438 | 1.6 | 1.9 | 4.2 | 1.146 |  | 1.2052
2014/03/05 05:46:02 | 99 | 39-UWSIF-UW Glut 220140009.17 | 39-UWSIF-UW Glut 2 |  | 0.7692 | 5 |  |  | 4092 | 0 | 57.2509 | 80.518 |  |  |  |  | 79.261 | 4.3 | 5.2 | 8.1 | 1.1057 |  | 1.1647
2014/03/05 05:54:32 | 100 | 39-UWSIF-UW Glut 220140009.18 | 39-UWSIF-UW Glut 2 |  | 0.7779 | 1 | 3862 | -0.072 |  |  | 11.0409 | 76.231 | 75.652 | 19.7 | 15.9 | 507.3 | [5 empty] | 0.36645 |  | 0.71773
2014/03/05 05:54:32 | 100 | 39-UWSIF-UW Glut 220140009.18 | 39-UWSIF-UW Glut 2 |  | 0.7779 | 2 | 3862 | 0 |  |  | 11.0291 | 76.15 | 75.571 | 19.3 | 15.4 | 521.2 | [5 empty] | 0.36647 |  | 0.71778
2014/03/05 05:54:32 | 100 | 39-UWSIF-UW Glut 220140009.18 | 39-UWSIF-UW Glut 2 |  | 0.7779 | 3 | 3112 | 28.325 |  |  | 9.8759 | 68.188 | 67.653 | 19.0 | 15.0 | 498.9 | [5 empty] | 0.37681 |  | 0.73811
2014/03/05 05:54:32 | 100 | 39-UWSIF-UW Glut 220140009.18 | 39-UWSIF-UW Glut 2 |  | 0.7779 | 4 |  |  | 2494 | 37.178 | 41.3991 | 58.883 |  |  |  |  | 57.937 | 1.6 | 2.0 | 4.2 | 1.1463 |  | 1.2055
2014/03/05 05:54:32 | 100 | 39-UWSIF-UW Glut 220140009.18 | 39-UWSIF-UW Glut 2 |  | 0.7779 | 5 |  |  | 4089 | 0 | 56.6814 | 80.619 |  |  |  |  | 79.36 | 4.3 | 5.2 | 8.1 | 1.1057 |  | 1.1647
2014/03/05 06:03:02 | 101 | 01-UWSIF-Liver 20140009.17 | 01-UWSIF-Liver |  | 0.7536 | 1 | 3862 | -0.055 |  |  | 11.3966 | 76.229 | 75.65 | 19.7 | 15.9 | 507.2 | [5 empty] | 0.36645 |  | 0.71763
2014/03/05 06:03:02 | 101 | 01-UWSIF-Liver 20140009.17 | 01-UWSIF-Liver |  | 0.7536 | 2 | 3863 | 0 |  |  | 11.3838 | 76.144 | 75.565 | 19.3 | 15.4 | 521.4 | [5 empty] | 0.36647 |  | 0.71767
2014/03/05 06:03:02 | 101 | 01-UWSIF-Liver 20140009.17 | 01-UWSIF-Liver |  | 0.7536 | 3 | 3832 | 7.15 |  |  | 12.4482 | 83.263 | 82.62 | 19.0 | 15.1 | 498.5 | [5 empty] | 0.36908 |  | 0.72281
2014/03/05 06:03:02 | 101 | 01-UWSIF-Liver 20140009.17 | 01-UWSIF-Liver |  | 0.7536 | 4 |  |  | 2744 | -3.864 | 46.9686 | 64.717 |  |  |  |  | 63.707 | 1.6 | 2.0 | 4.2 | 1.1014 |  | 1.1605
2014/03/05 06:03:02 | 101 | 01-UWSIF-Liver 20140009.17 | 01-UWSIF-Liver |  | 0.7536 | 5 |  |  | 4091 | 0 | 58.4605 | 80.552 |  |  |  |  | 79.295 | 4.5 | 5.3 | 8.2 | 1.1057 |  | 1.1642
2014/03/05 06:11:33 | 102 | 01-UWSIF-Liver 20140009.18 | 01-UWSIF-Liver |  | 0.7296 | 1 | 3863 | -0.011 |  |  | 11.7727 | 76.237 | 75.657 | 19.8 | 16.0 | 510.0 | [5 empty] | 0.36647 |  | 0.71782
2014/03/05 06:11:33 | 102 | 01-UWSIF-Liver 20140009.18 | 01-UWSIF-Liver |  | 0.7296 | 2 | 3864 | 0 |  |  | 11.7637 | 76.179 | 75.6 | 19.3 | 15.4 | 523.6 | [5 empty] | 0.36647 |  | 0.71782
2014/03/05 06:11:33 | 102 | 01-UWSIF-Liver 20140009.18 | 01-UWSIF-Liver |  | 0.7296 | 3 | 3700 | 7.222 |  |  | 12.4858 | 80.855 | 80.231 | 19.1 | 15.1 | 501.3 | [5 empty] | 0.36911 |  | 0.72301
2014/03/05 06:11:33 | 102 | 01-UWSIF-Liver 20140009.18 | 01-UWSIF-Liver |  | 0.7296 | 4 |  |  | 2657 | -3.959 | 47.0333 | 62.743 |  |  |  |  | 61.763 | 1.7 | 2.0 | 4.3 | 1.1013 |  | 1.1603
2014/03/05 06:11:33 | 102 | 01-UWSIF-Liver 20140009.18 | 01-UWSIF-Liver |  | 0.7296 | 5 |  |  | 4089 | 0 | 60.3832 | 80.551 |  |  |  |  | 79.294 | 4.4 | 5.2 | 8.2 | 1.1057 |  | 1.1641
Analyst: DEW
Sample Count: 0706
Sample Count:
Helium Pressure: 1700
Oxygen pressure: 600
Mass 28:21
Mass 29:16
Mass 30: 246
Peak Center:3.065
Instrument: ANNIE
Data analyst:
Comments:

## Analysis Information
Principal of operation
●The Finnigan DeltaPlus XP is run in continuous flow mode and connected to either a Carlo Erba 1110 or a Costech 4010 elemental analyzer via a Finnigan ConFlo III interface.
●A sample contained within a tin capsule is dropped into a combustion reactor held at 1020°C. For nitrogen and carbon analysis, the combustion reactor contains chromium oxide for oxidation and silvered cobaltous/cobaltic oxide for removal of sulfur. When the tin capsule is exposed to a gas flow temporarily enriched with ultra-high purity oxygen (25-30 mL min-1), flash combustion occurs which raises the temperature of the sample to >1700°C. The encapsulated sample, depending on its composition, combusts generating one or more of these gases: N2, NxOx, CO2, and H2O. The reduction reactor contains reduced copper wires for the reduction of nitrogen oxides to N2 and the removal of excess O2. An adsorption trap containing magnesium perchlorate removes the H2O. The remaining N2 and CO2 gases travel through a Porapak Q chromatographic column (80-100 mL min-1 at 50˚C) as the carrier gas and then moves to the ConFlo III open-split interface.
●The gas then moves through the ConFlo III to the Finnigan DeltaPlus XP mass spectrometer. Reference gases are pulsed by the ConFlo III into the mass spectrometer for proper mass balance.
QA/QC
Quality assurance of carbon and nitrogen isotope composition of solids is based on the standard uncertainty of the known value of the secondary laboratory reference material calculated on multiple analyses. For carbon isotope composition, if the standard uncertainty is greater than 0.15‰, the unknowns are re-analyzed (until the 2-sigma expanded standard uncertainty of the result is better than 0.3‰). The carbon isotopic composition is reported in per mil relative to VPDB scale such that USGS40 glutamic acid, USGS41 glutamic acid, and NIST 8542 sucrose ANU, respectively are -26.24‰, +37.76‰, and -10.05‰. For nitrogen isotope ratio composition, if the standard uncertainty is greater than 0.2‰, the unknowns are re-analyzed (until the 2-sigma expanded standard uncertainty of the result is better than 0.4‰). The nitrogen isotopic composition is reported in per mil relative to nitrogen in air on a scale such that USGS40 glutamic acid, USGS41 glutamic acid, and NIST8549 potassium nitrate, respectively are -4.52‰, +47.57‰, and 4.36‰.
Normalization
Whenever isotopic analyses are performed, reference materials must be included with the unknowns. In general two ‘bracketing’ reference materials should be used for correcting purposes. These reference materials ideally would be similar in chemical complexity to the unknowns but have very different isotopic compositions. Using two such reference materials, you can generate two linear equations, one for each referenece material. Since these linear equations will have the same slope and intercept, combining the two equations will generate a single linear equation that can be used to correct the unknowns.
The linear equation is derived as follows:
Slope (m): m=[δRM 1-known – δRM 2-known] / [δRM 1-measured – δRM 2-measured]
Intercept (b): b= δRM 1-known – [m* δRM 1-measured]
Normalization equation:
δSA-COR = m* δSA-measured + b
Where:
δRM 1-known = the accepted delta value of reference material 1
δRM 2-known = the accepted delta value of reference material 2
δRM 1-measured = the measured delta value of reference material 1
δRM 2-measured = the measured delta value of reference material 2
δSA-measured = measured delta value of sample
δSA-COR = corrected delta value of sample
13C and 15N analyses of solids

## Contact
For questions about the analysis, please contact:
|  | Chandelle Macdonald | Master Technician
| email: | cmacdon1@uwyo.edu
| Phone: | (307) 766-6373
Contact information for UW Stable Isotope Facility:
| Personnel: | Dr. David G. Williams | Faculty Director
dgw@uwyo.edu
|  | Craig Cook | Facility Director
ccook21@uwyo.edu
|  | Chandelle Macdonald | Master Technician
cmacdon1@uwyo.edu
| Address: | UWYO Stable Isotope Facility
University of Wyoming
Berry Biodiversity Center Rm 214
Laramie, WY 82071
| email: | uwyosif@uwyo.edu
| Phone: | (307) 766-6373
| Fax: | (307) 766-6403