## Final Report
Final Report
Job 2014-0020
|  | Isotope(s) requested: | δ13C, δ15N
|  | Instrument Used: | Costech 4010 Elemental Analyzer coupled to a Thermo Delta Plus XP IRMS
|  | Analytical Code: | 010 (combustion/reduction organics)
|  | Units: | δ13C and δ15N values are reported w.r.t. VPDB and AIR respectively in parts per thousand (per mil)
|  | Principal Investigator: | Jason Stockwell
|  | Job submission contact: | Peter Euchlide
|  | Sample Material(s): | Mysis Tissue
|  | Date Submited: | 2014-02-03
|  | Number of unknown samples analyzed: | 83
Number of reference samples analyzed:
|  | Quality Control Reference Material 1: | 36-UWSIF-Glutamic 1
|  | Quality Control Reference Material 2: | 39-UWSIF-Glutamic 2
|  | Quality Assurance Reference Material 3: | UWSIF01 (Liver)
Quality Control Data |  |  |  |  | Quality Assurance Data
Reference Material 1 | Known |  | Known |  | Reference Material 3 | Known | Normalized | Known | Normalized
36-UWSIF-Glutamic 1 | δ13C PDB |  | δ15N AIR |  | UWSIF01 (Liver) | δ13C PDB | δ13C PDB | δ15N AIR | δ15N AIR
average (N=12) | -28.28 |  | -4.62 |  | average (N=12) | -17.82 | -17.7282 | 6.78 | 6.7746
standard uncertainty | 0.073221 |  | 0.047382 |  | standard uncertainty |  | 0.041209 |  | 0.049253
Reference Material 2 | Known |  | Known |  | Long-Term | δ13C PDB | -18.12 | δ15N AIR | 6.38
39-UWSIF-Glutamic 2 | δ13C PDB |  | δ15N AIR |  | Acceptable Range | 2 σ = 0.3 | -17.52 | 2 σ = 0.4 | 7.18
average (N=12) | 24.36 |  | 27.89 |  | Weight Percent | Known | Measured | Known | Measured
standard uncertainty | 0.064352 |  | 0.058322 |  |  | Wt% C | Wt% C | Wt% N | Wt% N
[5 empty] | average (N=12) | 49.54 | 49.9452 | 12.9 | 12.7231
[5 empty] | standard uncertainty |  | 0.59565 |  | 0.18609
[5 empty] | relative error (%) |  | 1.2397 |  | 1.5631
Record Keeping |  |  |  |  | Quality Assurance Approval
Date Reported: | 2014-03-27 |  |  |  | Reviewer: | Craig Cook
Date Invoiced: | 2014-03-27 |  |  |  | Title: | Facility Director
Initial: | cjm |  |  |  | Date Reviewed: | 2014-03-27
Comments: |  |  |  |  | Comments: | Data meet all QAQC criteria in the SOP.
Analytical Comments:
*Sample weight percents are calculated using the sample weights reported by the user and are dependent upon the accuracy of these values.
Quality Control Color Legend
Yellow | Peak amplitudes too low for reliable results. Use with extreme caution or rerun the sample.
Green | Weight percent values suspect. Use with caution.
Pink | Potential outlier. Use with caution.
Blue | Sample lost during the analysis. Reload and rerun the sample.
Olive | Possible sample ID problem. Check loading documents.
SIF ID | Sample ID | δ15N | δ13C | Wt% N* | Wt% C* | C:N ratio | Comments
20140020.052 | 131 | 13.7541 | -30.0652 | 10.0769 | 50.5338 | 5.0148
20140020.022 | 133 | 12.6662 | -28.8024 | 10.6017 | 46.4309 | 4.3796
20140020.051 | 198 | 12.6472 | -30.3218 | 9.6791 | 47.9043 | 4.9492
20140020.021 | 203 | 14.6198 | -31.0896 | 9.8089 | 51.0789 | 5.2074
20140020.053 | 217 | 13.1446 | -30.5551 | 10.4486 | 50.0869 | 4.7937
20140020.023 | 246 | 12.9174 | -30.1899 | 9.0853 | 51.631 | 5.6829
20140020.054 | 275 | 13.3818 | -28.8095 | 10.7894 | 48.6571 | 4.5097
20140020.024 | 285 | 12.3219 | -28.847 | 10.7856 | 48.0366 | 4.4538
20140020.026 | 306 | [5 empty] | Sample lost while crimping.
20140020.025 | 355 | 12.5171 | -27.6958 | 8.7632 | 34.3445 | 3.9192 | Weight percent values suspect. Use with caution.
20140020.056 | 372 | 14.0563 | -27.9828 | 12.3161 | 47.8274 | 3.8833
20140020.055 | 392 | 13.2927 | -27.7657 | 13.056 | 47.1326 | 3.61
20140020.057 | 459 | 12.9254 | -28.0092 | 11.6543 | 44.4175 | 3.8112
20140020.028 | 467 | 12.9845 | -29.6341 | 10.2232 | 50.6777 | 4.9571
20140020.027 | 469 | 13.1726 | -28.917 | 11.0434 | 45.3892 | 4.1101
20140020.058 | 491 | 12.6051 | -27.5913 | 11.2719 | 42.4935 | 3.7699
20140020.059 | 510 | 13.7391 | -28.1562 | 12.5931 | 47.3112 | 3.7569
20140020.060 | 533 | 14.3065 | -28.7364 | 12.3404 | 47.098 | 3.8166
20140020.030 | 546 | 13.3988 | -28.9565 | 11.6421 | 47.704 | 4.0976
20140020.029 | 568 | 14.2815 | -28.7445 | 12.9558 | 47.1417 | 3.6386
20140020.037 | 610 | 14.1684 | -30.0378 | 11.2434 | 52.0636 | 4.6306
20140020.067 | 621 | 14.3035 | -28.9839 | 11.9985 | 46.9504 | 3.913
20140020.007 | 644 | 14.0873 | -29.4353 | 10.8188 | 49.0958 | 4.538
20140020.038 | 660 | 14.2405 | -28.8226 | 12.176 | 47.2357 | 3.8794
20140020.008 | 689 | 13.667 | -28.4088 | 13.3781 | 48.2396 | 3.6059
20140020.068 | 693 | 14.3756 | -29.1503 | 11.6949 | 49.0569 | 4.1947
20140020.035 | 702 | 14.1794 | -29.4221 | 10.5212 | 49.5654 | 4.711
20140020.066 | 712 | 14.0313 | -29.0072 | 11.3154 | 47.9515 | 4.2377
20140020.005 | 716 | 12.1308 | -28.2435 | 11.6694 | 43.5867 | 3.7351
20140020.036 | 750 | 14.0723 | -29.5174 | 10.3797 | 49.3609 | 4.7555
20140020.006 | 793 | 14.2685 | -28.7618 | 9.389 | 34.2162 | 3.6443 | Weight percent values suspect. Use with caution.
20140020.065 | 796 | 13.9252 | -28.3621 | 11.1377 | 44.5554 | 4.0004
20140020.063 | 841 | 14.0913 | -28.3743 | 12.5816 | 46.8848 | 3.7265
20140020.033 | 856 | 15.0471 | -28.5762 | 11.8076 | 45.726 | 3.8726
20140020.004 | 865 | 13.9442 | -28.2709 | 13.0408 | 46.3954 | 3.5577
20140020.064 | 871 | 13.8171 | -29.1837 | 10.9389 | 50.9797 | 4.6604
20140020.003 | 886 | 14.3225 | -28.635 | 11.7359 | 49.5208 | 4.2196
20140020.034 | 894 | 14.9951 | -28.1633 | 12.3008 | 43.729 | 3.555
20140020.062 | 902 | 13.8952 | -28.9018 | 9.7718 | 42.2447 | 4.3231
20140020.002 | 925 | 13.3708 | -29.0965 | 10.732 | 50.2187 | 4.6794
20140020.061 | 930 | 13.1086 | -28.3216 | 11.923 | 46.7132 | 3.9179
20140020.031 | 953 | 13.1566 | -29.0874 | 10.7973 | 47.1425 | 4.3661
20140020.001 | 969 | 12.1408 | -27.4655 | 4.229 | 15.6034 | 3.6897 | Weight percent values suspect. Use with caution.
20140020.032 | 978 | 13.5339 | -27.9645 | 5.0306 | 20.0788 | 3.9913
20140020.010 | 1005 | 12.8864 | -30.0205 | 10.2589 | 50.7798 | 4.9498
20140020.009 | 1009 | 14.4476 | -28.136 | 11.7413 | 46.6131 | 3.97
20140020.070 | 1068 | 14.2855 | -29.2304 | 11.0958 | 48.7666 | 4.3951
20140020.040 | 1070 | 14.1784 | -30.6961 | 10.4417 | 51.2409 | 4.9073
20140020.039 | 1089 | 13.1116 | -28.9677 | 10.8034 | 47.1221 | 4.3618
20140020.069 | 1095 | 14.3195 | -29.1756 | 12.1314 | 47.2476 | 3.8946
20140020.041 | 1103 | 13.4458 | -28.7395 | 11.7457 | 48.9128 | 4.1643
20140020.011 | 1119 | 12.8513 | -30.6048 | 10.238 | 47.1702 | 4.6074
20140020.072 | 1131 | 12.5451 | -28.1755 | 11.5218 | 45.9092 | 3.9845
20140020.071 | 1153 | 12.457 | -28.1025 | 11.7113 | 44.7104 | 3.8177
20140020.012 | 1175 | 13.3257 | -28.6269 | 10.6858 | 47.0421 | 4.4023
20140020.042 | 1186 | 12.6112 | -27.9808 | 5.3468 | 22.2918 | 4.1692 | Weight percent values suspect. Use with caution.
20140020.013 | 1202 | 14.2375 | -28.7669 | 11.9943 | 45.3699 | 3.7826
20140020.044 | 1232 | 13.7681 | -28.5386 | 11.7888 | 45.1819 | 3.8326
20140020.074 | 1235 | 14.2474 | -28.8655 | 12.3001 | 46.3065 | 3.7647
20140020.014 | 1258 | 14.3525 | -28.1238 | 10.8087 | 39.4901 | 3.6535
20140020.073 | 1261 | 13.276 | -29.1234 | 11.3477 | 46.2578 | 4.0764
20140020.043 | 1296 | 14.3175 | -28.5224 | 12.1993 | 45.5554 | 3.7343
20140020.015 | 1332 | 14.7959 | -28.3682 | 6.7784 | 25.3571 | 3.7409 | Weight percent values suspect. Use with caution.
20140020.045 | 1337 | 14.4546 | -28.1065 | 12.4503 | 45.9038 | 3.687
20140020.076 | 1340 | 14.6217 | -28.6156 | 12.3579 | 45.5469 | 3.6857
20140020.046 | 1341 | 14.0313 | -28.9352 | 12.0882 | 45.9477 | 3.801
20140020.075 | 1360 | 14.9931 | -28.2949 | 13.4433 | 49.0365 | 3.6477
20140020.016 | 1395 | 14.3325 | -28.6817 | 11.3993 | 48.0728 | 4.2172
20140020.049 | 1410 | 13.1906 | -29.0377 | 11.9036 | 48.1654 | 4.0463
20140020.020 | 1423 | 13.7961 | -28.8176 | 11.4949 | 45.9083 | 3.9938
20140020.079 | 1441 | 14.1176 | -29.3288 | 12.1027 | 51.5886 | 4.2626
20140020.019 | 1456 | 12.7032 | -27.6227 | 11.2151 | 40.7688 | 3.6352
20140020.083 | 1472 | 12.6251 | -27.3784 | 11.7614 | 43.8819 | 3.731
20140020.080 | 1477 | 14.3532 | -28.8078 | 12.8733 | 46.9847 | 3.6498
20140020.018 | 1502 | 13.2136 | -27.5761 | 11.3977 | 41.6458 | 3.6539
20140020.078 | 1540 | 13.5406 | -29.3703 | 11.0238 | 49.5289 | 4.4929
20140020.017 | 1577 | 12.7112 | -28.1573 | 10.8432 | 43.8908 | 4.0478
20140020.077 | 1580 | 13.3559 | -28.5043 | 10.845 | 45.545 | 4.1996
20140020.047 | 1596 | 12.2298 | -27.8114 | 11.3171 | 41.009 | 3.6236
20140020.048 | 1598 | 12.7022 | -30.1331 | 10.3181 | 47.657 | 4.6188
20140020.081 | 1603 | 14.1396 | -29.0567 | 12.6284 | 51.3701 | 4.0678
20140020.082 | 1604 | 13.2241 | -28.9434 | 10.149 | 49.0944 | 4.8374
20140020.050 | 1605 | 12.6352 | -29.4414 | 10.7667 | 46.3256 | 4.3027
Stable Isotope Facility

## QAQC, calculations
Line | Identifier 1 | δ15N | δ13C | δ15N | δ13C | Line | Identifier 1 | δ15N | δ13C | δ15N | δ13C | Line | Identifier 1 | Absolute Difference |  | Relative Error |  | δ15N |  | δ13C
|  | Normalized | Normalized | known | known |  |  | Normalized | Normalized | known | known |  |  | % Nitrogen | % Carbon | % Nitrogen | % Carbon | normalized |  | normalized
QC Reference Materials | [5 empty] | QC Reference Materials | [5 empty] | QA Reference Material
4 | 36-UWSIF-UT Glut 120140020.24 | -4.5886 | -28.1349 |  |  | 6 | 39-UWSIF-UW Glut 220140020.21 | 27.8103 | 24.3176 |  |  | 8 | 01-UWSIF-Liver 20140020.21 | 0.12174 | 0.61171 | 13.0217 | 50.1517 |  | 6.8315 | -17.6978
5 | 36-UWSIF-UT Glut 120140020.25 | -4.5746 | -28.1958 |  |  | 7 | 39-UWSIF-UW Glut 220140020.22 | 27.8643 | 24.42 |  |  | 9 | 01-UWSIF-Liver 20140020.22 | 0.026827 | 0.28058 | 12.9268 | 49.8206 |  | 6.7044 | -17.7566
34 | 36-UWSIF-UT Glut 120140020.26 | -4.5125 | -28.3033 |  |  | 36 | 39-UWSIF-UW Glut 220140020.23 | 27.9424 | 24.3297 |  |  | 38 | 01-UWSIF-Liver 20140020.23 | 0.16792 | 0.097358 | 12.7321 | 49.4426 |  | 6.8345 | -17.6735
35 | 36-UWSIF-UT Glut 120140020.27 | -4.6076 | -28.2952 |  |  | 37 | 39-UWSIF-UW Glut 220140020.24 | 27.9354 | 24.3886 |  |  | 39 | 01-UWSIF-Liver 20140020.24 | 0.15029 | 0.14266 | 12.7497 | 49.3973 |  | 6.8295 | -17.7434
64 | 36-UWSIF-UT Glut 120140020.28 | -4.5816 | -28.3104 |  |  | 66 | 39-UWSIF-UW Glut 220140020.25 | 27.9184 | 24.4139 |  |  | 68 | 01-UWSIF-Liver 20140020.25 | 0.46906 | 0.55516 | 12.4309 | 48.9848 |  | 6.7765 | -17.712
65 | 36-UWSIF-UT Glut 120140020.29 | -4.6587 | -28.3084 |  |  | 67 | 39-UWSIF-UW Glut 220140020.26 | 27.8804 | 24.4291 |  |  | 69 | 01-UWSIF-Liver 20140020.26 | 0.43919 | 0.45877 | 12.4608 | 49.0812 |  | 6.7965 | -17.7576
94 | 36-UWSIF-UT Glut 120140020.211 | -4.6887 | -28.353 |  |  | 96 | 39-UWSIF-UW Glut 220140020.27 | 27.8573 | 24.2709 |  |  | 98 | 01-UWSIF-Liver 20140020.27 | 0.38052 | 0.51067 | 12.5195 | 50.0507 |  | 6.7605 | -17.7495
95 | 36-UWSIF-UT Glut 120140020.212 | -4.7167 | -28.3378 |  |  | 97 | 39-UWSIF-UW Glut 220140020.28 | 27.9004 | 24.3399 |  |  | 99 | 01-UWSIF-Liver 20140020.28 | 0.32027 | 0.64225 | 12.5797 | 50.1823 |  | 6.6974 | -17.8144
4 | 36-UWSIF-UT Glut 120140020.34 | -4.6544 | -28.2312 |  |  | 6 | 39-UWSIF-UW Glut 220140020.31 | 27.972 | 24.4024 |  |  | 8 | 01-UWSIF-Liver 20140020.31 | 0.0815 | 1.2053 | 12.8185 | 50.7453 |  | 6.7521 | -17.7075
5 | 36-UWSIF-UT Glut 120140020.35 | -4.5586 | -28.3243 |  |  | 7 | 39-UWSIF-UW Glut 220140020.32 | 27.8822 | 24.4216 |  |  | 9 | 01-UWSIF-Liver 20140020.32 | 0.0959 | 0.8798 | 12.8041 | 50.4198 |  | 6.814 | -17.7419
21 | 36-UWSIF-UT Glut 120140020.36 | -4.6055 | -28.2494 |  |  | 23 | 39-UWSIF-UW Glut 220140020.33 | 27.8682 | 24.199 |  |  | 25 | 01-UWSIF-Liver 20140020.33 | 0.0892 | 0.9877 | 12.8108 | 50.5277 |  | 6.7801 | -17.7187
22 | 36-UWSIF-UT Glut 120140020.37 | -4.6455 | -28.3111 |  |  | 24 | 39-UWSIF-UW Glut 220140020.34 | 27.8302 | 24.4266 |  |  | 26 | 01-UWSIF-Liver 20140020.34 | 0.0772 | 0.9978 | 12.8228 | 50.5378 |  | 6.7182 | -17.665
12 | average | -4.6161 | -28.2796 | -4.62 | -28.28 | 12 | average | 27.8885 | 24.3633 | 27.89 | 24.36 | 12 |  |  | Known % | 12.9 | 49.54 | δ15N known | 6.78 | δ13C known | -17.82
| std uncertainty | 0.058322 | 0.064352 |  |  |  | std uncertainty | 0.047382 | 0.073221 | [5 empty] | Average | 12.7231 | 49.9452 | average δ15N | 6.7746 | average δ13C | -17.7282
[15 empty] | Stdev | 0.18609 | 0.59565 | stdev δ15N | 0.049253 | stdev δ13C | 0.041209
[15 empty] | Absolute Error | 0.20163 | 0.61415
[15 empty] | Relative Error | 1.5631 | 1.2397 | Long-Term | 7.18 | Long-Term | -17.52
[18 empty] | Acceptable Range | 6.38 | Acceptable Range | -18.12

## Run 1
Date | Line | Identifier 1 | Identifier 2 | Amount | Ampl 28 | d 15N/14N | Ampl 44 | d 13C/12C | Nitrogen % | Carbon % | Drift Corr for C% | C:N | Corr. d15N | Corr. d13C
2014/03/16 14:46:40 | 10 | 20140020.001 | 969 | 0.793 | 1052 | 12.596 | 711 | -13.462 | 4.229 | 15.423 | 15.6034 | 3.6897 | 12.1408 | -27.4655 | Some sample fell out of tin, weigh percents not accurate
2014/03/16 14:55:08 | 11 | 20140020.002 | 925 | 0.765 | 2617 | 13.825 | 2178 | -15.07 | 10.732 | 50.0202 | 50.2187 | 4.6794 | 13.3708 | -29.0965
2014/03/16 15:03:36 | 12 | 20140020.003 | 886 | 0.849 | 3206 | 14.776 | 2383 | -14.615 | 11.7359 | 49.3042 | 49.5208 | 4.2196 | 14.3225 | -28.635
2014/03/16 15:12:05 | 13 | 20140020.004 | 865 | 0.839 | 3515 | 14.398 | 2213 | -14.256 | 13.0408 | 46.1608 | 46.3954 | 3.5577 | 13.9442 | -28.2709
2014/03/16 15:20:33 | 14 | 20140020.005 | 716 | 0.823 | 3082 | 12.586 | 2041 | -14.229 | 11.6694 | 43.334 | 43.5867 | 3.7351 | 12.1308 | -28.2435
2014/03/16 15:29:02 | 15 | 20140020.006 | 793 | 0.852 | 2551 | 14.722 | 1664 | -14.74 | 9.389 | 33.9454 | 34.2162 | 3.6443 | 14.2685 | -28.7618 | Some sample fell out of tin, weigh percents not accurate
2014/03/16 15:37:30 | 16 | 20140020.007 | 644 | 0.811 | 2810 | 14.541 | 2254 | -15.404 | 10.8188 | 48.807 | 49.0958 | 4.538 | 14.0873 | -29.4353
2014/03/16 15:45:59 | 17 | 20140020.008 | 689 | 0.782 | 3347 | 14.121 | 2140 | -14.392 | 13.3781 | 47.9328 | 48.2396 | 3.6059 | 13.667 | -28.4088
2014/03/16 15:54:27 | 18 | 20140020.009 | 1009 | 0.785 | 2949 | 14.901 | 2074 | -14.123 | 11.7413 | 46.2882 | 46.6131 | 3.97 | 14.4476 | -28.136
2014/03/16 16:02:56 | 19 | 20140020.010 | 1005 | 0.794 | 2626 | 13.341 | 2292 | -15.981 | 10.2589 | 50.4369 | 50.7798 | 4.9498 | 12.8864 | -30.0205
2014/03/16 16:11:24 | 20 | 20140020.011 | 1119 | 0.81 | 2658 | 13.306 | 2170 | -16.557 | 10.238 | 46.8092 | 47.1702 | 4.6074 | 12.8513 | -30.6048
2014/03/16 16:19:53 | 21 | 20140020.012 | 1175 | 0.839 | 2887 | 13.78 | 2242 | -14.607 | 10.6858 | 46.663 | 47.0421 | 4.4023 | 13.3257 | -28.6269
2014/03/16 16:28:21 | 22 | 20140020.013 | 1202 | 0.804 | 3093 | 14.691 | 2068 | -14.745 | 11.9943 | 44.9728 | 45.3699 | 3.7826 | 14.2375 | -28.7669
2014/03/16 16:36:50 | 23 | 20140020.014 | 1258 | 0.813 | 2817 | 14.806 | 1830 | -14.111 | 10.8087 | 39.075 | 39.4901 | 3.6535 | 14.3525 | -28.1238
2014/03/16 16:45:19 | 24 | 20140020.015 | 1332 | 0.776 | 1672 | 15.249 | 1129 | -14.352 | 6.7784 | 24.9239 | 25.3571 | 3.7409 | 14.7959 | -28.3682 | Some sample fell out of tin, weigh percents not accurate
2014/03/16 16:53:47 | 25 | 20140020.016 | 1395 | 0.802 | 2942 | 14.786 | 2189 | -14.661 | 11.3993 | 47.6215 | 48.0728 | 4.2172 | 14.3325 | -28.6817
2014/03/16 17:02:16 | 26 | 20140020.017 | 1577 | 0.796 | 2775 | 13.166 | 1990 | -14.144 | 10.8432 | 43.4215 | 43.8908 | 4.0478 | 12.7112 | -28.1573
2014/03/16 17:10:45 | 27 | 20140020.018 | 1502 | 0.775 | 2829 | 13.668 | 1842 | -13.571 | 11.3977 | 41.1585 | 41.6458 | 3.6539 | 13.2136 | -27.5761
2014/03/16 17:19:13 | 28 | 20140020.019 | 1456 | 0.741 | 2652 | 13.158 | 1717 | -13.617 | 11.2151 | 40.2634 | 40.7688 | 3.6352 | 12.7032 | -27.6227
2014/03/16 17:27:42 | 29 | 20140020.020 | 1423 | 0.638 | 2344 | 14.25 | 1673 | -14.795 | 11.4949 | 45.3849 | 45.9083 | 3.9938 | 13.7961 | -28.8176
2014/03/16 17:36:11 | 30 | 20140020.021 | 203 | 0.761 | 2405 | 15.073 | 2211 | -17.035 | 9.8089 | 50.5374 | 51.0789 | 5.2074 | 14.6198 | -31.0896
2014/03/16 17:44:41 | 31 | 20140020.022 | 133 | 0.76 | 2601 | 13.121 | 2025 | -14.78 | 10.6017 | 45.8714 | 46.4309 | 4.3796 | 12.6662 | -28.8024
2014/03/16 17:53:10 | 32 | 20140020.023 | 246 | 0.816 | 2394 | 13.372 | 2390 | -16.148 | 9.0853 | 51.0534 | 51.631 | 5.6829 | 12.9174 | -30.1899
2014/03/16 18:01:39 | 33 | 20140020.024 | 285 | 0.806 | 2788 | 12.777 | 2187 | -14.824 | 10.7856 | 47.441 | 48.0366 | 4.4538 | 12.3219 | -28.847
2014/03/16 19:01:03 | 40 | 20140020.025 | 355 | 0.851 | 2382 | 12.972 | 1651 | -13.689 | 8.7632 | 33.6226 | 34.3445 | 3.9192 | 12.5171 | -27.6958 | Some sample fell out of tin, weigh percents not accurate
2014/03/16 19:09:32 | 41 | 20140020.026 | 306 | 0.856 | [10 empty] | Sample lost while crimping.
2014/03/16 19:18:01 | 42 | 20140020.027 | 469 | 0.808 | 2878 | 13.627 | 2080 | -14.893 | 11.0434 | 44.6311 | 45.3892 | 4.1101 | 13.1726 | -28.917
2014/03/16 19:26:30 | 43 | 20140020.028 | 467 | 0.816 | 2687 | 13.439 | 2337 | -15.6 | 10.2232 | 49.9015 | 50.6777 | 4.9571 | 12.9845 | -29.6341
2014/03/16 19:35:00 | 44 | 20140020.029 | 568 | 0.725 | 3000 | 14.735 | 1930 | -14.723 | 12.9558 | 46.3476 | 47.1417 | 3.6386 | 14.2815 | -28.7445
2014/03/16 19:43:29 | 45 | 20140020.030 | 546 | 0.824 | 3068 | 13.853 | 2198 | -14.932 | 11.6421 | 46.8918 | 47.704 | 4.0976 | 13.3988 | -28.9565
2014/03/16 19:51:59 | 46 | 20140020.031 | 953 | 0.829 | 2877 | 13.611 | 2198 | -15.061 | 10.7973 | 46.3123 | 47.1425 | 4.3661 | 13.1566 | -29.0874
2014/03/16 20:00:28 | 47 | 20140020.032 | 978 | 0.806 | 1282 | 13.988 | 906 | -13.954 | 5.0306 | 19.2305 | 20.0788 | 3.9913 | 13.5339 | -27.9645
2014/03/16 20:08:57 | 48 | 20140020.033 | 856 | 0.819 | 3099 | 15.5 | 2101 | -14.557 | 11.8076 | 44.8597 | 45.726 | 3.8726 | 15.0471 | -28.5762
2014/03/16 20:17:26 | 49 | 20140020.034 | 894 | 0.76 | 2985 | 15.448 | 1870 | -14.15 | 12.3008 | 42.8446 | 43.729 | 3.555 | 14.9951 | -28.1633
2014/03/16 20:25:55 | 50 | 20140020.035 | 702 | 0.805 | 2727 | 14.633 | 2243 | -15.391 | 10.5212 | 48.6629 | 49.5654 | 4.711 | 14.1794 | -29.4221
2014/03/16 20:34:25 | 51 | 20140020.036 | 750 | 0.771 | 2562 | 14.526 | 2134 | -15.485 | 10.3797 | 48.4404 | 49.3609 | 4.7555 | 14.0723 | -29.5174
2014/03/16 20:42:54 | 52 | 20140020.037 | 610 | 0.818 | 2952 | 14.622 | 2390 | -15.998 | 11.2434 | 51.1251 | 52.0636 | 4.6306 | 14.1684 | -30.0378
2014/03/16 20:51:23 | 53 | 20140020.038 | 660 | 0.775 | 3017 | 14.694 | 2052 | -14.8 | 12.176 | 46.2791 | 47.2357 | 3.8794 | 14.2405 | -28.8226
2014/03/16 20:59:53 | 54 | 20140020.039 | 1089 | 0.779 | 2697 | 13.566 | 2062 | -14.943 | 10.8034 | 46.1475 | 47.1221 | 4.3618 | 13.1116 | -28.9677
2014/03/16 21:08:23 | 55 | 20140020.040 | 1070 | 0.777 | 2607 | 14.632 | 2235 | -16.647 | 10.4417 | 50.2482 | 51.2409 | 4.9073 | 14.1784 | -30.6961
2014/03/16 21:16:52 | 56 | 20140020.041 | 1103 | 0.805 | 3031 | 13.9 | 2204 | -14.718 | 11.7457 | 47.902 | 48.9128 | 4.1643 | 13.4458 | -28.7395
2014/03/16 21:25:22 | 57 | 20140020.042 | 1186 | 0.792 | 1333 | 13.066 | 979 | -13.97 | 5.3468 | 21.263 | 22.2918 | 4.1692 | 12.6112 | -27.9808 | Some sample fell out of tin, weigh percents not accurate
2014/03/16 21:33:51 | 58 | 20140020.043 | 1296 | 0.803 | 3135 | 14.771 | 2046 | -14.504 | 12.1993 | 44.5086 | 45.5554 | 3.7343 | 14.3175 | -28.5224
2014/03/16 21:42:21 | 59 | 20140020.044 | 1232 | 0.75 | 2831 | 14.222 | 1906 | -14.52 | 11.7888 | 44.117 | 45.1819 | 3.8326 | 13.7681 | -28.5386
2014/03/16 21:50:50 | 60 | 20140020.045 | 1337 | 0.842 | 3367 | 14.908 | 2166 | -14.094 | 12.4503 | 44.8209 | 45.9038 | 3.687 | 14.4546 | -28.1065
2014/03/16 21:59:20 | 61 | 20140020.046 | 1341 | 0.805 | 3088 | 14.485 | 2048 | -14.911 | 12.0882 | 44.8467 | 45.9477 | 3.801 | 14.0313 | -28.9352
2014/03/16 22:07:49 | 62 | 20140020.047 | 1596 | 0.805 | 2911 | 12.685 | 1847 | -13.803 | 11.3171 | 39.89 | 41.009 | 3.6236 | 12.2298 | -27.8114
2014/03/16 22:16:19 | 63 | 20140020.048 | 1598 | 0.845 | 2793 | 13.157 | 2248 | -16.092 | 10.3181 | 46.5199 | 47.657 | 4.6188 | 12.7022 | -30.1331
2014/03/16 23:15:48 | 70 | 20140020.049 | 1410 | 0.768 | 2928 | 13.645 | 2067 | -15.012 | 11.9036 | 46.902 | 48.1654 | 4.0463 | 13.1906 | -29.0377
2014/03/16 23:24:17 | 71 | 20140020.050 | 1605 | 0.766 | 2638 | 13.09 | 1983 | -15.41 | 10.7667 | 45.0441 | 46.3256 | 4.3027 | 12.6352 | -29.4414
2014/03/16 23:32:47 | 72 | 20140020.051 | 198 | 0.843 | 2612 | 13.102 | 2244 | -16.278 | 9.6791 | 46.6047 | 47.9043 | 4.9492 | 12.6472 | -30.3218
2014/03/16 23:41:17 | 73 | 20140020.052 | 131 | 0.813 | 2626 | 14.208 | 2285 | -16.025 | 10.0769 | 49.2163 | 50.5338 | 5.0148 | 13.7541 | -30.0652
2014/03/16 23:49:47 | 74 | 20140020.053 | 217 | 0.743 | 2494 | 13.599 | 2088 | -16.508 | 10.4486 | 48.7512 | 50.0869 | 4.7937 | 13.1446 | -30.5551
2014/03/16 23:58:17 | 75 | 20140020.054 | 275 | 0.796 | 2758 | 13.836 | 2159 | -14.787 | 10.7894 | 47.3034 | 48.6571 | 4.5097 | 13.3818 | -28.8095
2014/03/17 00:06:47 | 76 | 20140020.055 | 392 | 0.808 | 3376 | 13.747 | 2116 | -13.758 | 13.056 | 45.7608 | 47.1326 | 3.61 | 13.2927 | -27.7657
2014/03/17 00:15:16 | 77 | 20140020.056 | 372 | 0.78 | 3068 | 14.51 | 2072 | -13.972 | 12.3161 | 46.4376 | 47.8274 | 3.8833 | 14.0563 | -27.9828
2014/03/17 00:23:46 | 78 | 20140020.057 | 459 | 0.845 | 3163 | 13.38 | 2084 | -13.998 | 11.6543 | 43.0097 | 44.4175 | 3.8112 | 12.9254 | -28.0092
2014/03/17 00:32:16 | 79 | 20140020.058 | 491 | 0.807 | 2931 | 13.06 | 1921 | -13.586 | 11.2719 | 41.0676 | 42.4935 | 3.7699 | 12.6051 | -27.5913
2014/03/17 00:40:45 | 80 | 20140020.059 | 510 | 0.81 | 3275 | 14.193 | 2135 | -14.143 | 12.5931 | 45.8672 | 47.3112 | 3.7569 | 13.7391 | -28.1562
2014/03/17 00:49:15 | 81 | 20140020.060 | 533 | 0.802 | 3194 | 14.76 | 2113 | -14.715 | 12.3404 | 45.636 | 47.098 | 3.8166 | 14.3065 | -28.7364
2014/03/17 00:57:44 | 82 | 20140020.061 | 930 | 0.792 | 3017 | 13.563 | 2046 | -14.306 | 11.923 | 45.2331 | 46.7132 | 3.9179 | 13.1086 | -28.3216
2014/03/17 01:06:14 | 83 | 20140020.062 | 902 | 0.753 | 2350 | 14.349 | 1766 | -14.878 | 9.7718 | 40.7466 | 42.2447 | 4.3231 | 13.8952 | -28.9018
2014/03/17 01:14:44 | 84 | 20140020.063 | 841 | 0.777 | 3135 | 14.545 | 2026 | -14.358 | 12.5816 | 45.3687 | 46.8848 | 3.7265 | 14.0913 | -28.3743
2014/03/17 01:23:14 | 85 | 20140020.064 | 871 | 0.819 | 2903 | 14.271 | 2338 | -15.156 | 10.9389 | 49.4456 | 50.9797 | 4.6604 | 13.8171 | -29.1837
2014/03/17 01:31:44 | 86 | 20140020.065 | 796 | 0.841 | 2992 | 14.379 | 2069 | -14.346 | 11.1377 | 43.0031 | 44.5554 | 4.0004 | 13.9252 | -28.3621
2014/03/17 01:40:14 | 87 | 20140020.066 | 712 | 0.811 | 2945 | 14.485 | 2160 | -14.982 | 11.3154 | 46.3813 | 47.9515 | 4.2377 | 14.0313 | -29.0072
2014/03/17 01:48:44 | 88 | 20140020.067 | 621 | 0.749 | 2867 | 14.757 | 1948 | -14.959 | 11.9985 | 45.3621 | 46.9504 | 3.913 | 14.3035 | -28.9839
2014/03/17 01:57:14 | 89 | 20140020.068 | 693 | 0.834 | 3129 | 14.829 | 2259 | -15.123 | 11.6949 | 47.4505 | 49.0569 | 4.1947 | 14.3756 | -29.1503
2014/03/17 02:05:43 | 90 | 20140020.069 | 1095 | 0.786 | 3064 | 14.773 | 2067 | -15.148 | 12.1314 | 45.6232 | 47.2476 | 3.8946 | 14.3195 | -29.1756
2014/03/17 02:14:13 | 91 | 20140020.070 | 1068 | 0.814 | 2905 | 14.739 | 2198 | -15.202 | 11.0958 | 47.1242 | 48.7666 | 4.3951 | 14.2855 | -29.2304
2014/03/17 02:22:43 | 92 | 20140020.071 | 1153 | 0.836 | 3137 | 12.912 | 2065 | -14.09 | 11.7113 | 43.0499 | 44.7104 | 3.8177 | 12.457 | -28.1025
2014/03/17 02:31:14 | 93 | 20140020.072 | 1131 | 0.827 | 3051 | 13 | 2095 | -14.162 | 11.5218 | 44.2307 | 45.9092 | 3.9845 | 12.5451 | -28.1755
Reference Check
2014/03/16 14:29:43 | 8 | 01-UWSIF-Liver 20140020.21 | 01-UWSIF-Liver | 0.7682 | 3203 | 7.291 | 2195 | -3.832 | 13.0217 | 50.0073 | 50.1517 | 3.8514 | 6.8315 | -17.6978
2014/03/16 14:38:12 | 9 | 01-UWSIF-Liver 20140020.22 | 01-UWSIF-Liver | 0.73 | 3034 | 7.164 | 2085 | -3.89 | 12.9268 | 49.6581 | 49.8206 | 3.854 | 6.7044 | -17.7566
2014/03/16 18:44:05 | 38 | 01-UWSIF-Liver 20140020.23 | 01-UWSIF-Liver | 0.7589 | 3121 | 7.294 | 2142 | -3.808 | 12.7321 | 48.7568 | 49.4426 | 3.8833 | 6.8345 | -17.6735
2014/03/16 18:52:34 | 39 | 01-UWSIF-Liver 20140020.24 | 01-UWSIF-Liver | 0.7939 | 3267 | 7.289 | 2228 | -3.877 | 12.7497 | 48.6934 | 49.3973 | 3.8744 | 6.8295 | -17.7434
2014/03/16 22:58:48 | 68 | 01-UWSIF-Liver 20140020.25 | 01-UWSIF-Liver | 0.7591 | 3049 | 7.236 | 2095 | -3.846 | 12.4309 | 47.7575 | 48.9848 | 3.9406 | 6.7765 | -17.712
2014/03/16 23:07:18 | 69 | 01-UWSIF-Liver 20140020.26 | 01-UWSIF-Liver | 0.7337 | 2946 | 7.256 | 2031 | -3.891 | 12.4608 | 47.8358 | 49.0812 | 3.9388 | 6.7965 | -17.7576
2014/03/17 03:13:46 | 98 | 01-UWSIF-Liver 20140020.27 | 01-UWSIF-Liver | 0.777 | 3137 | 7.22 | 2169 | -3.883 | 12.5195 | 48.2819 | 50.0507 | 3.9978 | 6.7605 | -17.7495
2014/03/17 03:22:16 | 99 | 01-UWSIF-Liver 20140020.28 | 01-UWSIF-Liver | 0.7235 | 2935 | 7.157 | 2031 | -3.947 | 12.5797 | 48.3954 | 50.1823 | 3.9891 | 6.6974 | -17.8144
[5 empty] | average | 7.2384 |  | -3.8717 | 12.6777 | 48.6733 | 49.6389 |  | 6.7789 | -17.7381
[5 empty] | std. dev. | 0.055076 |  | 0.042808 | 0.2175 | 0.80422 | 0.47756 |  | 0.05512 | 0.04342
Reference Material
2014/03/16 13:55:50 | 4 | 36-UWSIF-UT Glut 120140020.24 | 36-UWSIF-UT Glut 1 | 0.764 | 2302 | -4.12 | 1786 | -14.122 | 9.4962 | 40.703 | 40.7752 | 4.2939 | -4.5886 | -28.1349
2014/03/16 14:04:18 | 5 | 36-UWSIF-UT Glut 120140020.25 | 36-UWSIF-UT Glut 1 | 1.027 | 3138 | -4.106 | 2390 | -14.182 | 9.5206 | 40.7522 | 40.8424 | 4.2899 | -4.5746 | -28.1958
2014/03/16 18:10:09 | 34 | 36-UWSIF-UT Glut 120140020.26 | 36-UWSIF-UT Glut 1 | 0.7673 | 2309 | -4.044 | 1784 | -14.288 | 9.3482 | 39.9031 | 40.5168 | 4.3342 | -4.5125 | -28.3033
2014/03/16 18:18:38 | 35 | 36-UWSIF-UT Glut 120140020.27 | 36-UWSIF-UT Glut 1 | 0.757 | 2262 | -4.139 | 1746 | -14.28 | 9.3256 | 39.6972 | 40.3289 | 4.3246 | -4.6076 | -28.2952
2014/03/16 22:24:49 | 64 | 36-UWSIF-UT Glut 120140020.28 | 36-UWSIF-UT Glut 1 | 0.7298 | 2126 | -4.113 | 1649 | -14.295 | 9.1328 | 39.0134 | 40.1685 | 4.3983 | -4.5816 | -28.3104
2014/03/16 22:33:18 | 65 | 36-UWSIF-UT Glut 120140020.29 | 36-UWSIF-UT Glut 1 | 0.7903 | 2316 | -4.19 | 1789 | -14.293 | 9.1316 | 39.0109 | 40.184 | 4.4005 | -4.6587 | -28.3084
2014/03/17 02:39:44 | 94 | 36-UWSIF-UT Glut 120140020.211 | 36-UWSIF-UT Glut 1 | 0.7732 | 2269 | -4.22 | 1766 | -14.337 | 9.1336 | 39.2406 | 40.9372 | 4.482 | -4.6887 | -28.353
2014/03/17 02:48:14 | 95 | 36-UWSIF-UT Glut 120140020.212 | 36-UWSIF-UT Glut 1 | 0.7356 | 2155 | -4.248 | 1681 | -14.322 | 9.1169 | 39.1323 | 40.847 | 4.4804 | -4.7167 | -28.3378
[5 empty] | average | -4.1475 |  | -14.2649 | 9.2757 | 39.6816 | 40.575 |  | -4.6161 | -28.2799
[5 empty] | std. dev. | 0.067205 |  | 0.073881 | 0.17026 | 0.72026 | 0.31596 |  | 0.067259 | 0.074938
2014/03/16 14:12:46 | 6 | 39-UWSIF-UW Glut 220140020.21 | 39-UWSIF-UW Glut 2 | 0.7555 | 2471 | 28.253 | 1908 | 37.591 | 10.306 | 44.0456 | 44.1539 | 4.2843 | 27.8103 | 24.3176
2014/03/16 14:21:15 | 7 | 39-UWSIF-UW Glut 220140020.22 | 39-UWSIF-UW Glut 2 | 0.7701 | 2514 | 28.307 | 1934 | 37.692 | 10.2805 | 43.9325 | 44.0588 | 4.2857 | 27.8643 | 24.42
2014/03/16 18:27:07 | 36 | 39-UWSIF-UW Glut 220140020.23 | 39-UWSIF-UW Glut 2 | 0.7252 | 2353 | 28.385 | 1817 | 37.603 | 10.1271 | 43.2379 | 43.8877 | 4.3337 | 27.9424 | 24.3297
2014/03/16 18:35:36 | 37 | 39-UWSIF-UW Glut 220140020.24 | 39-UWSIF-UW Glut 2 | 0.7205 | 2342 | 28.378 | 1806 | 37.661 | 10.1503 | 43.2157 | 43.8835 | 4.3234 | 27.9354 | 24.3886
2014/03/16 22:41:48 | 66 | 39-UWSIF-UW Glut 220140020.25 | 39-UWSIF-UW Glut 2 | 0.7327 | 2309 | 28.361 | 1782 | 37.686 | 9.8719 | 42.1528 | 43.344 | 4.3906 | 27.9184 | 24.4139
2014/03/16 22:50:18 | 67 | 39-UWSIF-UW Glut 220140020.26 | 39-UWSIF-UW Glut 2 | 0.7273 | 2290 | 28.323 | 1772 | 37.701 | 9.8528 | 42.0967 | 43.306 | 4.3953 | 27.8804 | 24.4291
2014/03/17 02:56:45 | 96 | 39-UWSIF-UW Glut 220140020.27 | 39-UWSIF-UW Glut 2 | 0.714 | 2269 | 28.3 | 1761 | 37.545 | 9.9287 | 42.5804 | 44.3131 | 4.4632 | 27.8573 | 24.2709
2014/03/17 03:05:15 | 97 | 39-UWSIF-UW Glut 220140020.28 | 39-UWSIF-UW Glut 2 | 0.7284 | 2315 | 28.343 | 1799 | 37.613 | 9.9442 | 42.6414 | 44.3922 | 4.4641 | 27.9004 | 24.3399
[5 empty] | average | 28.3313 |  | 37.6365 | 10.0577 | 42.9879 | 43.9174 |  | 27.8886 | 24.3637
[5 empty] | std. dev. | 0.044487 |  | 0.056579 | 0.18155 | 0.74672 | 0.40753 |  | 0.044522 | 0.057388
15N Normilization
|  |  | Meas. | Actual
|  | UWSIF 36- UT Glutamic | -4.1475 | -4.616
|  | UWSIF 39- UW Glutamic 2 | 28.3313 | 27.888
Lab QC Check
|  | Ref. Check | Corrected | Actual | 0.0011343
|  | UWSIF01 (Liver) | 6.7789 | 6.78
Percentage
|  | UWSIF01 (Liver) | 12.6777 | 12.897
13C Normilization
|  |  | Meas. | Actual
|  | UWSIF 36- UT Glutamic | -14.2649 | -28.279
|  | UWSIF 39- UW Glutamic 2 | 37.6365 | 24.362
Lab QC Check
|  | Ref. Check | Corrected | Actual | 0.081884
|  | UWSIF01 (Liver) | -17.7381 | -17.82
Percentage
|  | UWSIF01 (Liver) | 49.6389 | 49.537
Date Analyzed:3/17/14
Analyst:DEW
[6 empty] | Instrument: | Costech EA 4010

## Run 2
Date | Line | Identifier 1 | Identifier 2 | Amount | Ampl 28 | d 15N/14N | Ampl 44 | d 13C/12C | Nitrogen % | Carbon % | C:N | Corr. d15N | Corr. d13C
2014/03/24 11:10:45 | 10 | 20140020.073 | 1261 | 0.828 | 2628 | 13.81 | 1878 | -15.037 | 11.3477 | 46.2578 | 4.0764 | 13.276 | -29.1234
2014/03/24 11:19:12 | 11 | 20140020.074 | 1235 | 0.848 | 2915 | 14.783 | 1930 | -14.782 | 12.3001 | 46.3065 | 3.7647 | 14.2474 | -28.8655
2014/03/24 11:27:40 | 12 | 20140020.075 | 1360 | 0.756 | 2819 | 15.53 | 1815 | -14.218 | 13.4433 | 49.0365 | 3.6477 | 14.9931 | -28.2949
2014/03/24 11:36:08 | 13 | 20140020.076 | 1340 | 0.846 | 2925 | 15.158 | 1896 | -14.535 | 12.3579 | 45.5469 | 3.6857 | 14.6217 | -28.6156
2014/03/24 11:44:36 | 14 | 20140020.077 | 1580 | 0.781 | 2363 | 13.89 | 1760 | -14.425 | 10.845 | 45.545 | 4.1996 | 13.3559 | -28.5043
2014/03/24 11:53:04 | 15 | 20140020.078 | 1540 | 0.811 | 2508 | 14.075 | 1983 | -15.281 | 11.0238 | 49.5289 | 4.4929 | 13.5406 | -29.3703
2014/03/24 12:01:31 | 16 | 20140020.079 | 1441 | 0.84 | 2867 | 14.653 | 2145 | -15.24 | 12.1027 | 51.5886 | 4.2626 | 14.1176 | -29.3288
2014/03/24 12:09:59 | 17 | 20140020.080 | 1477 | 0.77 | 2765 | 14.889 | 1776 | -14.725 | 12.8733 | 46.9847 | 3.6498 | 14.3532 | -28.8078
2014/03/24 12:18:27 | 18 | 20140020.081 | 1603 | 0.76 | 2675 | 14.675 | 1915 | -14.971 | 12.6284 | 51.3701 | 4.0678 | 14.1396 | -29.0567
2014/03/24 12:26:55 | 19 | 20140020.082 | 1604 | 0.809 | 2384 | 13.758 | 1949 | -14.859 | 10.149 | 49.0944 | 4.8374 | 13.2241 | -28.9434
2014/03/24 12:35:23 | 20 | 20140020.083 | 1472 | 0.794 | 2593 | 13.158 | 1707 | -13.312 | 11.7614 | 43.8819 | 3.731 | 12.6251 | -27.3784
Reference Check
2014/03/24 10:53:49 | 8 | 01-UWSIF-Liver 20140020.31 | 01-UWSIF-Liver | 0.68 | 2420 | 7.275 | 1695 | -3.752 | 12.8185 | 50.7453 | 3.9588 | 6.7521 | -17.7075
2014/03/24 11:02:17 | 9 | 01-UWSIF-Liver 20140020.32 | 01-UWSIF-Liver | 0.759 | 2723 | 7.337 | 1890 | -3.786 | 12.8041 | 50.4198 | 3.9378 | 6.814 | -17.7419
2014/03/24 13:17:44 | 25 | 01-UWSIF-Liver 20140020.33 | 01-UWSIF-Liver | 0.7543 | 2720 | 7.303 | 1893 | -3.763 | 12.8108 | 50.5277 | 3.9441 | 6.7801 | -17.7187
2014/03/24 13:26:12 | 26 | 01-UWSIF-Liver 20140020.34 | 01-UWSIF-Liver | 0.731 | 2630 | 7.241 | 1834 | -3.71 | 12.8228 | 50.5378 | 3.9412 | 6.7182 | -17.665
[5 empty] | average | 7.289 |  | -3.7527 | 12.814 | 50.5577 |  | 6.7661 | -17.7083
[5 empty] | std. dev. | 0.040825 |  | 0.031826 | 0.0082851 | 0.13602 |  | 0.040755 | 0.032196
Reference Material
2014/03/24 10:19:57 | 4 | 36-UWSIF-UT Glut 120140020.34 | 36-UWSIF-UT Glut 1 | 0.813 | 2140 | -4.151 | 1669 | -14.155 | 9.3728 | 41.3882 | 4.4158 | -4.6544 | -28.2312
2014/03/24 10:28:25 | 5 | 36-UWSIF-UT Glut 120140020.35 | 36-UWSIF-UT Glut 1 | 1.067 | 2835 | -4.055 | 2182 | -14.247 | 9.375 | 41.0831 | 4.3822 | -4.5586 | -28.3243
2014/03/24 12:43:51 | 21 | 36-UWSIF-UT Glut 120140020.36 | 36-UWSIF-UT Glut 1 | 0.788 | 2049 | -4.102 | 1603 | -14.173 | 9.3331 | 41.2854 | 4.4235 | -4.6055 | -28.2494
2014/03/24 12:52:20 | 22 | 36-UWSIF-UT Glut 120140020.37 | 36-UWSIF-UT Glut 1 | 0.814 | 2108 | -4.142 | 1653 | -14.234 | 9.361 | 41.3127 | 4.4133 | -4.6455 | -28.3111
[5 empty] | average | -4.1125 |  | -14.2023 | 9.3605 | 41.2674 |  | -4.616 | -28.279
[5 empty] | std. dev. | 0.043852 |  | 0.045088 | 0.019257 | 0.1303 |  | 0.043777 | 0.045611
2014/03/24 10:36:53 | 6 | 39-UWSIF-UW Glut 220140020.31 | 39-UWSIF-UW Glut 2 | 0.7874 | 2235 | 28.531 | 1739 | 37.875 | 10.1665 | 44.7878 | 4.4054 | 27.972 | 24.4024
2014/03/24 10:45:21 | 7 | 39-UWSIF-UW Glut 220140020.32 | 39-UWSIF-UW Glut 2 | 0.7924 | 2255 | 28.441 | 1754 | 37.894 | 10.1707 | 44.8455 | 4.4093 | 27.8822 | 24.4216
2014/03/24 13:00:48 | 23 | 39-UWSIF-UW Glut 220140020.33 | 39-UWSIF-UW Glut 2 | 0.78 | 2215 | 28.427 | 1729 | 37.674 | 10.2071 | 44.9183 | 4.4007 | 27.8682 | 24.199
2014/03/24 13:09:16 | 24 | 39-UWSIF-UW Glut 220140020.34 | 39-UWSIF-UW Glut 2 | 0.7282 | 2044 | 28.389 | 1610 | 37.899 | 10.1229 | 44.7772 | 4.4234 | 27.8302 | 24.4266
[5 empty] | average | 28.447 |  | 37.8355 | 10.1668 | 44.8322 |  | 27.8881 | 24.3624
[5 empty] | std. dev. | 0.060155 |  | 0.10816 | 0.03448 | 0.064773 |  | 0.060053 | 0.10942
15N Normilization
|  |  | Meas. | Actual
|  | UWSIF 36- UT Glutamic | -4.1125 | -4.616
|  | UWSIF 39- UW Glutamic 2 | 28.447 | 27.888
Lab QC Check
|  | Ref. Check | Corrected | Actual | 0.013891
|  | UWSIF01 (Liver) | 6.7661 | 6.78
Percentage
|  | UWSIF01 (Liver) | 12.814 | 12.897
13C Normilization
|  |  | Meas. | Actual
|  | UWSIF 36- UT Glutamic | -14.2023 | -28.279
|  | UWSIF 39- UW Glutamic 2 | 37.8355 | 24.362
Lab QC Check
|  | Ref. Check | Corrected | Actual | 0.11172
|  | UWSIF01 (Liver) | -17.7083 | -17.82
Percentage
|  | UWSIF01 (Liver) | 50.5577 | 49.537
Date Analyzed:3/24/14
Analyst:DEW
[6 empty] | Instrument: | Costech EA 4010

## Original 1
Time Code | Row | Identifier 1 | Identifier 2 | Comment | Amount | Peak Nr | Ampl 28 | d 15N/14N | Ampl 44 | d 13C/12C | Amt% | Area All | Area 28 | BGD 28 | BGD 29 | BGD 30 | Area 44 | BGD 44 | BGD 45 | BGD 46 | AT% 13C/12C | AT% 15N/14N | rR 45CO2/44CO2 | rR 29N2/28N2
2014/03/16 13:30:25 | 1 | 36-UWSIF-UT Glut 120140020.21 | 36-UWSIF-UT Glut 1 | Job: 2014-0020 Run 1 Stockwell | 0.714 | 1 | 3054 | -0.019 |  |  | 12.1531 | 60.291 | 59.831 | 17.0 | 13.7 | 175.7 | [5 empty] | 0.36646 |  | 0.72843
2014/03/16 13:30:25 | 1 | 36-UWSIF-UT Glut 120140020.21 | 36-UWSIF-UT Glut 1 | Job: 2014-0020 Run 1 Stockwell | 0.714 | 2 | 3050 | 0 |  |  | 12.1328 | 60.19 | 59.731 | 16.9 | 13.7 | 181.0 | [5 empty] | 0.36647 |  | 0.72845
2014/03/16 13:30:25 | 1 | 36-UWSIF-UT Glut 120140020.21 | 36-UWSIF-UT Glut 1 | Job: 2014-0020 Run 1 Stockwell | 0.714 | 3 | 2139 | -4.156 |  |  | 10.2236 | 50.719 | 50.332 | 16.8 | 13.7 | 167.2 | [5 empty] | 0.36495 |  | 0.72542
2014/03/16 13:30:25 | 1 | 36-UWSIF-UT Glut 120140020.21 | 36-UWSIF-UT Glut 1 | Job: 2014-0020 Run 1 Stockwell | 0.714 | 4 |  |  | 1661 | -14.172 | 46.2615 | 42.861 |  |  |  |  | 42.193 | 2.8 | 3.3 | 5.7 | 1.0902 |  | 1.1611
2014/03/16 13:30:25 | 1 | 36-UWSIF-UT Glut 120140020.21 | 36-UWSIF-UT Glut 1 | Job: 2014-0020 Run 1 Stockwell | 0.714 | 5 |  |  | 3398 | 0 | 72.5899 | 67.254 |  |  |  |  | 66.196 | 4.8 | 5.6 | 8.5 | 1.1057 |  | 1.1766
2014/03/16 13:38:54 | 2 | 36-UWSIF-UT Glut 120140020.22 | 36-UWSIF-UT Glut 1 | Analyst: DEw | 0.79 | 1 | 3046 | -0.017 |  |  |  | 60.118 | 59.66 | 16.6 | 13.5 | 167.2 | [5 empty] | 0.36647 |  | 0.72848
2014/03/16 13:38:54 | 2 | 36-UWSIF-UT Glut 120140020.22 | 36-UWSIF-UT Glut 1 | Analyst: DEw | 0.79 | 2 | 3039 | 0 |  |  |  | 60.01 | 59.553 | 16.5 | 13.4 | 174.6 | [5 empty] | 0.36647 |  | 0.72849
2014/03/16 13:38:54 | 2 | 36-UWSIF-UT Glut 120140020.22 | 36-UWSIF-UT Glut 1 | Analyst: DEw | 0.79 | 3 | 2385 | -4.172 |  |  | 9.52 | 55.951 | 55.524 | 16.5 | 13.3 | 162.2 | [5 empty] | 0.36495 |  | 0.72545
2014/03/16 13:38:54 | 2 | 36-UWSIF-UT Glut 120140020.22 | 36-UWSIF-UT Glut 1 | Analyst: DEw | 0.79 | 4 |  |  | 1843 | -14.072 | 40.81 | 47.158 |  |  |  |  | 46.423 | 2.4 | 2.9 | 5.3 | 1.0903 |  | 1.1612
2014/03/16 13:38:54 | 2 | 36-UWSIF-UT Glut 120140020.22 | 36-UWSIF-UT Glut 1 | Analyst: DEw | 0.79 | 5 |  |  | 3380 | 0 |  | 66.91 |  |  |  |  | 65.858 | 4.2 | 5.0 | 7.8 | 1.1057 |  | 1.1766
2014/03/16 13:47:21 | 3 | 36-UWSIF-UT Glut 120140020.23 | 36-UWSIF-UT Glut 1 | Sample Count:376 | 0.402 | 1 | 3020 | 0.003 |  |  | 19.9236 | 59.585 | 59.132 | 16.2 | 13.0 | 164.6 | [5 empty] | 0.36647 |  | 0.72858
2014/03/16 13:47:21 | 3 | 36-UWSIF-UT Glut 120140020.23 | 36-UWSIF-UT Glut 1 | Sample Count:376 | 0.402 | 2 | 3024 | 0 |  |  | 19.937 | 59.625 | 59.171 | 16.1 | 12.9 | 172.3 | [5 empty] | 0.36647 |  | 0.72857
2014/03/16 13:47:21 | 3 | 36-UWSIF-UT Glut 120140020.23 | 36-UWSIF-UT Glut 1 | Sample Count:376 | 0.402 | 3 | 1185 | -4.232 |  |  | 9.411 | 28.145 | 27.932 | 16.1 | 12.8 | 159.9 | [5 empty] | 0.36493 |  | 0.72549
2014/03/16 13:47:21 | 3 | 36-UWSIF-UT Glut 120140020.23 | 36-UWSIF-UT Glut 1 | Sample Count:376 | 0.402 | 4 |  |  | 930 | -13.904 | 39.987 | 23.513 |  |  |  |  | 23.146 | 2.2 | 2.6 | 4.9 | 1.0905 |  | 1.1615
2014/03/16 13:47:21 | 3 | 36-UWSIF-UT Glut 120140020.23 | 36-UWSIF-UT Glut 1 | Sample Count:376 | 0.402 | 5 |  |  | 3364 | 0 | 113.2295 | 66.58 |  |  |  |  | 65.533 | 3.4 | 4.0 | 6.6 | 1.1057 |  | 1.1767
2014/03/16 13:55:50 | 4 | 36-UWSIF-UT Glut 120140020.24 | 36-UWSIF-UT Glut 1 | Helium Pressure: 1200 | 0.764 | 1 | 3019 | 0.058 |  |  | 10.4784 | 59.557 | 59.104 | 15.8 | 12.6 | 161.9 | [5 empty] | 0.36649 |  | 0.7286
2014/03/16 13:55:50 | 4 | 36-UWSIF-UT Glut 120140020.24 | 36-UWSIF-UT Glut 1 | Helium Pressure: 1200 | 0.764 | 2 | 3019 | 0 |  |  | 10.4823 | 59.579 | 59.126 | 15.8 | 12.5 | 170.0 | [5 empty] | 0.36647 |  | 0.72856
2014/03/16 13:55:50 | 4 | 36-UWSIF-UT Glut 120140020.24 | 36-UWSIF-UT Glut 1 | Helium Pressure: 1200 | 0.764 | 3 | 2302 | -4.12 |  |  | 9.4962 | 53.974 | 53.563 | 15.8 | 12.5 | 159.3 | [5 empty] | 0.36497 |  | 0.72556
2014/03/16 13:55:50 | 4 | 36-UWSIF-UT Glut 120140020.24 | 36-UWSIF-UT Glut 1 | Helium Pressure: 1200 | 0.764 | 4 |  |  | 1786 | -14.122 | 40.703 | 45.486 |  |  |  |  | 44.777 | 2.0 | 2.4 | 4.7 | 1.0902 |  | 1.1613
2014/03/16 13:55:50 | 4 | 36-UWSIF-UT Glut 120140020.24 | 36-UWSIF-UT Glut 1 | Helium Pressure: 1200 | 0.764 | 5 |  |  | 3364 | 0 | 59.6017 | 66.606 |  |  |  |  | 65.558 | 3.6 | 4.3 | 7.0 | 1.1057 |  | 1.1766
2014/03/16 14:04:18 | 5 | 36-UWSIF-UT Glut 120140020.25 | 36-UWSIF-UT Glut 1 | Oxygen pressure: 2400 | 1.027 | 1 | 3028 | 0.034 |  |  | 7.8219 | 59.762 | 59.307 | 15.7 | 12.4 | 165.0 | [5 empty] | 0.36649 |  | 0.72867
2014/03/16 14:04:18 | 5 | 36-UWSIF-UT Glut 120140020.25 | 36-UWSIF-UT Glut 1 | Oxygen pressure: 2400 | 1.027 | 2 | 3027 | 0 |  |  | 7.8203 | 59.75 | 59.296 | 15.6 | 12.3 | 172.6 | [5 empty] | 0.36647 |  | 0.72864
2014/03/16 14:04:18 | 5 | 36-UWSIF-UT Glut 120140020.25 | 36-UWSIF-UT Glut 1 | Oxygen pressure: 2400 | 1.027 | 3 | 3138 | -4.106 |  |  | 9.5206 | 72.741 | 72.186 | 15.6 | 12.2 | 161.4 | [5 empty] | 0.36497 |  | 0.72565
2014/03/16 14:04:18 | 5 | 36-UWSIF-UT Glut 120140020.25 | 36-UWSIF-UT Glut 1 | Oxygen pressure: 2400 | 1.027 | 4 |  |  | 2390 | -14.182 | 40.7522 | 61.218 |  |  |  |  | 60.263 | 2.0 | 2.4 | 4.6 | 1.0901 |  | 1.1612
2014/03/16 14:04:18 | 5 | 36-UWSIF-UT Glut 120140020.25 | 36-UWSIF-UT Glut 1 | Oxygen pressure: 2400 | 1.027 | 5 |  |  | 3368 | 0 | 44.3807 | 66.669 |  |  |  |  | 65.62 | 3.8 | 4.5 | 7.2 | 1.1057 |  | 1.1766
2014/03/16 14:12:46 | 6 | 39-UWSIF-UW Glut 220140020.21 | 39-UWSIF-UW Glut 2 | Mass 28:17 | 0.7555 | 1 | 3033 | 0.012 |  |  | 10.6476 | 59.845 | 59.39 | 15.6 | 12.3 | 168.7 | [5 empty] | 0.36648 |  | 0.72868
2014/03/16 14:12:46 | 6 | 39-UWSIF-UW Glut 220140020.21 | 39-UWSIF-UW Glut 2 | Mass 28:17 | 0.7555 | 2 | 3029 | 0 |  |  | 10.6294 | 59.743 | 59.288 | 15.5 | 12.1 | 176.2 | [5 empty] | 0.36647 |  | 0.72868
2014/03/16 14:12:46 | 6 | 39-UWSIF-UW Glut 220140020.21 | 39-UWSIF-UW Glut 2 | Mass 28:17 | 0.7555 | 3 | 2471 | 28.253 |  |  | 10.306 | 57.925 | 57.471 | 15.5 | 12.1 | 164.7 | [5 empty] | 0.37679 |  | 0.74926
2014/03/16 14:12:46 | 6 | 39-UWSIF-UW Glut 220140020.21 | 39-UWSIF-UW Glut 2 | Mass 28:17 | 0.7555 | 4 |  |  | 1908 | 37.591 | 44.0456 | 48.674 |  |  |  |  | 47.888 | 1.9 | 2.3 | 4.4 | 1.1467 |  | 1.2184
2014/03/16 14:12:46 | 6 | 39-UWSIF-UW Glut 220140020.21 | 39-UWSIF-UW Glut 2 | Mass 28:17 | 0.7555 | 5 |  |  | 3358 | 0 | 60.1911 | 66.516 |  |  |  |  | 65.469 | 3.5 | 4.2 | 6.7 | 1.1057 |  | 1.1769
2014/03/16 14:21:15 | 7 | 39-UWSIF-UW Glut 220140020.22 | 39-UWSIF-UW Glut 2 | Mass 29:14 | 0.7701 | 1 | 3022 | 0.033 |  |  | 10.4161 | 59.675 | 59.221 | 15.3 | 12.0 | 169.0 | [5 empty] | 0.36648 |  | 0.72871
2014/03/16 14:21:15 | 7 | 39-UWSIF-UW Glut 220140020.22 | 39-UWSIF-UW Glut 2 | Mass 29:14 | 0.7701 | 2 | 3020 | 0 |  |  | 10.3925 | 59.54 | 59.087 | 15.2 | 11.9 | 176.5 | [5 empty] | 0.36647 |  | 0.72869
2014/03/16 14:21:15 | 7 | 39-UWSIF-UW Glut 220140020.22 | 39-UWSIF-UW Glut 2 | Mass 29:14 | 0.7701 | 3 | 2514 | 28.307 |  |  | 10.2805 | 58.898 | 58.436 | 15.3 | 11.9 | 164.6 | [5 empty] | 0.37681 |  | 0.74931
2014/03/16 14:21:15 | 7 | 39-UWSIF-UW Glut 220140020.22 | 39-UWSIF-UW Glut 2 | Mass 29:14 | 0.7701 | 4 |  |  | 1934 | 37.692 | 43.9325 | 49.487 |  |  |  |  | 48.688 | 1.8 | 2.2 | 4.3 | 1.1469 |  | 1.2186
2014/03/16 14:21:15 | 7 | 39-UWSIF-UW Glut 220140020.22 | 39-UWSIF-UW Glut 2 | Mass 29:14 | 0.7701 | 5 |  |  | 3360 | 0 | 59.0692 | 66.538 |  |  |  |  | 65.491 | 3.4 | 4.1 | 6.7 | 1.1057 |  | 1.177
2014/03/16 14:29:43 | 8 | 01-UWSIF-Liver 20140020.21 | 01-UWSIF-Liver | Mass 30: 129 | 0.7682 | 1 | 3024 | -0.006 |  |  | 10.4411 | 59.671 | 59.217 | 15.2 | 11.9 | 170.8 | [5 empty] | 0.36647 |  | 0.72868
2014/03/16 14:29:43 | 8 | 01-UWSIF-Liver 20140020.21 | 01-UWSIF-Liver | Mass 30: 129 | 0.7682 | 2 | 3025 | 0 |  |  | 10.4388 | 59.658 | 59.204 | 15.1 | 11.7 | 178.5 | [5 empty] | 0.36647 |  | 0.72868
2014/03/16 14:29:43 | 8 | 01-UWSIF-Liver 20140020.21 | 01-UWSIF-Liver | Mass 30: 129 | 0.7682 | 3 | 3203 | 7.291 |  |  | 13.0217 | 74.419 | 73.846 | 15.1 | 11.8 | 166.5 | [5 empty] | 0.36913 |  | 0.73399
2014/03/16 14:29:43 | 8 | 01-UWSIF-Liver 20140020.21 | 01-UWSIF-Liver | Mass 30: 129 | 0.7682 | 4 |  |  | 2195 | -3.832 | 50.0073 | 56.191 |  |  |  |  | 55.308 | 1.8 | 2.1 | 4.3 | 1.1015 |  | 1.1728
2014/03/16 14:29:43 | 8 | 01-UWSIF-Liver 20140020.21 | 01-UWSIF-Liver | Mass 30: 129 | 0.7682 | 5 |  |  | 3357 | 0 | 59.0982 | 66.406 |  |  |  |  | 65.361 | 3.5 | 4.2 | 6.7 | 1.1057 |  | 1.1768
2014/03/16 14:38:12 | 9 | 01-UWSIF-Liver 20140020.22 | 01-UWSIF-Liver | Peak Center:3.065 | 0.73 | 1 | 3017 | -0.045 |  |  | 10.9629 | 59.538 | 59.085 | 15.0 | 11.7 | 173.3 | [5 empty] | 0.36646 |  | 0.72874
2014/03/16 14:38:12 | 9 | 01-UWSIF-Liver 20140020.22 | 01-UWSIF-Liver | Peak Center:3.065 | 0.73 | 2 | 3016 | 0 |  |  | 10.9611 | 59.528 | 59.075 | 14.9 | 11.6 | 180.8 | [5 empty] | 0.36647 |  | 0.72877
2014/03/16 14:38:12 | 9 | 01-UWSIF-Liver 20140020.22 | 01-UWSIF-Liver | Peak Center:3.065 | 0.73 | 3 | 3034 | 7.164 |  |  | 12.9268 | 70.203 | 69.662 | 15.0 | 11.6 | 168.7 | [5 empty] | 0.36909 |  | 0.73399
2014/03/16 14:38:12 | 9 | 01-UWSIF-Liver 20140020.22 | 01-UWSIF-Liver | Peak Center:3.065 | 0.73 | 4 |  |  | 2085 | -3.89 | 49.6581 | 53.024 |  |  |  |  | 52.191 | 1.7 | 2.1 | 4.2 | 1.1014 |  | 1.1728
2014/03/16 14:38:12 | 9 | 01-UWSIF-Liver 20140020.22 | 01-UWSIF-Liver | Peak Center:3.065 | 0.73 | 5 |  |  | 3349 | 0 | 62.1178 | 66.328 |  |  |  |  | 65.285 | 3.4 | 4.1 | 6.6 | 1.1057 |  | 1.1768
2014/03/16 14:46:40 | 10 | 20140020.001 | 969 | Instrument: ANNIE | 0.793 | 1 | 3013 | 0.024 |  |  | 10.0788 | 59.46 | 59.007 | 14.9 | 11.6 | 175.4 | [5 empty] | 0.36648 |  | 0.72879
2014/03/16 14:46:40 | 10 | 20140020.001 | 969 | Instrument: ANNIE | 0.793 | 2 | 3013 | 0 |  |  | 10.0874 | 59.511 | 59.058 | 14.8 | 11.5 | 182.7 | [5 empty] | 0.36647 |  | 0.72877
2014/03/16 14:46:40 | 10 | 20140020.001 | 969 | Instrument: ANNIE | 0.793 | 3 | 1052 | 12.596 |  |  | 4.229 | 24.949 | 24.757 | 14.8 | 11.5 | 170.3 | [5 empty] | 0.37107 |  | 0.73795
2014/03/16 14:46:40 | 10 | 20140020.001 | 969 | Instrument: ANNIE | 0.793 | 4 |  |  | 711 | -13.462 | 15.423 | 17.89 |  |  |  |  | 17.61 | 1.6 | 2.0 | 4.1 | 1.0909 |  | 1.1623
2014/03/16 14:46:40 | 10 | 20140020.001 | 969 | Instrument: ANNIE | 0.793 | 5 |  |  | 3352 | 0 | 57.123 | 66.259 |  |  |  |  | 65.216 | 2.7 | 3.2 | 5.6 | 1.1057 |  | 1.1769
2014/03/16 14:55:08 | 11 | 20140020.002 | 925 | Data analyst: | 0.765 | 1 | 3013 | -0.026 |  |  | 10.4458 | 59.449 | 58.997 | 14.7 | 11.4 | 173.6 | [5 empty] | 0.36646 |  | 0.72876
2014/03/16 14:55:08 | 11 | 20140020.002 | 925 | Data analyst: | 0.765 | 2 | 3009 | 0 |  |  | 10.4351 | 59.389 | 58.937 | 14.6 | 11.3 | 181.0 | [5 empty] | 0.36647 |  | 0.72878
2014/03/16 14:55:08 | 11 | 20140020.002 | 925 | Data analyst: | 0.765 | 3 | 2617 | 13.825 |  |  | 10.732 | 61.078 | 60.605 | 14.6 | 11.3 | 169.8 | [5 empty] | 0.37152 |  | 0.73885
2014/03/16 14:55:08 | 11 | 20140020.002 | 925 | Data analyst: | 0.765 | 4 |  |  | 2178 | -15.07 | 50.0202 | 55.972 |  |  |  |  | 55.098 | 1.6 | 1.9 | 3.9 | 1.0892 |  | 1.1605
2014/03/16 14:55:08 | 11 | 20140020.002 | 925 | Data analyst: | 0.765 | 5 |  |  | 3341 | 0 | 59.1011 | 66.133 |  |  |  |  | 65.092 | 3.3 | 3.9 | 6.5 | 1.1057 |  | 1.1767
2014/03/16 15:03:36 | 12 | 20140020.003 | 886 | Comments: | 0.849 | 1 | 3006 | 0.044 |  |  | 9.3967 | 59.351 | 58.899 | 14.6 | 11.4 | 176.6 | [5 empty] | 0.36649 |  | 0.72873
2014/03/16 15:03:36 | 12 | 20140020.003 | 886 | Comments: | 0.849 | 2 | 3005 | 0 |  |  | 9.3926 | 59.325 | 58.874 | 14.5 | 11.3 | 183.8 | [5 empty] | 0.36647 |  | 0.7287
2014/03/16 15:03:36 | 12 | 20140020.003 | 886 | Comments: | 0.849 | 3 | 3206 | 14.776 |  |  | 11.7359 | 74.125 | 73.55 | 14.6 | 11.3 | 172.0 | [5 empty] | 0.37187 |  | 0.73947
2014/03/16 15:03:36 | 12 | 20140020.003 | 886 | Comments: | 0.849 | 4 |  |  | 2383 | -14.615 | 49.3042 | 61.228 |  |  |  |  | 60.272 | 1.6 | 1.9 | 4.0 | 1.0897 |  | 1.161
2014/03/16 15:03:36 | 12 | 20140020.003 | 886 | Comments: | 0.849 | 5 |  |  | 3336 | 0 | 53.1454 | 65.998 |  |  |  |  | 64.96 | 3.4 | 4.0 | 6.6 | 1.1057 |  | 1.1767
2014/03/16 15:12:05 | 13 | 20140020.004 | 865 |  | 0.839 | 1 | 3006 | 0.029 |  |  | 9.5064 | 59.336 | 58.885 | 14.6 | 11.3 | 179.7 | [5 empty] | 0.36648 |  | 0.72879
2014/03/16 15:12:05 | 13 | 20140020.004 | 865 |  | 0.839 | 2 | 3005 | 0 |  |  | 9.5061 | 59.334 | 58.883 | 14.5 | 11.2 | 187.0 | [5 empty] | 0.36647 |  | 0.72877
2014/03/16 15:12:05 | 13 | 20140020.004 | 865 |  | 0.839 | 3 | 3515 | 14.398 |  |  | 13.0408 | 81.397 | 80.765 | 14.5 | 11.2 | 174.9 | [5 empty] | 0.37173 |  | 0.73926
2014/03/16 15:12:05 | 13 | 20140020.004 | 865 |  | 0.839 | 4 |  |  | 2213 | -14.256 | 46.1608 | 56.649 |  |  |  |  | 55.764 | 1.6 | 1.9 | 4.1 | 1.0901 |  | 1.1614
2014/03/16 15:12:05 | 13 | 20140020.004 | 865 |  | 0.839 | 5 |  |  | 3335 | 0 | 53.7943 | 66.017 |  |  |  |  | 64.979 | 3.3 | 3.9 | 6.5 | 1.1057 |  | 1.1767
2014/03/16 15:20:33 | 14 | 20140020.005 | 716 |  | 0.823 | 1 | 2996 | -0.021 |  |  | 9.6603 | 59.147 | 58.697 | 14.5 | 11.2 | 181.7 | [5 empty] | 0.36646 |  | 0.72872
2014/03/16 15:20:33 | 14 | 20140020.005 | 716 |  | 0.823 | 2 | 2999 | 0 |  |  | 9.657 | 59.127 | 58.677 | 14.4 | 11.1 | 188.8 | [5 empty] | 0.36647 |  | 0.72874
2014/03/16 15:20:33 | 14 | 20140020.005 | 716 |  | 0.823 | 3 | 3082 | 12.586 |  |  | 11.6694 | 71.448 | 70.895 | 14.4 | 11.1 | 177.2 | [5 empty] | 0.37107 |  | 0.73791
2014/03/16 15:20:33 | 14 | 20140020.005 | 716 |  | 0.823 | 4 |  |  | 2041 | -14.229 | 43.334 | 52.166 |  |  |  |  | 51.351 | 1.5 | 1.8 | 3.9 | 1.0901 |  | 1.1615
2014/03/16 15:20:33 | 14 | 20140020.005 | 716 |  | 0.823 | 5 |  |  | 3329 | 0 | 54.7446 | 65.902 |  |  |  |  | 64.865 | 3.2 | 3.7 | 6.3 | 1.1057 |  | 1.1767
2014/03/16 15:29:02 | 15 | 20140020.006 | 793 |  | 0.852 | 1 | 2993 | 0.005 |  |  | 9.3257 | 59.11 | 58.661 | 14.4 | 11.1 | 183.0 | [5 empty] | 0.36647 |  | 0.72877
2014/03/16 15:29:02 | 15 | 20140020.006 | 793 |  | 0.852 | 2 | 2993 | 0 |  |  | 9.3164 | 59.052 | 58.602 | 14.3 | 11.0 | 189.9 | [5 empty] | 0.36647 |  | 0.72877
2014/03/16 15:29:02 | 15 | 20140020.006 | 793 |  | 0.852 | 3 | 2551 | 14.722 |  |  | 9.389 | 59.512 | 59.05 | 14.3 | 11.0 | 178.2 | [5 empty] | 0.37185 |  | 0.7395
2014/03/16 15:29:02 | 15 | 20140020.006 | 793 |  | 0.852 | 4 |  |  | 1664 | -14.74 | 33.9454 | 42.304 |  |  |  |  | 41.643 | 1.5 | 1.8 | 3.9 | 1.0895 |  | 1.161
2014/03/16 15:29:02 | 15 | 20140020.006 | 793 |  | 0.852 | 5 |  |  | 3324 | 0 | 52.7688 | 65.762 |  |  |  |  | 64.727 | 2.9 | 3.5 | 5.9 | 1.1057 |  | 1.1768
2014/03/16 15:37:30 | 16 | 20140020.007 | 644 |  | 0.811 | 1 | 2987 | 0.019 |  |  | 9.7686 | 58.938 | 58.49 | 14.2 | 11.0 | 183.1 | [5 empty] | 0.36648 |  | 0.72876
2014/03/16 15:37:30 | 16 | 20140020.007 | 644 |  | 0.811 | 2 | 2984 | 0 |  |  | 9.7578 | 58.873 | 58.425 | 14.1 | 10.9 | 190.2 | [5 empty] | 0.36647 |  | 0.72874
2014/03/16 15:37:30 | 16 | 20140020.007 | 644 |  | 0.811 | 3 | 2810 | 14.541 |  |  | 10.8188 | 65.274 | 64.768 | 14.2 | 10.9 | 178.2 | [5 empty] | 0.37178 |  | 0.73934
2014/03/16 15:37:30 | 16 | 20140020.007 | 644 |  | 0.811 | 4 |  |  | 2254 | -15.404 | 48.807 | 57.898 |  |  |  |  | 56.994 | 1.4 | 1.7 | 3.8 | 1.0888 |  | 1.1603
2014/03/16 15:37:30 | 16 | 20140020.007 | 644 |  | 0.811 | 5 |  |  | 3315 | 0 | 55.309 | 65.611 |  |  |  |  | 64.579 | 3.1 | 3.6 | 6.3 | 1.1057 |  | 1.1767
2014/03/16 15:45:59 | 17 | 20140020.008 | 689 |  | 0.782 | 1 | 2983 | 0.034 |  |  | 10.1195 | 58.872 | 58.425 | 14.2 | 10.9 | 184.7 | [5 empty] | 0.36649 |  | 0.72879
2014/03/16 15:45:59 | 17 | 20140020.008 | 689 |  | 0.782 | 2 | 2986 | 0 |  |  | 10.1255 | 58.907 | 58.459 | 14.1 | 10.8 | 192.0 | [5 empty] | 0.36647 |  | 0.72877
2014/03/16 15:45:59 | 17 | 20140020.008 | 689 |  | 0.782 | 3 | 3347 | 14.121 |  |  | 13.3781 | 77.829 | 77.225 | 14.2 | 10.9 | 180.6 | [5 empty] | 0.37163 |  | 0.73906
2014/03/16 15:45:59 | 17 | 20140020.008 | 689 |  | 0.782 | 4 |  |  | 2140 | -14.392 | 47.9328 | 54.828 |  |  |  |  | 53.971 | 1.4 | 1.7 | 3.8 | 1.0899 |  | 1.1613
2014/03/16 15:45:59 | 17 | 20140020.008 | 689 |  | 0.782 | 5 |  |  | 3314 | 0 | 57.3133 | 65.557 |  |  |  |  | 64.526 | 3.0 | 3.6 | 6.2 | 1.1057 |  | 1.1767
2014/03/16 15:54:27 | 18 | 20140020.009 | 1009 |  | 0.785 | 1 | 2981 | -0.016 |  |  | 10.0764 | 58.846 | 58.399 | 14.1 | 10.9 | 187.1 | [5 empty] | 0.36647 |  | 0.72879
2014/03/16 15:54:27 | 18 | 20140020.009 | 1009 |  | 0.785 | 2 | 2977 | 0 |  |  | 10.0497 | 58.69 | 58.244 | 14.0 | 10.8 | 193.7 | [5 empty] | 0.36647 |  | 0.72881
2014/03/16 15:54:27 | 18 | 20140020.009 | 1009 |  | 0.785 | 3 | 2949 | 14.901 |  |  | 11.7413 | 68.569 | 68.037 | 14.1 | 10.8 | 181.9 | [5 empty] | 0.37191 |  | 0.73967
2014/03/16 15:54:27 | 18 | 20140020.009 | 1009 |  | 0.785 | 4 |  |  | 2074 | -14.123 | 46.2882 | 53.15 |  |  |  |  | 52.319 | 1.4 | 1.6 | 3.8 | 1.0902 |  | 1.1617
2014/03/16 15:54:27 | 18 | 20140020.009 | 1009 |  | 0.785 | 5 |  |  | 3311 | 0 | 57.0164 | 65.468 |  |  |  |  | 64.438 | 2.9 | 3.5 | 6.0 | 1.1057 |  | 1.1767
2014/03/16 16:02:56 | 19 | 20140020.010 | 1005 |  | 0.794 | 1 | 2972 | 0.011 |  |  | 9.9339 | 58.679 | 58.233 | 14.0 | 10.8 | 187.5 | [5 empty] | 0.36648 |  | 0.7288
2014/03/16 16:02:56 | 19 | 20140020.010 | 1005 |  | 0.794 | 2 | 2968 | 0 |  |  | 9.9134 | 58.558 | 58.113 | 13.9 | 10.7 | 193.9 | [5 empty] | 0.36647 |  | 0.7288
2014/03/16 16:02:56 | 19 | 20140020.010 | 1005 |  | 0.794 | 3 | 2626 | 13.341 |  |  | 10.2589 | 60.599 | 60.13 | 14.0 | 10.7 | 182.3 | [5 empty] | 0.37134 |  | 0.73852
2014/03/16 16:02:56 | 19 | 20140020.010 | 1005 |  | 0.794 | 4 |  |  | 2292 | -15.981 | 50.4369 | 58.577 |  |  |  |  | 57.662 | 1.3 | 1.6 | 3.7 | 1.0882 |  | 1.1596
2014/03/16 16:02:56 | 19 | 20140020.010 | 1005 |  | 0.794 | 5 |  |  | 3300 | 0 | 56.235 | 65.311 |  |  |  |  | 64.283 | 3.0 | 3.5 | 6.0 | 1.1057 |  | 1.1767
2014/03/16 16:11:24 | 20 | 20140020.011 | 1119 |  | 0.81 | 1 | 2966 | 0.015 |  |  | 9.7258 | 58.608 | 58.162 | 13.9 | 10.8 | 188.1 | [5 empty] | 0.36648 |  | 0.72881
2014/03/16 16:11:24 | 20 | 20140020.011 | 1119 |  | 0.81 | 2 | 2967 | 0 |  |  | 9.714 | 58.537 | 58.091 | 13.8 | 10.7 | 195.4 | [5 empty] | 0.36647 |  | 0.72879
2014/03/16 16:11:24 | 20 | 20140020.011 | 1119 |  | 0.81 | 3 | 2658 | 13.306 |  |  | 10.238 | 61.694 | 61.216 | 13.9 | 10.7 | 183.4 | [5 empty] | 0.37133 |  | 0.73849
2014/03/16 16:11:24 | 20 | 20140020.011 | 1119 |  | 0.81 | 4 |  |  | 2170 | -16.557 | 46.8092 | 55.46 |  |  |  |  | 54.594 | 1.3 | 1.6 | 3.7 | 1.0876 |  | 1.1591
2014/03/16 16:11:24 | 20 | 20140020.011 | 1119 |  | 0.81 | 5 |  |  | 3289 | 0 | 54.9139 | 65.062 |  |  |  |  | 64.038 | 2.9 | 3.4 | 5.9 | 1.1057 |  | 1.1767
2014/03/16 16:19:53 | 21 | 20140020.012 | 1175 |  | 0.839 | 1 | 2959 | 0.02 |  |  | 9.3538 | 58.384 | 57.94 | 13.8 | 10.7 | 188.6 | [5 empty] | 0.36648 |  | 0.72882
2014/03/16 16:19:53 | 21 | 20140020.012 | 1175 |  | 0.839 | 2 | 2954 | 0 |  |  | 9.3402 | 58.299 | 57.856 | 13.7 | 10.6 | 195.3 | [5 empty] | 0.36647 |  | 0.72881
2014/03/16 16:19:53 | 21 | 20140020.012 | 1175 |  | 0.839 | 3 | 2887 | 13.78 |  |  | 10.6858 | 66.698 | 66.181 | 13.8 | 10.6 | 183.6 | [5 empty] | 0.3715 |  | 0.73885
2014/03/16 16:19:53 | 21 | 20140020.012 | 1175 |  | 0.839 | 4 |  |  | 2242 | -14.607 | 46.663 | 57.266 |  |  |  |  | 56.37 | 1.3 | 1.5 | 3.6 | 1.0897 |  | 1.1612
2014/03/16 16:19:53 | 21 | 20140020.012 | 1175 |  | 0.839 | 5 |  |  | 3285 | 0 | 53.0154 | 65.062 |  |  |  |  | 64.038 | 2.9 | 3.4 | 5.9 | 1.1057 |  | 1.1767
2014/03/16 16:28:21 | 22 | 20140020.013 | 1202 |  | 0.804 | 1 | 2960 | -0.027 |  |  | 9.7653 | 58.41 | 57.966 | 13.8 | 10.7 | 190.5 | [5 empty] | 0.36646 |  | 0.72881
2014/03/16 16:28:21 | 22 | 20140020.013 | 1202 |  | 0.804 | 2 | 2962 | 0 |  |  | 9.772 | 58.45 | 58.005 | 13.7 | 10.6 | 197.5 | [5 empty] | 0.36647 |  | 0.72883
2014/03/16 16:28:21 | 22 | 20140020.013 | 1202 |  | 0.804 | 3 | 3093 | 14.691 |  |  | 11.9943 | 71.742 | 71.185 | 13.8 | 10.6 | 185.7 | [5 empty] | 0.37184 |  | 0.73954
2014/03/16 16:28:21 | 22 | 20140020.013 | 1202 |  | 0.804 | 4 |  |  | 2068 | -14.745 | 44.9728 | 52.889 |  |  |  |  | 52.062 | 1.2 | 1.5 | 3.5 | 1.0895 |  | 1.1611
2014/03/16 16:28:21 | 22 | 20140020.013 | 1202 |  | 0.804 | 5 |  |  | 3296 | 0 | 55.4273 | 65.184 |  |  |  |  | 64.158 | 2.8 | 3.3 | 5.8 | 1.1057 |  | 1.1767
2014/03/16 16:36:50 | 23 | 20140020.014 | 1258 |  | 0.813 | 1 | 2963 | -0.032 |  |  | 9.6723 | 58.501 | 58.056 | 13.8 | 10.7 | 192.5 | [5 empty] | 0.36646 |  | 0.72883
2014/03/16 16:36:50 | 23 | 20140020.014 | 1258 |  | 0.813 | 2 | 2961 | 0 |  |  | 9.6619 | 58.438 | 57.994 | 13.7 | 10.6 | 199.5 | [5 empty] | 0.36647 |  | 0.72885
2014/03/16 16:36:50 | 23 | 20140020.014 | 1258 |  | 0.813 | 3 | 2817 | 14.806 |  |  | 10.8087 | 65.374 | 64.867 | 13.8 | 10.7 | 187.3 | [5 empty] | 0.37188 |  | 0.73964
2014/03/16 16:36:50 | 23 | 20140020.014 | 1258 |  | 0.813 | 4 |  |  | 1830 | -14.111 | 39.075 | 46.468 |  |  |  |  | 45.741 | 1.2 | 1.5 | 3.5 | 1.0902 |  | 1.1618
2014/03/16 16:36:50 | 23 | 20140020.014 | 1258 |  | 0.813 | 5 |  |  | 3292 | 0 | 54.7658 | 65.127 |  |  |  |  | 64.102 | 2.7 | 3.2 | 5.6 | 1.1057 |  | 1.1767
2014/03/16 16:45:19 | 24 | 20140020.015 | 1332 |  | 0.776 | 1 | 2964 | 0.027 |  |  | 10.1306 | 58.484 | 58.039 | 13.7 | 10.7 | 193.4 | [5 empty] | 0.36648 |  | 0.72882
2014/03/16 16:45:19 | 24 | 20140020.015 | 1332 |  | 0.776 | 2 | 2960 | 0 |  |  | 10.1206 | 58.427 | 57.982 | 13.7 | 10.6 | 200.2 | [5 empty] | 0.36647 |  | 0.7288
2014/03/16 16:45:19 | 24 | 20140020.015 | 1332 |  | 0.776 | 3 | 1672 | 15.249 |  |  | 6.7784 | 39.132 | 38.829 | 13.7 | 10.6 | 187.9 | [5 empty] | 0.37204 |  | 0.73991
2014/03/16 16:45:19 | 24 | 20140020.015 | 1332 |  | 0.776 | 4 |  |  | 1129 | -14.352 | 24.9239 | 28.29 |  |  |  |  | 27.848 | 1.2 | 1.4 | 3.5 | 1.09 |  | 1.1616
2014/03/16 16:45:19 | 24 | 20140020.015 | 1332 |  | 0.776 | 5 |  |  | 3285 | 0 | 57.2549 | 64.988 |  |  |  |  | 63.965 | 2.3 | 2.8 | 5.1 | 1.1057 |  | 1.1768
2014/03/16 16:53:47 | 25 | 20140020.016 | 1395 |  | 0.802 | 1 | 2957 | 0.002 |  |  | 9.7802 | 58.353 | 57.91 | 13.6 | 10.6 | 191.7 | [5 empty] | 0.36647 |  | 0.72882
2014/03/16 16:53:47 | 25 | 20140020.016 | 1395 |  | 0.802 | 2 | 2956 | 0 |  |  | 9.7727 | 58.308 | 57.865 | 13.6 | 10.5 | 199.0 | [5 empty] | 0.36647 |  | 0.72881
2014/03/16 16:53:47 | 25 | 20140020.016 | 1395 |  | 0.802 | 3 | 2942 | 14.786 |  |  | 11.3993 | 68.013 | 67.486 | 13.6 | 10.5 | 187.2 | [5 empty] | 0.37187 |  | 0.73959
2014/03/16 16:53:47 | 25 | 20140020.016 | 1395 |  | 0.802 | 4 |  |  | 2189 | -14.661 | 47.6215 | 55.865 |  |  |  |  | 54.991 | 1.1 | 1.3 | 3.4 | 1.0896 |  | 1.1612
2014/03/16 16:53:47 | 25 | 20140020.016 | 1395 |  | 0.802 | 5 |  |  | 3282 | 0 | 55.371 | 64.956 |  |  |  |  | 63.933 | 2.7 | 3.2 | 5.8 | 1.1057 |  | 1.1768
2014/03/16 17:02:16 | 26 | 20140020.017 | 1577 |  | 0.796 | 1 | 2960 | 0.014 |  |  | 9.8651 | 58.419 | 57.975 | 13.7 | 10.6 | 194.9 | [5 empty] | 0.36648 |  | 0.72883
2014/03/16 17:02:16 | 26 | 20140020.017 | 1577 |  | 0.796 | 2 | 2958 | 0 |  |  | 9.8554 | 58.362 | 57.918 | 13.6 | 10.5 | 202.2 | [5 empty] | 0.36647 |  | 0.72882
2014/03/16 17:02:16 | 26 | 20140020.017 | 1577 |  | 0.796 | 3 | 2775 | 13.166 |  |  | 10.8432 | 64.212 | 63.714 | 13.7 | 10.5 | 190.1 | [5 empty] | 0.37128 |  | 0.73841
2014/03/16 17:02:16 | 26 | 20140020.017 | 1577 |  | 0.796 | 4 |  |  | 1990 | -14.144 | 43.4215 | 50.557 |  |  |  |  | 49.766 | 1.1 | 1.4 | 3.4 | 1.0902 |  | 1.1618
2014/03/16 17:02:16 | 26 | 20140020.017 | 1577 |  | 0.796 | 5 |  |  | 3290 | 0 | 55.9407 | 65.133 |  |  |  |  | 64.108 | 2.7 | 3.2 | 5.6 | 1.1057 |  | 1.1768
2014/03/16 17:10:45 | 27 | 20140020.018 | 1502 |  | 0.775 | 1 | 2957 | 0.016 |  |  | 10.1268 | 58.387 | 57.943 | 13.6 | 10.6 | 196.2 | [5 empty] | 0.36648 |  | 0.7288
2014/03/16 17:10:45 | 27 | 20140020.018 | 1502 |  | 0.775 | 2 | 2957 | 0 |  |  | 10.119 | 58.342 | 57.899 | 13.6 | 10.5 | 203.1 | [5 empty] | 0.36647 |  | 0.72879
2014/03/16 17:10:45 | 27 | 20140020.018 | 1502 |  | 0.775 | 3 | 2829 | 13.668 |  |  | 11.3977 | 65.715 | 65.205 | 13.7 | 10.5 | 191.8 | [5 empty] | 0.37146 |  | 0.73875
2014/03/16 17:10:45 | 27 | 20140020.018 | 1502 |  | 0.775 | 4 |  |  | 1842 | -13.571 | 41.1585 | 46.658 |  |  |  |  | 45.927 | 1.1 | 1.4 | 3.4 | 1.0908 |  | 1.1624
2014/03/16 17:10:45 | 27 | 20140020.018 | 1502 |  | 0.775 | 5 |  |  | 3294 | 0 | 57.4916 | 65.173 |  |  |  |  | 64.147 | 2.6 | 3.1 | 5.5 | 1.1057 |  | 1.1768
2014/03/16 17:19:13 | 28 | 20140020.019 | 1456 |  | 0.741 | 1 | 2967 | 0.063 |  |  | 10.6249 | 58.572 | 58.126 | 13.7 | 10.6 | 199.1 | [5 empty] | 0.3665 |  | 0.72886
2014/03/16 17:19:13 | 28 | 20140020.019 | 1456 |  | 0.741 | 2 | 2962 | 0 |  |  | 10.6054 | 58.464 | 58.019 | 13.6 | 10.5 | 205.7 | [5 empty] | 0.36647 |  | 0.72881
2014/03/16 17:19:13 | 28 | 20140020.019 | 1456 |  | 0.741 | 3 | 2652 | 13.158 |  |  | 11.2151 | 61.825 | 61.347 | 13.7 | 10.5 | 194.2 | [5 empty] | 0.37128 |  | 0.7384
2014/03/16 17:19:13 | 28 | 20140020.019 | 1456 |  | 0.741 | 4 |  |  | 1717 | -13.617 | 40.2634 | 43.64 |  |  |  |  | 42.957 | 1.1 | 1.4 | 3.4 | 1.0908 |  | 1.1625
2014/03/16 17:19:13 | 28 | 20140020.019 | 1456 |  | 0.741 | 5 |  |  | 3291 | 0 | 60.0577 | 65.095 |  |  |  |  | 64.071 | 2.5 | 2.9 | 5.4 | 1.1057 |  | 1.1768
2014/03/16 17:27:42 | 29 | 20140020.020 | 1423 |  | 0.638 | 1 | 2962 | 0.026 |  |  | 12.3212 | 58.481 | 58.037 | 13.6 | 10.5 | 199.4 | [5 empty] | 0.36648 |  | 0.72884
2014/03/16 17:27:42 | 29 | 20140020.020 | 1423 |  | 0.638 | 2 | 2958 | 0 |  |  | 12.2945 | 58.355 | 57.911 | 13.5 | 10.4 | 206.1 | [5 empty] | 0.36647 |  | 0.72882
2014/03/16 17:27:42 | 29 | 20140020.020 | 1423 |  | 0.638 | 3 | 2344 | 14.25 |  |  | 11.4949 | 54.559 | 54.136 | 13.6 | 10.5 | 193.8 | [5 empty] | 0.37167 |  | 0.7392
2014/03/16 17:27:42 | 29 | 20140020.020 | 1423 |  | 0.638 | 4 |  |  | 1673 | -14.795 | 45.3849 | 42.354 |  |  |  |  | 41.691 | 1.1 | 1.3 | 3.4 | 1.0895 |  | 1.1611
2014/03/16 17:27:42 | 29 | 20140020.020 | 1423 |  | 0.638 | 5 |  |  | 3292 | 0 | 69.8263 | 65.163 |  |  |  |  | 64.137 | 2.5 | 2.9 | 5.4 | 1.1057 |  | 1.1768
2014/03/16 17:36:11 | 30 | 20140020.021 | 203 |  | 0.761 | 1 | 2961 | 0.008 |  |  | 10.3268 | 58.465 | 58.02 | 13.6 | 10.5 | 199.8 | [5 empty] | 0.36647 |  | 0.72882
2014/03/16 17:36:11 | 30 | 20140020.021 | 203 |  | 0.761 | 2 | 2959 | 0 |  |  | 10.3209 | 58.431 | 57.987 | 13.5 | 10.4 | 206.7 | [5 empty] | 0.36647 |  | 0.72882
2014/03/16 17:36:11 | 30 | 20140020.021 | 203 |  | 0.761 | 3 | 2405 | 15.073 |  |  | 9.8089 | 55.532 | 55.102 | 13.6 | 10.5 | 194.5 | [5 empty] | 0.37198 |  | 0.7398
2014/03/16 17:36:11 | 30 | 20140020.021 | 203 |  | 0.761 | 4 |  |  | 2211 | -17.035 | 50.5374 | 56.255 |  |  |  |  | 55.376 | 1.1 | 1.3 | 3.3 | 1.087 |  | 1.1587
2014/03/16 17:36:11 | 30 | 20140020.021 | 203 |  | 0.761 | 5 |  |  | 3289 | 0 | 58.5048 | 65.123 |  |  |  |  | 64.099 | 2.7 | 3.2 | 5.6 | 1.1057 |  | 1.1768
2014/03/16 17:44:41 | 31 | 20140020.022 | 133 |  | 0.76 | 1 | 2959 | 0.009 |  |  | 10.3279 | 58.394 | 57.95 | 13.6 | 10.5 | 200.8 | [5 empty] | 0.36647 |  | 0.72884
2014/03/16 17:44:41 | 31 | 20140020.022 | 133 |  | 0.76 | 2 | 2960 | 0 |  |  | 10.3372 | 58.447 | 58.002 | 13.5 | 10.4 | 207.7 | [5 empty] | 0.36647 |  | 0.72883
2014/03/16 17:44:41 | 31 | 20140020.022 | 133 |  | 0.76 | 3 | 2601 | 13.121 |  |  | 10.6017 | 59.942 | 59.478 | 13.6 | 10.5 | 195.9 | [5 empty] | 0.37126 |  | 0.73839
2014/03/16 17:44:41 | 31 | 20140020.022 | 133 |  | 0.76 | 4 |  |  | 2025 | -14.78 | 45.8714 | 50.994 |  |  |  |  | 50.196 | 1.1 | 1.3 | 3.4 | 1.0895 |  | 1.1611
2014/03/16 17:44:41 | 31 | 20140020.022 | 133 |  | 0.76 | 5 |  |  | 3294 | 0 | 58.56 | 65.099 |  |  |  |  | 64.075 | 2.6 | 3.1 | 5.5 | 1.1057 |  | 1.1768
2014/03/16 17:53:10 | 32 | 20140020.023 | 246 |  | 0.816 | 1 | 2961 | 0.014 |  |  | 9.6334 | 58.481 | 58.036 | 13.6 | 10.5 | 202.1 | [5 empty] | 0.36648 |  | 0.72881
2014/03/16 17:53:10 | 32 | 20140020.023 | 246 |  | 0.816 | 2 | 2961 | 0 |  |  | 9.6208 | 58.404 | 57.96 | 13.5 | 10.4 | 209.1 | [5 empty] | 0.36647 |  | 0.7288
2014/03/16 17:53:10 | 32 | 20140020.023 | 246 |  | 0.816 | 3 | 2394 | 13.372 |  |  | 9.0853 | 55.154 | 54.727 | 13.6 | 10.4 | 196.4 | [5 empty] | 0.37135 |  | 0.73854
2014/03/16 17:53:10 | 32 | 20140020.023 | 246 |  | 0.816 | 4 |  |  | 2390 | -16.148 | 51.0534 | 60.936 |  |  |  |  | 59.984 | 1.1 | 1.3 | 3.3 | 1.088 |  | 1.1596
2014/03/16 17:53:10 | 32 | 20140020.023 | 246 |  | 0.816 | 5 |  |  | 3290 | 0 | 54.5562 | 65.117 |  |  |  |  | 64.092 | 2.7 | 3.2 | 5.7 | 1.1057 |  | 1.1768
2014/03/16 18:01:39 | 33 | 20140020.024 | 285 |  | 0.806 | 1 | 2961 | 0.011 |  |  | 9.7377 | 58.39 | 57.945 | 13.6 | 10.5 | 202.9 | [5 empty] | 0.36648 |  | 0.72883
2014/03/16 18:01:39 | 33 | 20140020.024 | 285 |  | 0.806 | 2 | 2957 | 0 |  |  | 9.7147 | 58.252 | 57.809 | 13.5 | 10.4 | 209.4 | [5 empty] | 0.36647 |  | 0.72882
2014/03/16 18:01:39 | 33 | 20140020.024 | 285 |  | 0.806 | 3 | 2788 | 12.777 |  |  | 10.7856 | 64.673 | 64.172 | 13.6 | 10.4 | 197.8 | [5 empty] | 0.37114 |  | 0.73814
2014/03/16 18:01:39 | 33 | 20140020.024 | 285 |  | 0.806 | 4 |  |  | 2187 | -14.824 | 47.441 | 55.931 |  |  |  |  | 55.056 | 1.1 | 1.3 | 3.4 | 1.0894 |  | 1.1611
2014/03/16 18:01:39 | 33 | 20140020.024 | 285 |  | 0.806 | 5 |  |  | 3285 | 0 | 55.1581 | 65.029 |  |  |  |  | 64.005 | 2.6 | 3.1 | 5.6 | 1.1057 |  | 1.1768
2014/03/16 18:10:09 | 34 | 36-UWSIF-UT Glut 120140020.26 | 36-UWSIF-UT Glut 1 |  | 0.7673 | 1 | 2960 | 0.021 |  |  | 10.2385 | 58.445 | 58 | 13.6 | 10.5 | 204.2 | [5 empty] | 0.36648 |  | 0.72885
2014/03/16 18:10:09 | 34 | 36-UWSIF-UT Glut 120140020.26 | 36-UWSIF-UT Glut 1 |  | 0.7673 | 2 | 2960 | 0 |  |  | 10.2254 | 58.37 | 57.926 | 13.5 | 10.4 | 211.0 | [5 empty] | 0.36647 |  | 0.72884
2014/03/16 18:10:09 | 34 | 36-UWSIF-UT Glut 120140020.26 | 36-UWSIF-UT Glut 1 |  | 0.7673 | 3 | 2309 | -4.044 |  |  | 9.3482 | 53.363 | 52.956 | 13.6 | 10.4 | 198.7 | [5 empty] | 0.365 |  | 0.72589
2014/03/16 18:10:09 | 34 | 36-UWSIF-UT Glut 120140020.26 | 36-UWSIF-UT Glut 1 |  | 0.7673 | 4 |  |  | 1784 | -14.288 | 39.9031 | 44.785 |  |  |  |  | 44.084 | 1.1 | 1.3 | 3.3 | 1.09 |  | 1.1617
2014/03/16 18:10:09 | 34 | 36-UWSIF-UT Glut 120140020.26 | 36-UWSIF-UT Glut 1 |  | 0.7673 | 5 |  |  | 3297 | 0 | 58.0855 | 65.192 |  |  |  |  | 64.166 | 2.5 | 3.0 | 5.4 | 1.1057 |  | 1.1768
2014/03/16 18:18:38 | 35 | 36-UWSIF-UT Glut 120140020.27 | 36-UWSIF-UT Glut 1 |  | 0.757 | 1 | 2965 | -0.02 |  |  | 10.3946 | 58.539 | 58.094 | 13.6 | 10.4 | 205.3 | [5 empty] | 0.36646 |  | 0.7288
2014/03/16 18:18:38 | 35 | 36-UWSIF-UT Glut 120140020.27 | 36-UWSIF-UT Glut 1 |  | 0.757 | 2 | 2966 | 0 |  |  | 10.4026 | 58.584 | 58.138 | 13.5 | 10.4 | 212.5 | [5 empty] | 0.36647 |  | 0.72882
2014/03/16 18:18:38 | 35 | 36-UWSIF-UT Glut 120140020.27 | 36-UWSIF-UT Glut 1 |  | 0.757 | 3 | 2262 | -4.139 |  |  | 9.3256 | 52.519 | 52.119 | 13.6 | 10.4 | 200.6 | [5 empty] | 0.36496 |  | 0.7258
2014/03/16 18:18:38 | 35 | 36-UWSIF-UT Glut 120140020.27 | 36-UWSIF-UT Glut 1 |  | 0.757 | 4 |  |  | 1746 | -14.28 | 39.6972 | 43.956 |  |  |  |  | 43.268 | 1.0 | 1.3 | 3.3 | 1.09 |  | 1.1617
2014/03/16 18:18:38 | 35 | 36-UWSIF-UT Glut 120140020.27 | 36-UWSIF-UT Glut 1 |  | 0.757 | 5 |  |  | 3293 | 0 | 58.885 | 65.202 |  |  |  |  | 64.176 | 2.4 | 2.9 | 5.3 | 1.1057 |  | 1.1768
2014/03/16 18:27:07 | 36 | 39-UWSIF-UW Glut 220140020.23 | 39-UWSIF-UW Glut 2 |  | 0.7252 | 1 | 2960 | 0.022 |  |  | 10.8266 | 58.411 | 57.967 | 13.5 | 10.4 | 205.7 | [5 empty] | 0.36648 |  | 0.72883
2014/03/16 18:27:07 | 36 | 39-UWSIF-UW Glut 220140020.23 | 39-UWSIF-UW Glut 2 |  | 0.7252 | 2 | 2963 | 0 |  |  | 10.8363 | 58.463 | 58.018 | 13.4 | 10.3 | 212.5 | [5 empty] | 0.36647 |  | 0.72881
2014/03/16 18:27:07 | 36 | 39-UWSIF-UW Glut 220140020.23 | 39-UWSIF-UW Glut 2 |  | 0.7252 | 3 | 2353 | 28.385 |  |  | 10.1271 | 54.637 | 54.208 | 13.6 | 10.4 | 200.5 | [5 empty] | 0.37683 |  | 0.7495
2014/03/16 18:27:07 | 36 | 39-UWSIF-UW Glut 220140020.23 | 39-UWSIF-UW Glut 2 |  | 0.7252 | 4 |  |  | 1817 | 37.603 | 43.2379 | 45.865 |  |  |  |  | 45.122 | 1.0 | 1.3 | 3.3 | 1.1468 |  | 1.219
2014/03/16 18:27:07 | 36 | 39-UWSIF-UW Glut 220140020.23 | 39-UWSIF-UW Glut 2 |  | 0.7252 | 5 |  |  | 3295 | 0 | 61.4184 | 65.15 |  |  |  |  | 64.125 | 2.5 | 3.0 | 5.3 | 1.1057 |  | 1.1771
2014/03/16 18:35:36 | 37 | 39-UWSIF-UW Glut 220140020.24 | 39-UWSIF-UW Glut 2 |  | 0.7205 | 1 | 2962 | 0.049 |  |  | 10.9067 | 58.461 | 58.017 | 13.5 | 10.4 | 206.7 | [5 empty] | 0.36649 |  | 0.72884
2014/03/16 18:35:36 | 37 | 39-UWSIF-UW Glut 220140020.24 | 39-UWSIF-UW Glut 2 |  | 0.7205 | 2 | 2964 | 0 |  |  | 10.9106 | 58.482 | 58.038 | 13.4 | 10.3 | 213.8 | [5 empty] | 0.36647 |  | 0.7288
2014/03/16 18:35:36 | 37 | 39-UWSIF-UW Glut 220140020.24 | 39-UWSIF-UW Glut 2 |  | 0.7205 | 3 | 2342 | 28.378 |  |  | 10.1503 | 54.407 | 53.98 | 13.6 | 10.4 | 202.0 | [5 empty] | 0.37683 |  | 0.74949
2014/03/16 18:35:36 | 37 | 39-UWSIF-UW Glut 220140020.24 | 39-UWSIF-UW Glut 2 |  | 0.7205 | 4 |  |  | 1806 | 37.661 | 43.2157 | 45.545 |  |  |  |  | 44.806 | 1.0 | 1.3 | 3.3 | 1.1468 |  | 1.2191
2014/03/16 18:35:36 | 37 | 39-UWSIF-UW Glut 220140020.24 | 39-UWSIF-UW Glut 2 |  | 0.7205 | 5 |  |  | 3295 | 0 | 61.8074 | 65.138 |  |  |  |  | 64.113 | 2.4 | 3.0 | 5.3 | 1.1057 |  | 1.1771
2014/03/16 18:44:05 | 38 | 01-UWSIF-Liver 20140020.23 | 01-UWSIF-Liver |  | 0.7589 | 1 | 2961 | -0.006 |  |  | 10.3435 | 58.398 | 57.954 | 13.5 | 10.4 | 207.1 | [5 empty] | 0.36647 |  | 0.72882
2014/03/16 18:44:05 | 38 | 01-UWSIF-Liver 20140020.23 | 01-UWSIF-Liver |  | 0.7589 | 2 | 2964 | 0 |  |  | 10.3568 | 58.473 | 58.028 | 13.4 | 10.3 | 214.4 | [5 empty] | 0.36647 |  | 0.72883
2014/03/16 18:44:05 | 38 | 01-UWSIF-Liver 20140020.23 | 01-UWSIF-Liver |  | 0.7589 | 3 | 3121 | 7.294 |  |  | 12.7321 | 71.883 | 71.329 | 13.5 | 10.4 | 202.5 | [5 empty] | 0.36913 |  | 0.73414
2014/03/16 18:44:05 | 38 | 01-UWSIF-Liver 20140020.23 | 01-UWSIF-Liver |  | 0.7589 | 4 |  |  | 2142 | -3.808 | 48.7568 | 54.123 |  |  |  |  | 53.27 | 1.0 | 1.2 | 3.2 | 1.1015 |  | 1.1733
2014/03/16 18:44:05 | 38 | 01-UWSIF-Liver 20140020.23 | 01-UWSIF-Liver |  | 0.7589 | 5 |  |  | 3288 | 0 | 58.6599 | 65.116 |  |  |  |  | 64.091 | 2.6 | 3.1 | 5.5 | 1.1057 |  | 1.1769
2014/03/16 18:52:34 | 39 | 01-UWSIF-Liver 20140020.24 | 01-UWSIF-Liver |  | 0.7939 | 1 | 2959 | 0.041 |  |  | 9.8855 | 58.386 | 57.942 | 13.5 | 10.4 | 209.2 | [5 empty] | 0.36649 |  | 0.72883
2014/03/16 18:52:34 | 39 | 01-UWSIF-Liver 20140020.24 | 01-UWSIF-Liver |  | 0.7939 | 2 | 2959 | 0 |  |  | 9.8824 | 58.368 | 57.924 | 13.4 | 10.3 | 216.0 | [5 empty] | 0.36647 |  | 0.7288
2014/03/16 18:52:34 | 39 | 01-UWSIF-Liver 20140020.24 | 01-UWSIF-Liver |  | 0.7939 | 3 | 3267 | 7.289 |  |  | 12.7497 | 75.303 | 74.722 | 13.5 | 10.3 | 203.9 | [5 empty] | 0.36913 |  | 0.73412
2014/03/16 18:52:34 | 39 | 01-UWSIF-Liver 20140020.24 | 01-UWSIF-Liver |  | 0.7939 | 4 |  |  | 2228 | -3.877 | 48.6934 | 56.545 |  |  |  |  | 55.654 | 1.0 | 1.2 | 3.2 | 1.1014 |  | 1.1732
2014/03/16 18:52:34 | 39 | 01-UWSIF-Liver 20140020.24 | 01-UWSIF-Liver |  | 0.7939 | 5 |  |  | 3273 | 0 | 55.7733 | 64.767 |  |  |  |  | 63.748 | 2.6 | 3.1 | 5.5 | 1.1057 |  | 1.1769
2014/03/16 19:01:03 | 40 | 20140020.025 | 355 |  | 0.851 | 1 | 2942 | 0.025 |  |  | 9.173 | 58.074 | 57.633 | 13.4 | 10.4 | 209.2 | [5 empty] | 0.36648 |  | 0.72884
2014/03/16 19:01:03 | 40 | 20140020.025 | 355 |  | 0.851 | 2 | 2942 | 0 |  |  | 9.1654 | 58.026 | 57.585 | 13.4 | 10.2 | 215.7 | [5 empty] | 0.36647 |  | 0.72882
2014/03/16 19:01:03 | 40 | 20140020.025 | 355 |  | 0.851 | 3 | 2382 | 12.972 |  |  | 8.7632 | 55.48 | 55.051 | 13.5 | 10.3 | 203.4 | [5 empty] | 0.37121 |  | 0.73827
2014/03/16 19:01:03 | 40 | 20140020.025 | 355 |  | 0.851 | 4 |  |  | 1651 | -13.689 | 33.6226 | 41.852 |  |  |  |  | 41.197 | 1.0 | 1.2 | 3.2 | 1.0907 |  | 1.1624
2014/03/16 19:01:03 | 40 | 20140020.025 | 355 |  | 0.851 | 5 |  |  | 3262 | 0 | 51.8619 | 64.556 |  |  |  |  | 63.54 | 2.4 | 2.8 | 5.2 | 1.1057 |  | 1.1769
2014/03/16 19:09:32 | 41 | 20140020.026 | 306 |  | 0.856 | 1 | 2933 | 0.02 |  |  | 9.0896 | 57.884 | 57.444 | 13.3 | 10.3 | 207.8 | [5 empty] | 0.36648 |  | 0.72884 | Sample lost while crimping, picked up sample off the floor, but was unsure if was correct sample. Since # is drastically different from other samples, decided sample was lost.
2014/03/16 19:09:32 | 41 | 20140020.026 | 306 |  | 0.856 | 2 | 2930 | 0 |  |  | 9.0877 | 57.872 | 57.432 | 13.3 | 10.1 | 214.4 | [5 empty] | 0.36647 |  | 0.72882
2014/03/16 19:09:32 | 41 | 20140020.026 | 306 |  | 0.856 | 3 | 827 | 7.963 |  |  | 3.0848 | 19.645 | 19.495 | 13.4 | 10.2 | 202.1 | [5 empty] | 0.36938 |  | 0.73463
2014/03/16 19:09:32 | 41 | 20140020.026 | 306 |  | 0.856 | 4 |  |  | 1056 | -5.282 | 21.4864 | 26.903 |  |  |  |  | 26.479 | 0.9 | 1.2 | 3.2 | 1.0999 |  | 1.1717
2014/03/16 19:09:32 | 41 | 20140020.026 | 306 |  | 0.856 | 5 |  |  | 3264 | 0 | 51.5434 | 64.537 |  |  |  |  | 63.521 | 2.1 | 2.5 | 4.8 | 1.1057 |  | 1.1769
2014/03/16 19:18:01 | 42 | 20140020.027 | 469 |  | 0.808 | 1 | 2934 | 0.026 |  |  | 9.6383 | 57.937 | 57.496 | 13.3 | 10.2 | 205.4 | [5 empty] | 0.36648 |  | 0.72884
2014/03/16 19:18:01 | 42 | 20140020.027 | 469 |  | 0.808 | 2 | 2930 | 0 |  |  | 9.6256 | 57.86 | 57.421 | 13.2 | 10.1 | 212.5 | [5 empty] | 0.36647 |  | 0.72882
2014/03/16 19:18:01 | 42 | 20140020.027 | 469 |  | 0.808 | 3 | 2878 | 13.627 |  |  | 11.0434 | 66.383 | 65.869 | 13.3 | 10.2 | 200.4 | [5 empty] | 0.37145 |  | 0.73875
2014/03/16 19:18:01 | 42 | 20140020.027 | 469 |  | 0.808 | 4 |  |  | 2080 | -14.893 | 44.6311 | 52.748 |  |  |  |  | 51.923 | 0.9 | 1.1 | 3.1 | 1.0894 |  | 1.1611
2014/03/16 19:18:01 | 42 | 20140020.027 | 469 |  | 0.808 | 5 |  |  | 3247 | 0 | 54.4604 | 64.365 |  |  |  |  | 63.352 | 2.5 | 2.9 | 5.4 | 1.1057 |  | 1.1769
2014/03/16 19:26:30 | 43 | 20140020.028 | 467 |  | 0.816 | 1 | 2928 | 0.027 |  |  | 9.5045 | 57.699 | 57.26 | 13.3 | 10.2 | 208.1 | [5 empty] | 0.36648 |  | 0.72882
2014/03/16 19:26:30 | 43 | 20140020.028 | 467 |  | 0.816 | 2 | 2925 | 0 |  |  | 9.5153 | 57.764 | 57.325 | 13.2 | 10.1 | 214.8 | [5 empty] | 0.36647 |  | 0.7288
2014/03/16 19:26:30 | 43 | 20140020.028 | 467 |  | 0.816 | 3 | 2687 | 13.439 |  |  | 10.2232 | 62.061 | 61.58 | 13.3 | 10.2 | 202.7 | [5 empty] | 0.37138 |  | 0.7386
2014/03/16 19:26:30 | 43 | 20140020.028 | 467 |  | 0.816 | 4 |  |  | 2337 | -15.6 | 49.9015 | 59.561 |  |  |  |  | 58.63 | 0.9 | 1.1 | 3.2 | 1.0886 |  | 1.1603
2014/03/16 19:26:30 | 43 | 20140020.028 | 467 |  | 0.816 | 5 |  |  | 3247 | 0 | 53.9121 | 64.348 |  |  |  |  | 63.336 | 2.6 | 3.0 | 5.6 | 1.1057 |  | 1.1768
2014/03/16 19:35:00 | 44 | 20140020.029 | 568 |  | 0.725 | 1 | 2925 | -0.012 |  |  | 10.713 | 57.782 | 57.342 | 13.3 | 10.2 | 209.2 | [5 empty] | 0.36647 |  | 0.72885
2014/03/16 19:35:00 | 44 | 20140020.029 | 568 |  | 0.725 | 2 | 2924 | 0 |  |  | 10.6904 | 57.66 | 57.222 | 13.2 | 10.1 | 215.9 | [5 empty] | 0.36647 |  | 0.72886
2014/03/16 19:35:00 | 44 | 20140020.029 | 568 |  | 0.725 | 3 | 3000 | 14.735 |  |  | 12.9558 | 69.879 | 69.337 | 13.3 | 10.2 | 203.7 | [5 empty] | 0.37185 |  | 0.7396
2014/03/16 19:35:00 | 44 | 20140020.029 | 568 |  | 0.725 | 4 |  |  | 1930 | -14.723 | 46.3476 | 49.15 |  |  |  |  | 48.381 | 1.0 | 1.2 | 3.2 | 1.0896 |  | 1.1612
2014/03/16 19:35:00 | 44 | 20140020.029 | 568 |  | 0.725 | 5 |  |  | 3242 | 0 | 60.5552 | 64.217 |  |  |  |  | 63.206 | 2.4 | 2.9 | 5.3 | 1.1057 |  | 1.1768
2014/03/16 19:43:29 | 45 | 20140020.030 | 546 |  | 0.824 | 1 | 2925 | -0.032 |  |  | 9.4159 | 57.721 | 57.282 | 13.3 | 10.2 | 210.8 | [5 empty] | 0.36646 |  | 0.72879
2014/03/16 19:43:29 | 45 | 20140020.030 | 546 |  | 0.824 | 2 | 2921 | 0 |  |  | 9.3937 | 57.585 | 57.147 | 13.2 | 10.1 | 217.0 | [5 empty] | 0.36647 |  | 0.72882
2014/03/16 19:43:29 | 45 | 20140020.030 | 546 |  | 0.824 | 3 | 3068 | 13.853 |  |  | 11.6421 | 71.368 | 70.814 | 13.3 | 10.1 | 205.1 | [5 empty] | 0.37153 |  | 0.73891
2014/03/16 19:43:29 | 45 | 20140020.030 | 546 |  | 0.824 | 4 |  |  | 2198 | -14.932 | 46.8918 | 56.518 |  |  |  |  | 55.634 | 1.0 | 1.2 | 3.2 | 1.0893 |  | 1.161
2014/03/16 19:43:29 | 45 | 20140020.030 | 546 |  | 0.824 | 5 |  |  | 3247 | 0 | 53.3188 | 64.264 |  |  |  |  | 63.253 | 2.5 | 3.0 | 5.5 | 1.1057 |  | 1.1768
2014/03/16 19:51:59 | 46 | 20140020.031 | 953 |  | 0.829 | 1 | 2922 | -0.014 |  |  | 9.3499 | 57.664 | 57.225 | 13.3 | 10.2 | 211.7 | [5 empty] | 0.36647 |  | 0.72882
2014/03/16 19:51:59 | 46 | 20140020.031 | 953 |  | 0.829 | 2 | 2923 | 0 |  |  | 9.3448 | 57.632 | 57.194 | 13.2 | 10.1 | 218.6 | [5 empty] | 0.36647 |  | 0.72883
2014/03/16 19:51:59 | 46 | 20140020.031 | 953 |  | 0.829 | 3 | 2877 | 13.611 |  |  | 10.7973 | 66.591 | 66.075 | 13.3 | 10.2 | 206.7 | [5 empty] | 0.37144 |  | 0.73875
2014/03/16 19:51:59 | 46 | 20140020.031 | 953 |  | 0.829 | 4 |  |  | 2198 | -15.061 | 46.3123 | 56.158 |  |  |  |  | 55.28 | 1.0 | 1.2 | 3.2 | 1.0892 |  | 1.1608
2014/03/16 19:51:59 | 46 | 20140020.031 | 953 |  | 0.829 | 5 |  |  | 3243 | 0 | 52.9442 | 64.2 |  |  |  |  | 63.189 | 2.5 | 3.0 | 5.5 | 1.1057 |  | 1.1768
2014/03/16 20:00:28 | 47 | 20140020.032 | 978 |  | 0.806 | 1 | 2919 | 0.031 |  |  | 9.6079 | 57.611 | 57.173 | 13.3 | 10.2 | 212.5 | [5 empty] | 0.36648 |  | 0.72882
2014/03/16 20:00:28 | 47 | 20140020.032 | 978 |  | 0.806 | 2 | 2920 | 0 |  |  | 9.6037 | 57.586 | 57.148 | 13.2 | 10.1 | 219.3 | [5 empty] | 0.36647 |  | 0.72879
2014/03/16 20:00:28 | 47 | 20140020.032 | 978 |  | 0.806 | 3 | 1282 | 13.988 |  |  | 5.0306 | 30.165 | 29.932 | 13.3 | 10.1 | 206.7 | [5 empty] | 0.37158 |  | 0.73899
2014/03/16 20:00:28 | 47 | 20140020.032 | 978 |  | 0.806 | 4 |  |  | 906 | -13.954 | 19.2305 | 22.672 |  |  |  |  | 22.317 | 0.9 | 1.1 | 3.1 | 1.0904 |  | 1.1621
2014/03/16 20:00:28 | 47 | 20140020.032 | 978 |  | 0.806 | 5 |  |  | 3233 | 0 | 54.2819 | 63.996 |  |  |  |  | 62.989 | 1.9 | 2.3 | 4.6 | 1.1057 |  | 1.1768
2014/03/16 20:08:57 | 48 | 20140020.033 | 856 |  | 0.819 | 1 | 2912 | 0.055 |  |  | 9.4308 | 57.461 | 57.025 | 13.2 | 10.1 | 209.6 | [5 empty] | 0.36649 |  | 0.72885
2014/03/16 20:08:57 | 48 | 20140020.033 | 856 |  | 0.819 | 2 | 2915 | 0 |  |  | 9.4334 | 57.477 | 57.04 | 13.1 | 10.0 | 216.1 | [5 empty] | 0.36647 |  | 0.72881
2014/03/16 20:08:57 | 48 | 20140020.033 | 856 |  | 0.819 | 3 | 3099 | 15.5 |  |  | 11.8076 | 71.943 | 71.385 | 13.2 | 10.0 | 205.0 | [5 empty] | 0.37213 |  | 0.7401
2014/03/16 20:08:57 | 48 | 20140020.033 | 856 |  | 0.819 | 4 |  |  | 2101 | -14.557 | 44.8597 | 53.74 |  |  |  |  | 52.899 | 0.9 | 1.1 | 3.1 | 1.0897 |  | 1.1614
2014/03/16 20:08:57 | 48 | 20140020.033 | 856 |  | 0.819 | 5 |  |  | 3231 | 0 | 53.3808 | 63.948 |  |  |  |  | 62.942 | 2.4 | 2.9 | 5.4 | 1.1057 |  | 1.1768
2014/03/16 20:17:26 | 49 | 20140020.034 | 894 |  | 0.76 | 1 | 2907 | 0.005 |  |  | 10.1499 | 57.388 | 56.951 | 13.2 | 10.1 | 212.0 | [5 empty] | 0.36647 |  | 0.72884
2014/03/16 20:17:26 | 49 | 20140020.034 | 894 |  | 0.76 | 2 | 2909 | 0 |  |  | 10.1583 | 57.435 | 56.999 | 13.1 | 10.1 | 218.8 | [5 empty] | 0.36647 |  | 0.72884
2014/03/16 20:17:26 | 49 | 20140020.034 | 894 |  | 0.76 | 3 | 2985 | 15.448 |  |  | 12.3008 | 69.549 | 69.009 | 13.2 | 10.1 | 206.6 | [5 empty] | 0.37211 |  | 0.7401
2014/03/16 20:17:26 | 49 | 20140020.034 | 894 |  | 0.76 | 4 |  |  | 1870 | -14.15 | 42.8446 | 47.629 |  |  |  |  | 46.883 | 0.9 | 1.1 | 3.1 | 1.0902 |  | 1.1619
2014/03/16 20:17:26 | 49 | 20140020.034 | 894 |  | 0.76 | 5 |  |  | 3226 | 0 | 57.4807 | 63.899 |  |  |  |  | 62.894 | 2.4 | 2.8 | 5.2 | 1.1057 |  | 1.1768
2014/03/16 20:25:55 | 50 | 20140020.035 | 702 |  | 0.805 | 1 | 2907 | 0.027 |  |  | 9.5754 | 57.345 | 56.909 | 13.2 | 10.1 | 213.1 | [5 empty] | 0.36648 |  | 0.72887
2014/03/16 20:25:55 | 50 | 20140020.035 | 702 |  | 0.805 | 2 | 2904 | 0 |  |  | 9.5704 | 57.315 | 56.88 | 13.1 | 10.0 | 219.7 | [5 empty] | 0.36647 |  | 0.72885
2014/03/16 20:25:55 | 50 | 20140020.035 | 702 |  | 0.805 | 3 | 2727 | 14.633 |  |  | 10.5212 | 63.01 | 62.521 | 13.2 | 10.0 | 207.3 | [5 empty] | 0.37182 |  | 0.73952
2014/03/16 20:25:55 | 50 | 20140020.035 | 702 |  | 0.805 | 4 |  |  | 2243 | -15.391 | 48.6629 | 57.3 |  |  |  |  | 56.404 | 0.9 | 1.1 | 3.1 | 1.0888 |  | 1.1605
2014/03/16 20:25:55 | 50 | 20140020.035 | 702 |  | 0.805 | 5 |  |  | 3222 | 0 | 54.1655 | 63.779 |  |  |  |  | 62.775 | 2.5 | 3.0 | 5.4 | 1.1057 |  | 1.1768
2014/03/16 20:34:25 | 51 | 20140020.036 | 750 |  | 0.771 | 1 | 2902 | 0.017 |  |  | 9.9874 | 57.286 | 56.851 | 13.2 | 10.1 | 213.3 | [5 empty] | 0.36648 |  | 0.72881
2014/03/16 20:34:25 | 51 | 20140020.036 | 750 |  | 0.771 | 2 | 2905 | 0 |  |  | 9.988 | 57.29 | 56.854 | 13.1 | 10.0 | 220.4 | [5 empty] | 0.36647 |  | 0.7288
2014/03/16 20:34:25 | 51 | 20140020.036 | 750 |  | 0.771 | 3 | 2562 | 14.526 |  |  | 10.3797 | 59.536 | 59.075 | 13.2 | 10.1 | 208.6 | [5 empty] | 0.37178 |  | 0.73939
2014/03/16 20:34:25 | 51 | 20140020.036 | 750 |  | 0.771 | 4 |  |  | 2134 | -15.485 | 48.4404 | 54.629 |  |  |  |  | 53.775 | 0.9 | 1.1 | 3.1 | 1.0887 |  | 1.1604
2014/03/16 20:34:25 | 51 | 20140020.036 | 750 |  | 0.771 | 5 |  |  | 3226 | 0 | 56.6056 | 63.837 |  |  |  |  | 62.833 | 2.5 | 2.9 | 5.4 | 1.1057 |  | 1.1768
2014/03/16 20:42:54 | 52 | 20140020.037 | 610 |  | 0.818 | 1 | 2906 | 0.009 |  |  | 9.4195 | 57.322 | 56.886 | 13.2 | 10.1 | 213.9 | [5 empty] | 0.36647 |  | 0.72882
2014/03/16 20:42:54 | 52 | 20140020.037 | 610 |  | 0.818 | 2 | 2903 | 0 |  |  | 9.4158 | 57.3 | 56.865 | 13.1 | 10.0 | 220.5 | [5 empty] | 0.36647 |  | 0.72881
2014/03/16 20:42:54 | 52 | 20140020.037 | 610 |  | 0.818 | 3 | 2952 | 14.622 |  |  | 11.2434 | 68.422 | 67.891 | 13.2 | 10.1 | 209.0 | [5 empty] | 0.37181 |  | 0.73947
2014/03/16 20:42:54 | 52 | 20140020.037 | 610 |  | 0.818 | 4 |  |  | 2390 | -15.998 | 51.1251 | 61.171 |  |  |  |  | 60.215 | 0.9 | 1.1 | 3.2 | 1.0882 |  | 1.1598
2014/03/16 20:42:54 | 52 | 20140020.037 | 610 |  | 0.818 | 5 |  |  | 3235 | 0 | 53.5196 | 64.036 |  |  |  |  | 63.029 | 2.6 | 3.1 | 5.5 | 1.1057 |  | 1.1768
2014/03/16 20:51:23 | 53 | 20140020.038 | 660 |  | 0.775 | 1 | 2912 | 0.034 |  |  | 9.9748 | 57.511 | 57.073 | 13.2 | 10.2 | 216.6 | [5 empty] | 0.36648 |  | 0.72886
2014/03/16 20:51:23 | 53 | 20140020.038 | 660 |  | 0.775 | 2 | 2911 | 0 |  |  | 9.9683 | 57.473 | 57.036 | 13.2 | 10.1 | 223.4 | [5 empty] | 0.36647 |  | 0.72884
2014/03/16 20:51:23 | 53 | 20140020.038 | 660 |  | 0.775 | 3 | 3017 | 14.694 |  |  | 12.176 | 70.202 | 69.657 | 13.3 | 10.1 | 211.3 | [5 empty] | 0.37184 |  | 0.73955
2014/03/16 20:51:23 | 53 | 20140020.038 | 660 |  | 0.775 | 4 |  |  | 2052 | -14.8 | 46.2791 | 52.462 |  |  |  |  | 51.642 | 0.9 | 1.1 | 3.2 | 1.0895 |  | 1.1611
2014/03/16 20:51:23 | 53 | 20140020.038 | 660 |  | 0.775 | 5 |  |  | 3233 | 0 | 56.3823 | 63.915 |  |  |  |  | 62.909 | 2.5 | 2.9 | 5.3 | 1.1057 |  | 1.1768
2014/03/16 20:59:53 | 54 | 20140020.039 | 1089 |  | 0.779 | 1 | 2913 | 0.028 |  |  | 9.9232 | 57.509 | 57.071 | 13.2 | 10.1 | 217.6 | [5 empty] | 0.36648 |  | 0.72885
2014/03/16 20:59:53 | 54 | 20140020.039 | 1089 |  | 0.779 | 2 | 2909 | 0 |  |  | 9.9065 | 57.412 | 56.975 | 13.1 | 10.0 | 223.7 | [5 empty] | 0.36647 |  | 0.72883
2014/03/16 20:59:53 | 54 | 20140020.039 | 1089 |  | 0.779 | 3 | 2697 | 13.566 |  |  | 10.8034 | 62.61 | 62.124 | 13.2 | 10.1 | 211.8 | [5 empty] | 0.37143 |  | 0.73872
2014/03/16 20:59:53 | 54 | 20140020.039 | 1089 |  | 0.779 | 4 |  |  | 2062 | -14.943 | 46.1475 | 52.583 |  |  |  |  | 51.761 | 0.9 | 1.1 | 3.2 | 1.0893 |  | 1.161
2014/03/16 20:59:53 | 54 | 20140020.039 | 1089 |  | 0.779 | 5 |  |  | 3243 | 0 | 56.2957 | 64.146 |  |  |  |  | 63.137 | 2.4 | 2.9 | 5.3 | 1.1057 |  | 1.1768
2014/03/16 21:08:23 | 55 | 20140020.040 | 1070 |  | 0.777 | 1 | 2918 | 0.014 |  |  | 9.9726 | 57.646 | 57.208 | 13.2 | 10.1 | 218.3 | [5 empty] | 0.36648 |  | 0.72883
2014/03/16 21:08:23 | 55 | 20140020.040 | 1070 |  | 0.777 | 2 | 2916 | 0 |  |  | 9.9617 | 57.583 | 57.145 | 13.2 | 10.0 | 225.1 | [5 empty] | 0.36647 |  | 0.72882
2014/03/16 21:08:23 | 55 | 20140020.040 | 1070 |  | 0.777 | 3 | 2607 | 14.632 |  |  | 10.4417 | 60.358 | 59.89 | 13.3 | 10.1 | 212.6 | [5 empty] | 0.37181 |  | 0.73948
2014/03/16 21:08:23 | 55 | 20140020.040 | 1070 |  | 0.777 | 4 |  |  | 2235 | -16.647 | 50.2482 | 57.109 |  |  |  |  | 56.216 | 0.9 | 1.1 | 3.1 | 1.0875 |  | 1.1591
2014/03/16 21:08:23 | 55 | 20140020.040 | 1070 |  | 0.777 | 5 |  |  | 3239 | 0 | 56.4151 | 64.118 |  |  |  |  | 63.109 | 2.5 | 3.0 | 5.5 | 1.1057 |  | 1.1768
2014/03/16 21:16:52 | 56 | 20140020.041 | 1103 |  | 0.805 | 1 | 2919 | 0.01 |  |  | 9.6107 | 57.556 | 57.118 | 13.2 | 10.1 | 218.9 | [5 empty] | 0.36648 |  | 0.72884
2014/03/16 21:16:52 | 56 | 20140020.041 | 1103 |  | 0.805 | 2 | 2917 | 0 |  |  | 9.6125 | 57.567 | 57.13 | 13.1 | 10.1 | 225.5 | [5 empty] | 0.36647 |  | 0.72884
2014/03/16 21:16:52 | 56 | 20140020.041 | 1103 |  | 0.805 | 3 | 3031 | 13.9 |  |  | 11.7457 | 70.342 | 69.797 | 13.3 | 10.1 | 214.0 | [5 empty] | 0.37155 |  | 0.73897
2014/03/16 21:16:52 | 56 | 20140020.041 | 1103 |  | 0.805 | 4 |  |  | 2204 | -14.718 | 47.902 | 56.404 |  |  |  |  | 55.522 | 0.9 | 1.1 | 3.1 | 1.0896 |  | 1.1612
2014/03/16 21:16:52 | 56 | 20140020.041 | 1103 |  | 0.805 | 5 |  |  | 3243 | 0 | 54.5078 | 64.182 |  |  |  |  | 63.172 | 2.5 | 3.0 | 5.5 | 1.1057 |  | 1.1768
2014/03/16 21:25:22 | 57 | 20140020.042 | 1186 |  | 0.792 | 1 | 2920 | 0.002 |  |  | 9.7873 | 57.667 | 57.229 | 13.2 | 10.1 | 220.5 | [5 empty] | 0.36647 |  | 0.72883
2014/03/16 21:25:22 | 57 | 20140020.042 | 1186 |  | 0.792 | 2 | 2920 | 0 |  |  | 9.7841 | 57.649 | 57.21 | 13.2 | 10.1 | 227.0 | [5 empty] | 0.36647 |  | 0.72883
2014/03/16 21:25:22 | 57 | 20140020.042 | 1186 |  | 0.792 | 3 | 1333 | 13.066 |  |  | 5.3468 | 31.504 | 31.261 | 13.3 | 10.1 | 214.8 | [5 empty] | 0.37124 |  | 0.73835
2014/03/16 21:25:22 | 57 | 20140020.042 | 1186 |  | 0.792 | 4 |  |  | 979 | -13.97 | 21.263 | 24.633 |  |  |  |  | 24.247 | 0.9 | 1.1 | 3.2 | 1.0904 |  | 1.1621
2014/03/16 21:25:22 | 57 | 20140020.042 | 1186 |  | 0.792 | 5 |  |  | 3239 | 0 | 55.2814 | 64.042 |  |  |  |  | 63.034 | 2.0 | 2.3 | 4.6 | 1.1057 |  | 1.1769
2014/03/16 21:33:51 | 58 | 20140020.043 | 1296 |  | 0.803 | 1 | 2915 | 0.007 |  |  | 9.6277 | 57.515 | 57.077 | 13.1 | 10.1 | 217.0 | [5 empty] | 0.36647 |  | 0.72883
2014/03/16 21:33:51 | 58 | 20140020.043 | 1296 |  | 0.803 | 2 | 2915 | 0 |  |  | 9.6279 | 57.516 | 57.079 | 13.1 | 10.0 | 224.3 | [5 empty] | 0.36647 |  | 0.72883
2014/03/16 21:33:51 | 58 | 20140020.043 | 1296 |  | 0.803 | 3 | 3135 | 14.771 |  |  | 12.1993 | 72.877 | 72.312 | 13.2 | 10.0 | 212.7 | [5 empty] | 0.37187 |  | 0.73959
2014/03/16 21:33:51 | 58 | 20140020.043 | 1296 |  | 0.803 | 4 |  |  | 2046 | -14.504 | 44.5086 | 52.278 |  |  |  |  | 51.46 | 0.8 | 1.1 | 3.1 | 1.0898 |  | 1.1615
2014/03/16 21:33:51 | 58 | 20140020.043 | 1296 |  | 0.803 | 5 |  |  | 3234 | 0 | 54.4727 | 63.981 |  |  |  |  | 62.975 | 2.4 | 2.9 | 5.3 | 1.1057 |  | 1.1768
2014/03/16 21:42:21 | 59 | 20140020.044 | 1232 |  | 0.75 | 1 | 2909 | 0.005 |  |  | 10.2898 | 57.413 | 56.977 | 13.2 | 10.1 | 219.8 | [5 empty] | 0.36647 |  | 0.72883
2014/03/16 21:42:21 | 59 | 20140020.044 | 1232 |  | 0.75 | 2 | 2913 | 0 |  |  | 10.2902 | 57.416 | 56.979 | 13.1 | 10.0 | 226.4 | [5 empty] | 0.36647 |  | 0.72882
2014/03/16 21:42:21 | 59 | 20140020.044 | 1232 |  | 0.75 | 3 | 2831 | 14.222 |  |  | 11.7888 | 65.777 | 65.267 | 13.2 | 10.0 | 213.9 | [5 empty] | 0.37167 |  | 0.73919
2014/03/16 21:42:21 | 59 | 20140020.044 | 1232 |  | 0.75 | 4 |  |  | 1906 | -14.52 | 44.117 | 48.398 |  |  |  |  | 47.641 | 0.9 | 1.1 | 3.1 | 1.0898 |  | 1.1614
2014/03/16 21:42:21 | 59 | 20140020.044 | 1232 |  | 0.75 | 5 |  |  | 3236 | 0 | 58.3517 | 64.014 |  |  |  |  | 63.007 | 2.4 | 2.8 | 5.3 | 1.1057 |  | 1.1768
2014/03/16 21:50:50 | 60 | 20140020.045 | 1337 |  | 0.842 | 1 | 2911 | 0.028 |  |  | 9.1756 | 57.476 | 57.039 | 13.2 | 10.1 | 220.3 | [5 empty] | 0.36648 |  | 0.72887
2014/03/16 21:50:50 | 60 | 20140020.045 | 1337 |  | 0.842 | 2 | 2909 | 0 |  |  | 9.1596 | 57.377 | 56.94 | 13.1 | 10.0 | 226.5 | [5 empty] | 0.36647 |  | 0.72885
2014/03/16 21:50:50 | 60 | 20140020.045 | 1337 |  | 0.842 | 3 | 3367 | 14.908 |  |  | 12.4503 | 77.989 | 77.384 | 13.2 | 10.0 | 214.4 | [5 empty] | 0.37191 |  | 0.73971
2014/03/16 21:50:50 | 60 | 20140020.045 | 1337 |  | 0.842 | 4 |  |  | 2166 | -14.094 | 44.8209 | 55.202 |  |  |  |  | 54.338 | 0.9 | 1.1 | 3.1 | 1.0902 |  | 1.1619
2014/03/16 21:50:50 | 60 | 20140020.045 | 1337 |  | 0.842 | 5 |  |  | 3236 | 0 | 51.9506 | 63.983 |  |  |  |  | 62.976 | 2.5 | 3.0 | 5.4 | 1.1057 |  | 1.1768
2014/03/16 21:59:20 | 61 | 20140020.046 | 1341 |  | 0.805 | 1 | 2912 | -0.001 |  |  | 9.5952 | 57.464 | 57.027 | 13.2 | 10.1 | 222.1 | [5 empty] | 0.36647 |  | 0.72885
2014/03/16 21:59:20 | 61 | 20140020.046 | 1341 |  | 0.805 | 2 | 2911 | 0 |  |  | 9.5804 | 57.375 | 56.939 | 13.1 | 10.0 | 228.6 | [5 empty] | 0.36647 |  | 0.72885
2014/03/16 21:59:20 | 61 | 20140020.046 | 1341 |  | 0.805 | 3 | 3088 | 14.485 |  |  | 12.0882 | 72.394 | 71.832 | 13.2 | 10.1 | 216.2 | [5 empty] | 0.37176 |  | 0.73941
2014/03/16 21:59:20 | 61 | 20140020.046 | 1341 |  | 0.805 | 4 |  |  | 2048 | -14.911 | 44.8467 | 52.806 |  |  |  |  | 51.981 | 0.9 | 1.1 | 3.1 | 1.0894 |  | 1.161
2014/03/16 21:59:20 | 61 | 20140020.046 | 1341 |  | 0.805 | 5 |  |  | 3233 | 0 | 54.3267 | 63.969 |  |  |  |  | 62.962 | 2.5 | 2.9 | 5.3 | 1.1057 |  | 1.1768
2014/03/16 22:07:49 | 62 | 20140020.047 | 1596 |  | 0.805 | 1 | 2909 | 0.019 |  |  | 9.5828 | 57.389 | 56.953 | 13.2 | 10.1 | 222.6 | [5 empty] | 0.36648 |  | 0.72882
2014/03/16 22:07:49 | 62 | 20140020.047 | 1596 |  | 0.805 | 2 | 2911 | 0 |  |  | 9.5854 | 57.405 | 56.969 | 13.1 | 10.0 | 229.0 | [5 empty] | 0.36647 |  | 0.7288
2014/03/16 22:07:49 | 62 | 20140020.047 | 1596 |  | 0.805 | 3 | 2911 | 12.685 |  |  | 11.3171 | 67.776 | 67.251 | 13.2 | 10.1 | 216.8 | [5 empty] | 0.3711 |  | 0.73805
2014/03/16 22:07:49 | 62 | 20140020.047 | 1596 |  | 0.805 | 4 |  |  | 1847 | -13.803 | 39.89 | 46.97 |  |  |  |  | 46.235 | 0.9 | 1.1 | 3.1 | 1.0906 |  | 1.1622
2014/03/16 22:07:49 | 62 | 20140020.047 | 1596 |  | 0.805 | 5 |  |  | 3228 | 0 | 54.2607 | 63.891 |  |  |  |  | 62.886 | 2.4 | 2.8 | 5.3 | 1.1057 |  | 1.1768
2014/03/16 22:16:19 | 63 | 20140020.048 | 1598 |  | 0.845 | 1 | 2904 | 0.025 |  |  | 9.0982 | 57.195 | 56.76 | 13.1 | 10.1 | 222.5 | [5 empty] | 0.36648 |  | 0.72885
2014/03/16 22:16:19 | 63 | 20140020.048 | 1598 |  | 0.845 | 2 | 2907 | 0 |  |  | 9.115 | 57.3 | 56.864 | 13.1 | 10.0 | 229.0 | [5 empty] | 0.36647 |  | 0.72883
2014/03/16 22:16:19 | 63 | 20140020.048 | 1598 |  | 0.845 | 3 | 2793 | 13.157 |  |  | 10.3181 | 64.863 | 64.361 | 13.2 | 10.0 | 216.8 | [5 empty] | 0.37128 |  | 0.73842
2014/03/16 22:16:19 | 63 | 20140020.048 | 1598 |  | 0.845 | 4 |  |  | 2248 | -16.092 | 46.5199 | 57.498 |  |  |  |  | 56.6 | 0.9 | 1.1 | 3.1 | 1.0881 |  | 1.1597
2014/03/16 22:16:19 | 63 | 20140020.048 | 1598 |  | 0.845 | 5 |  |  | 3231 | 0 | 51.6848 | 63.882 |  |  |  |  | 62.877 | 2.5 | 3.0 | 5.5 | 1.1057 |  | 1.1768
2014/03/16 22:24:49 | 64 | 36-UWSIF-UT Glut 120140020.28 | 36-UWSIF-UT Glut 1 |  | 0.7298 | 1 | 2910 | 0.057 |  |  | 10.5759 | 57.42 | 56.983 | 13.2 | 10.1 | 223.7 | [5 empty] | 0.36649 |  | 0.72887
2014/03/16 22:24:49 | 64 | 36-UWSIF-UT Glut 120140020.28 | 36-UWSIF-UT Glut 1 |  | 0.7298 | 2 | 2908 | 0 |  |  | 10.5595 | 57.331 | 56.895 | 13.1 | 10.0 | 229.7 | [5 empty] | 0.36647 |  | 0.72883
2014/03/16 22:24:49 | 64 | 36-UWSIF-UT Glut 120140020.28 | 36-UWSIF-UT Glut 1 |  | 0.7298 | 3 | 2126 | -4.113 |  |  | 9.1328 | 49.585 | 49.208 | 13.2 | 10.0 | 217.2 | [5 empty] | 0.36497 |  | 0.72583
2014/03/16 22:24:49 | 64 | 36-UWSIF-UT Glut 120140020.28 | 36-UWSIF-UT Glut 1 |  | 0.7298 | 4 |  |  | 1649 | -14.295 | 39.0134 | 41.646 |  |  |  |  | 40.995 | 0.9 | 1.1 | 3.1 | 1.09 |  | 1.1617
2014/03/16 22:24:49 | 64 | 36-UWSIF-UT Glut 120140020.28 | 36-UWSIF-UT Glut 1 |  | 0.7298 | 5 |  |  | 3224 | 0 | 59.8075 | 63.844 |  |  |  |  | 62.839 | 2.3 | 2.7 | 5.1 | 1.1057 |  | 1.1769
2014/03/16 22:33:18 | 65 | 36-UWSIF-UT Glut 120140020.29 | 36-UWSIF-UT Glut 1 |  | 0.7903 | 1 | 2902 | -0.014 |  |  | 9.7353 | 57.238 | 56.803 | 13.1 | 10.1 | 222.2 | [5 empty] | 0.36647 |  | 0.72883
2014/03/16 22:33:18 | 65 | 36-UWSIF-UT Glut 120140020.29 | 36-UWSIF-UT Glut 1 |  | 0.7903 | 2 | 2902 | 0 |  |  | 9.7337 | 57.229 | 56.794 | 13.1 | 10.0 | 228.8 | [5 empty] | 0.36647 |  | 0.72884
2014/03/16 22:33:18 | 65 | 36-UWSIF-UT Glut 120140020.29 | 36-UWSIF-UT Glut 1 |  | 0.7903 | 3 | 2316 | -4.19 |  |  | 9.1316 | 53.689 | 53.28 | 13.2 | 10.0 | 216.3 | [5 empty] | 0.36494 |  | 0.72578
2014/03/16 22:33:18 | 65 | 36-UWSIF-UT Glut 120140020.29 | 36-UWSIF-UT Glut 1 |  | 0.7903 | 4 |  |  | 1789 | -14.293 | 39.0109 | 45.096 |  |  |  |  | 44.39 | 0.8 | 1.1 | 3.0 | 1.09 |  | 1.1617
2014/03/16 22:33:18 | 65 | 36-UWSIF-UT Glut 120140020.29 | 36-UWSIF-UT Glut 1 |  | 0.7903 | 5 |  |  | 3218 | 0 | 55.1557 | 63.759 |  |  |  |  | 62.756 | 2.3 | 2.8 | 5.2 | 1.1057 |  | 1.1769
2014/03/16 22:41:48 | 66 | 39-UWSIF-UW Glut 220140020.25 | 39-UWSIF-UW Glut 2 |  | 0.7327 | 1 | 2895 | -0.014 |  |  | 10.4806 | 57.129 | 56.695 | 13.1 | 10.0 | 221.2 | [5 empty] | 0.36647 |  | 0.72881
2014/03/16 22:41:48 | 66 | 39-UWSIF-UW Glut 220140020.25 | 39-UWSIF-UW Glut 2 |  | 0.7327 | 2 | 2892 | 0 |  |  | 10.4676 | 57.058 | 56.624 | 13.0 | 9.9 | 227.4 | [5 empty] | 0.36647 |  | 0.72882
2014/03/16 22:41:48 | 66 | 39-UWSIF-UW Glut 220140020.25 | 39-UWSIF-UW Glut 2 |  | 0.7327 | 3 | 2309 | 28.361 |  |  | 9.8719 | 53.811 | 53.389 | 13.1 | 10.0 | 215.8 | [5 empty] | 0.37683 |  | 0.74949
2014/03/16 22:41:48 | 66 | 39-UWSIF-UW Glut 220140020.25 | 39-UWSIF-UW Glut 2 |  | 0.7327 | 4 |  |  | 1782 | 37.686 | 42.1528 | 45.177 |  |  |  |  | 44.444 | 0.8 | 1.1 | 3.1 | 1.1468 |  | 1.2191
2014/03/16 22:41:48 | 66 | 39-UWSIF-UW Glut 220140020.25 | 39-UWSIF-UW Glut 2 |  | 0.7327 | 5 |  |  | 3211 | 0 | 59.3203 | 63.576 |  |  |  |  | 62.575 | 2.3 | 2.8 | 5.2 | 1.1057 |  | 1.1771
2014/03/16 22:50:18 | 67 | 39-UWSIF-UW Glut 220140020.26 | 39-UWSIF-UW Glut 2 |  | 0.7273 | 1 | 2891 | 0.013 |  |  | 10.5385 | 57.021 | 56.587 | 13.1 | 10.0 | 221.2 | [5 empty] | 0.36648 |  | 0.72882
2014/03/16 22:50:18 | 67 | 39-UWSIF-UW Glut 220140020.26 | 39-UWSIF-UW Glut 2 |  | 0.7273 | 2 | 2886 | 0 |  |  | 10.5216 | 56.93 | 56.497 | 13.0 | 9.9 | 227.0 | [5 empty] | 0.36647 |  | 0.72881
2014/03/16 22:50:18 | 67 | 39-UWSIF-UW Glut 220140020.26 | 39-UWSIF-UW Glut 2 |  | 0.7273 | 3 | 2290 | 28.323 |  |  | 9.8528 | 53.311 | 52.892 | 13.1 | 10.0 | 215.2 | [5 empty] | 0.37681 |  | 0.74945
2014/03/16 22:50:18 | 67 | 39-UWSIF-UW Glut 220140020.26 | 39-UWSIF-UW Glut 2 |  | 0.7273 | 4 |  |  | 1772 | 37.701 | 42.0967 | 44.784 |  |  |  |  | 44.058 | 0.8 | 1.1 | 3.1 | 1.1469 |  | 1.2191
2014/03/16 22:50:18 | 67 | 39-UWSIF-UW Glut 220140020.26 | 39-UWSIF-UW Glut 2 |  | 0.7273 | 5 |  |  | 3204 | 0 | 59.5922 | 63.396 |  |  |  |  | 62.398 | 2.3 | 2.9 | 5.2 | 1.1057 |  | 1.1772
2014/03/16 22:58:48 | 68 | 01-UWSIF-Liver 20140020.25 | 01-UWSIF-Liver |  | 0.7591 | 1 | 2883 | -0.005 |  |  | 10.094 | 57.004 | 56.571 | 13.0 | 10.0 | 220.5 | [5 empty] | 0.36647 |  | 0.72888
2014/03/16 22:58:48 | 68 | 01-UWSIF-Liver 20140020.25 | 01-UWSIF-Liver |  | 0.7591 | 2 | 2880 | 0 |  |  | 10.0568 | 56.794 | 56.363 | 13.0 | 9.9 | 226.9 | [5 empty] | 0.36647 |  | 0.72888
2014/03/16 22:58:48 | 68 | 01-UWSIF-Liver 20140020.25 | 01-UWSIF-Liver |  | 0.7591 | 3 | 3049 | 7.236 |  |  | 12.4309 | 70.201 | 69.66 | 13.1 | 9.9 | 215.0 | [5 empty] | 0.36911 |  | 0.73415
2014/03/16 22:58:48 | 68 | 01-UWSIF-Liver 20140020.25 | 01-UWSIF-Liver |  | 0.7591 | 4 |  |  | 2095 | -3.846 | 47.7575 | 53.027 |  |  |  |  | 52.192 | 0.8 | 1.1 | 3.1 | 1.1015 |  | 1.1733
2014/03/16 22:58:48 | 68 | 01-UWSIF-Liver 20140020.25 | 01-UWSIF-Liver |  | 0.7591 | 5 |  |  | 3207 | 0 | 57.2183 | 63.532 |  |  |  |  | 62.532 | 2.5 | 3.0 | 5.4 | 1.1057 |  | 1.177
2014/03/16 23:07:18 | 69 | 01-UWSIF-Liver 20140020.26 | 01-UWSIF-Liver |  | 0.7337 | 1 | 2894 | 0.011 |  |  | 10.4561 | 57.073 | 56.639 | 13.1 | 10.1 | 223.3 | [5 empty] | 0.36648 |  | 0.72876
2014/03/16 23:07:18 | 69 | 01-UWSIF-Liver 20140020.26 | 01-UWSIF-Liver |  | 0.7337 | 2 | 2888 | 0 |  |  | 10.4337 | 56.951 | 56.518 | 13.0 | 10.0 | 229.8 | [5 empty] | 0.36647 |  | 0.72875
2014/03/16 23:07:18 | 69 | 01-UWSIF-Liver 20140020.26 | 01-UWSIF-Liver |  | 0.7337 | 3 | 2946 | 7.256 |  |  | 12.4608 | 68.016 | 67.491 | 13.1 | 10.0 | 217.5 | [5 empty] | 0.36912 |  | 0.73404
2014/03/16 23:07:18 | 69 | 01-UWSIF-Liver 20140020.26 | 01-UWSIF-Liver |  | 0.7337 | 4 |  |  | 2031 | -3.891 | 47.8358 | 51.337 |  |  |  |  | 50.528 | 0.9 | 1.1 | 3.1 | 1.1014 |  | 1.1732
2014/03/16 23:07:18 | 69 | 01-UWSIF-Liver 20140020.26 | 01-UWSIF-Liver |  | 0.7337 | 5 |  |  | 3222 | 0 | 59.3875 | 63.734 |  |  |  |  | 62.731 | 2.5 | 2.9 | 5.4 | 1.1057 |  | 1.1769
2014/03/16 23:15:48 | 70 | 20140020.049 | 1410 |  | 0.768 | 1 | 2900 | 0.033 |  |  | 10.0064 | 57.172 | 56.737 | 13.1 | 10.1 | 224.7 | [5 empty] | 0.36648 |  | 0.72885
2014/03/16 23:15:48 | 70 | 20140020.049 | 1410 |  | 0.768 | 2 | 2897 | 0 |  |  | 10.0057 | 57.168 | 56.733 | 13.1 | 10.0 | 231.4 | [5 empty] | 0.36647 |  | 0.72883
2014/03/16 23:15:48 | 70 | 20140020.049 | 1410 |  | 0.768 | 3 | 2928 | 13.645 |  |  | 11.9036 | 68.012 | 67.484 | 13.2 | 10.0 | 219.4 | [5 empty] | 0.37145 |  | 0.73877
2014/03/16 23:15:48 | 70 | 20140020.049 | 1410 |  | 0.768 | 4 |  |  | 2067 | -15.012 | 46.902 | 52.688 |  |  |  |  | 51.864 | 0.9 | 1.1 | 3.1 | 1.0892 |  | 1.1609
2014/03/16 23:15:48 | 70 | 20140020.049 | 1410 |  | 0.768 | 5 |  |  | 3220 | 0 | 56.7143 | 63.711 |  |  |  |  | 62.708 | 2.5 | 3.0 | 5.5 | 1.1057 |  | 1.1769
2014/03/16 23:24:17 | 71 | 20140020.050 | 1605 |  | 0.766 | 1 | 2901 | 0.061 |  |  | 10.0431 | 57.232 | 56.797 | 13.2 | 10.1 | 226.1 | [5 empty] | 0.36649 |  | 0.7288
2014/03/16 23:24:17 | 71 | 20140020.050 | 1605 |  | 0.766 | 2 | 2900 | 0 |  |  | 10.0451 | 57.243 | 56.808 | 13.1 | 10.0 | 232.6 | [5 empty] | 0.36647 |  | 0.72875
2014/03/16 23:24:17 | 71 | 20140020.050 | 1605 |  | 0.766 | 3 | 2638 | 13.09 |  |  | 10.7667 | 61.355 | 60.88 | 13.2 | 10.1 | 221.0 | [5 empty] | 0.37125 |  | 0.73829
2014/03/16 23:24:17 | 71 | 20140020.050 | 1605 |  | 0.766 | 4 |  |  | 1983 | -15.41 | 45.0441 | 50.469 |  |  |  |  | 49.68 | 0.9 | 1.1 | 3.1 | 1.0888 |  | 1.1605
2014/03/16 23:24:17 | 71 | 20140020.050 | 1605 |  | 0.766 | 5 |  |  | 3232 | 0 | 56.9879 | 63.852 |  |  |  |  | 62.847 | 2.5 | 2.9 | 5.5 | 1.1057 |  | 1.1768
2014/03/16 23:32:47 | 72 | 20140020.051 | 198 |  | 0.843 | 1 | 2912 | 0.008 |  |  | 9.1588 | 57.44 | 57.003 | 13.2 | 10.1 | 227.8 | [5 empty] | 0.36647 |  | 0.72882
2014/03/16 23:32:47 | 72 | 20140020.051 | 198 |  | 0.843 | 2 | 2911 | 0 |  |  | 9.1533 | 57.405 | 56.968 | 13.1 | 10.0 | 234.4 | [5 empty] | 0.36647 |  | 0.72881
2014/03/16 23:32:47 | 72 | 20140020.051 | 198 |  | 0.843 | 3 | 2612 | 13.102 |  |  | 9.6791 | 60.703 | 60.232 | 13.3 | 10.1 | 222.2 | [5 empty] | 0.37126 |  | 0.73836
2014/03/16 23:32:47 | 72 | 20140020.051 | 198 |  | 0.843 | 4 |  |  | 2244 | -16.278 | 46.6047 | 57.467 |  |  |  |  | 56.569 | 0.9 | 1.1 | 3.1 | 1.0879 |  | 1.1595
2014/03/16 23:32:47 | 72 | 20140020.051 | 198 |  | 0.843 | 5 |  |  | 3241 | 0 | 52.0496 | 64.181 |  |  |  |  | 63.171 | 2.6 | 3.1 | 5.6 | 1.1057 |  | 1.1769
2014/03/16 23:41:17 | 73 | 20140020.052 | 131 |  | 0.813 | 1 | 2912 | 0.033 |  |  | 9.4954 | 57.431 | 56.994 | 13.3 | 10.2 | 228.1 | [5 empty] | 0.36648 |  | 0.72884
2014/03/16 23:41:17 | 73 | 20140020.052 | 131 |  | 0.813 | 2 | 2914 | 0 |  |  | 9.5049 | 57.488 | 57.051 | 13.2 | 10.1 | 235.0 | [5 empty] | 0.36647 |  | 0.72881
2014/03/16 23:41:17 | 73 | 20140020.052 | 131 |  | 0.813 | 3 | 2626 | 14.208 |  |  | 10.0769 | 60.948 | 60.476 | 13.3 | 10.1 | 222.5 | [5 empty] | 0.37166 |  | 0.73917
2014/03/16 23:41:17 | 73 | 20140020.052 | 131 |  | 0.813 | 4 |  |  | 2285 | -16.025 | 49.2163 | 58.527 |  |  |  |  | 57.613 | 0.9 | 1.1 | 3.1 | 1.0881 |  | 1.1598
2014/03/16 23:41:17 | 73 | 20140020.052 | 131 |  | 0.813 | 5 |  |  | 3231 | 0 | 53.7819 | 63.957 |  |  |  |  | 62.95 | 2.6 | 3.1 | 5.6 | 1.1057 |  | 1.1768
2014/03/16 23:49:47 | 74 | 20140020.053 | 217 |  | 0.743 | 1 | 2911 | 0.033 |  |  | 10.3893 | 57.428 | 56.991 | 13.2 | 10.2 | 228.3 | [5 empty] | 0.36648 |  | 0.72881
2014/03/16 23:49:47 | 74 | 20140020.053 | 217 |  | 0.743 | 2 | 2918 | 0 |  |  | 10.402 | 57.497 | 57.06 | 13.2 | 10.1 | 235.6 | [5 empty] | 0.36647 |  | 0.72879
2014/03/16 23:49:47 | 74 | 20140020.053 | 217 |  | 0.743 | 3 | 2494 | 13.599 |  |  | 10.4486 | 57.755 | 57.307 | 13.3 | 10.1 | 223.4 | [5 empty] | 0.37144 |  | 0.7387
2014/03/16 23:49:47 | 74 | 20140020.053 | 217 |  | 0.743 | 4 |  |  | 2088 | -16.508 | 48.7512 | 52.983 |  |  |  |  | 52.155 | 0.9 | 1.1 | 3.2 | 1.0876 |  | 1.1592
2014/03/16 23:49:47 | 74 | 20140020.053 | 217 |  | 0.743 | 5 |  |  | 3265 | 0 | 59.3752 | 64.529 |  |  |  |  | 63.513 | 2.6 | 3.1 | 5.6 | 1.1057 |  | 1.1768
2014/03/16 23:58:17 | 75 | 20140020.054 | 275 |  | 0.796 | 1 | 2931 | 0.018 |  |  | 9.776 | 57.892 | 57.451 | 13.3 | 10.3 | 231.5 | [5 empty] | 0.36648 |  | 0.72883
2014/03/16 23:58:17 | 75 | 20140020.054 | 275 |  | 0.796 | 2 | 2935 | 0 |  |  | 9.7713 | 57.864 | 57.424 | 13.3 | 10.1 | 238.0 | [5 empty] | 0.36647 |  | 0.72882
2014/03/16 23:58:17 | 75 | 20140020.054 | 275 |  | 0.796 | 3 | 2758 | 13.836 |  |  | 10.7894 | 63.893 | 63.398 | 13.4 | 10.2 | 226.6 | [5 empty] | 0.37152 |  | 0.7389
2014/03/16 23:58:17 | 75 | 20140020.054 | 275 |  | 0.796 | 4 |  |  | 2159 | -14.787 | 47.3034 | 55.076 |  |  |  |  | 54.215 | 0.9 | 1.1 | 3.1 | 1.0895 |  | 1.1611
2014/03/16 23:58:17 | 75 | 20140020.054 | 275 |  | 0.796 | 5 |  |  | 3269 | 0 | 55.5579 | 64.687 |  |  |  |  | 63.669 | 2.6 | 3.1 | 5.6 | 1.1057 |  | 1.1768
2014/03/17 00:06:47 | 76 | 20140020.055 | 392 |  | 0.808 | 1 | 2939 | -0.028 |  |  | 9.6472 | 57.99 | 57.549 | 13.4 | 10.3 | 232.4 | [5 empty] | 0.36646 |  | 0.72879
2014/03/17 00:06:47 | 76 | 20140020.055 | 392 |  | 0.808 | 2 | 2938 | 0 |  |  | 9.6395 | 57.944 | 57.503 | 13.3 | 10.2 | 239.0 | [5 empty] | 0.36647 |  | 0.72881
2014/03/17 00:06:47 | 76 | 20140020.055 | 392 |  | 0.808 | 3 | 3376 | 13.747 |  |  | 13.056 | 78.481 | 77.871 | 13.4 | 10.2 | 226.9 | [5 empty] | 0.37149 |  | 0.73883
2014/03/17 00:06:47 | 76 | 20140020.055 | 392 |  | 0.808 | 4 |  |  | 2116 | -13.758 | 45.7608 | 54.084 |  |  |  |  | 53.237 | 0.9 | 1.1 | 3.2 | 1.0906 |  | 1.1623
2014/03/17 00:06:47 | 76 | 20140020.055 | 392 |  | 0.808 | 5 |  |  | 3260 | 0 | 54.5652 | 64.489 |  |  |  |  | 63.474 | 2.6 | 3.1 | 5.6 | 1.1057 |  | 1.1768
2014/03/17 00:15:16 | 77 | 20140020.056 | 372 |  | 0.78 | 1 | 2933 | -0.004 |  |  | 9.9812 | 57.919 | 57.478 | 13.4 | 10.3 | 233.4 | [5 empty] | 0.36647 |  | 0.7288
2014/03/17 00:15:16 | 77 | 20140020.056 | 372 |  | 0.78 | 2 | 2933 | 0 |  |  | 9.9671 | 57.837 | 57.397 | 13.3 | 10.2 | 239.4 | [5 empty] | 0.36647 |  | 0.72881
2014/03/17 00:15:16 | 77 | 20140020.056 | 372 |  | 0.78 | 3 | 3068 | 14.51 |  |  | 12.3161 | 71.468 | 70.913 | 13.4 | 10.2 | 227.5 | [5 empty] | 0.37177 |  | 0.73938
2014/03/17 00:15:16 | 77 | 20140020.056 | 372 |  | 0.78 | 4 |  |  | 2072 | -13.972 | 46.4376 | 52.982 |  |  |  |  | 52.153 | 0.9 | 1.1 | 3.2 | 1.0904 |  | 1.162
2014/03/17 00:15:16 | 77 | 20140020.056 | 372 |  | 0.78 | 5 |  |  | 3254 | 0 | 56.4751 | 64.434 |  |  |  |  | 63.42 | 2.6 | 3.1 | 5.5 | 1.1057 |  | 1.1768
2014/03/17 00:23:46 | 78 | 20140020.057 | 459 |  | 0.845 | 1 | 2927 | 0.027 |  |  | 9.1776 | 57.694 | 57.255 | 13.3 | 10.2 | 232.7 | [5 empty] | 0.36648 |  | 0.72879
2014/03/17 00:23:46 | 78 | 20140020.057 | 459 |  | 0.845 | 2 | 2927 | 0 |  |  | 9.1765 | 57.687 | 57.248 | 13.3 | 10.2 | 239.2 | [5 empty] | 0.36647 |  | 0.72878
2014/03/17 00:23:46 | 78 | 20140020.057 | 459 |  | 0.845 | 3 | 3163 | 13.38 |  |  | 11.6543 | 73.263 | 72.695 | 13.4 | 10.2 | 227.0 | [5 empty] | 0.37136 |  | 0.73853
2014/03/17 00:23:46 | 78 | 20140020.057 | 459 |  | 0.845 | 4 |  |  | 2084 | -13.998 | 43.0097 | 53.16 |  |  |  |  | 52.328 | 0.9 | 1.1 | 3.2 | 1.0903 |  | 1.162
2014/03/17 00:23:46 | 78 | 20140020.057 | 459 |  | 0.845 | 5 |  |  | 3252 | 0 | 52.0288 | 64.307 |  |  |  |  | 63.295 | 2.6 | 3.1 | 5.6 | 1.1057 |  | 1.1768
2014/03/17 00:32:16 | 79 | 20140020.058 | 491 |  | 0.807 | 1 | 2922 | 0.061 |  |  | 9.5983 | 57.625 | 57.186 | 13.3 | 10.3 | 232.9 | [5 empty] | 0.36649 |  | 0.72885
2014/03/17 00:32:16 | 79 | 20140020.058 | 491 |  | 0.807 | 2 | 2921 | 0 |  |  | 9.5849 | 57.545 | 57.107 | 13.3 | 10.1 | 239.4 | [5 empty] | 0.36647 |  | 0.7288
2014/03/17 00:32:16 | 79 | 20140020.058 | 491 |  | 0.807 | 3 | 2931 | 13.06 |  |  | 11.2719 | 67.673 | 67.148 | 13.4 | 10.2 | 226.8 | [5 empty] | 0.37124 |  | 0.73832
2014/03/17 00:32:16 | 79 | 20140020.058 | 491 |  | 0.807 | 4 |  |  | 1921 | -13.586 | 41.0676 | 48.477 |  |  |  |  | 47.718 | 0.9 | 1.1 | 3.2 | 1.0908 |  | 1.1625
2014/03/17 00:32:16 | 79 | 20140020.058 | 491 |  | 0.807 | 5 |  |  | 3246 | 0 | 54.4443 | 64.267 |  |  |  |  | 63.255 | 2.5 | 3.0 | 5.4 | 1.1057 |  | 1.1768
2014/03/17 00:40:45 | 80 | 20140020.059 | 510 |  | 0.81 | 1 | 2916 | 0.008 |  |  | 9.5499 | 57.548 | 57.11 | 13.3 | 10.2 | 232.3 | [5 empty] | 0.36647 |  | 0.72883
2014/03/17 00:40:45 | 80 | 20140020.059 | 510 |  | 0.81 | 2 | 2915 | 0 |  |  | 9.5323 | 57.442 | 57.005 | 13.2 | 10.1 | 238.7 | [5 empty] | 0.36647 |  | 0.72882
2014/03/17 00:40:45 | 80 | 20140020.059 | 510 |  | 0.81 | 3 | 3275 | 14.193 |  |  | 12.5931 | 75.886 | 75.296 | 13.3 | 10.1 | 226.9 | [5 empty] | 0.37165 |  | 0.73917
2014/03/17 00:40:45 | 80 | 20140020.059 | 510 |  | 0.81 | 4 |  |  | 2135 | -14.143 | 45.8672 | 54.344 |  |  |  |  | 53.493 | 0.9 | 1.1 | 3.1 | 1.0902 |  | 1.1618
2014/03/17 00:40:45 | 80 | 20140020.059 | 510 |  | 0.81 | 5 |  |  | 3245 | 0 | 54.2075 | 64.225 |  |  |  |  | 63.214 | 2.6 | 3.1 | 5.5 | 1.1057 |  | 1.1768
2014/03/17 00:49:15 | 81 | 20140020.060 | 533 |  | 0.802 | 1 | 2919 | 0.039 |  |  | 9.659 | 57.63 | 57.192 | 13.4 | 10.3 | 234.2 | [5 empty] | 0.36649 |  | 0.7288
2014/03/17 00:49:15 | 81 | 20140020.060 | 533 |  | 0.802 | 2 | 2920 | 0 |  |  | 9.6438 | 57.539 | 57.101 | 13.3 | 10.1 | 240.5 | [5 empty] | 0.36647 |  | 0.72877
2014/03/17 00:49:15 | 81 | 20140020.060 | 533 |  | 0.802 | 3 | 3194 | 14.76 |  |  | 12.3404 | 73.628 | 73.056 | 13.3 | 10.2 | 227.8 | [5 empty] | 0.37186 |  | 0.73953
2014/03/17 00:49:15 | 81 | 20140020.060 | 533 |  | 0.802 | 4 |  |  | 2113 | -14.715 | 45.636 | 53.536 |  |  |  |  | 52.698 | 0.9 | 1.1 | 3.1 | 1.0896 |  | 1.1612
2014/03/17 00:49:15 | 81 | 20140020.060 | 533 |  | 0.802 | 5 |  |  | 3241 | 0 | 54.7244 | 64.197 |  |  |  |  | 63.187 | 2.6 | 3.1 | 5.6 | 1.1057 |  | 1.1768
2014/03/17 00:57:44 | 82 | 20140020.061 | 930 |  | 0.792 | 1 | 2921 | -0.004 |  |  | 9.782 | 57.636 | 57.198 | 13.4 | 10.3 | 234.5 | [5 empty] | 0.36647 |  | 0.72879
2014/03/17 00:57:44 | 82 | 20140020.061 | 930 |  | 0.792 | 2 | 2922 | 0 |  |  | 9.7774 | 57.609 | 57.171 | 13.3 | 10.2 | 241.4 | [5 empty] | 0.36647 |  | 0.72879
2014/03/17 00:57:44 | 82 | 20140020.061 | 930 |  | 0.792 | 3 | 3017 | 13.563 |  |  | 11.923 | 70.251 | 69.706 | 13.4 | 10.2 | 229.4 | [5 empty] | 0.37142 |  | 0.73868
2014/03/17 00:57:44 | 82 | 20140020.061 | 930 |  | 0.792 | 4 |  |  | 2046 | -14.306 | 45.2331 | 52.401 |  |  |  |  | 51.582 | 0.9 | 1.1 | 3.1 | 1.09 |  | 1.1617
2014/03/17 00:57:44 | 82 | 20140020.061 | 930 |  | 0.792 | 5 |  |  | 3247 | 0 | 55.4969 | 64.292 |  |  |  |  | 63.28 | 2.6 | 3.1 | 5.5 | 1.1057 |  | 1.1768
2014/03/17 01:06:14 | 83 | 20140020.062 | 902 |  | 0.753 | 1 | 2922 | 0.004 |  |  | 10.2822 | 57.6 | 57.162 | 13.4 | 10.3 | 235.1 | [5 empty] | 0.36647 |  | 0.72881
2014/03/17 01:06:14 | 83 | 20140020.062 | 902 |  | 0.753 | 2 | 2923 | 0 |  |  | 10.2758 | 57.564 | 57.126 | 13.3 | 10.1 | 241.6 | [5 empty] | 0.36647 |  | 0.72881
2014/03/17 01:06:14 | 83 | 20140020.062 | 902 |  | 0.753 | 3 | 2350 | 14.349 |  |  | 9.7718 | 54.741 | 54.317 | 13.4 | 10.2 | 229.3 | [5 empty] | 0.37171 |  | 0.73927
2014/03/17 01:06:14 | 83 | 20140020.062 | 902 |  | 0.753 | 4 |  |  | 1766 | -14.878 | 40.7466 | 44.879 |  |  |  |  | 44.178 | 0.9 | 1.1 | 3.2 | 1.0894 |  | 1.161
2014/03/17 01:06:14 | 83 | 20140020.062 | 902 |  | 0.753 | 5 |  |  | 3255 | 0 | 58.4858 | 64.418 |  |  |  |  | 63.404 | 2.4 | 2.9 | 5.3 | 1.1057 |  | 1.1768
2014/03/17 01:14:44 | 84 | 20140020.063 | 841 |  | 0.777 | 1 | 2930 | 0.04 |  |  | 9.9954 | 57.778 | 57.338 | 13.4 | 10.3 | 236.1 | [5 empty] | 0.36649 |  | 0.72882
2014/03/17 01:14:44 | 84 | 20140020.063 | 841 |  | 0.777 | 2 | 2929 | 0 |  |  | 9.9864 | 57.726 | 57.287 | 13.3 | 10.1 | 242.2 | [5 empty] | 0.36647 |  | 0.72879
2014/03/17 01:14:44 | 84 | 20140020.063 | 841 |  | 0.777 | 3 | 3135 | 14.545 |  |  | 12.5816 | 72.728 | 72.163 | 13.4 | 10.2 | 229.7 | [5 empty] | 0.37178 |  | 0.73939
2014/03/17 01:14:44 | 84 | 20140020.063 | 841 |  | 0.777 | 4 |  |  | 2026 | -14.358 | 45.3687 | 51.563 |  |  |  |  | 50.756 | 0.9 | 1.1 | 3.1 | 1.09 |  | 1.1616
2014/03/17 01:14:44 | 84 | 20140020.063 | 841 |  | 0.777 | 5 |  |  | 3258 | 0 | 56.7376 | 64.484 |  |  |  |  | 63.469 | 2.6 | 3.0 | 5.5 | 1.1057 |  | 1.1768
2014/03/17 01:23:14 | 85 | 20140020.064 | 871 |  | 0.819 | 1 | 2930 | -0.01 |  |  | 9.4883 | 57.812 | 57.372 | 13.4 | 10.3 | 236.8 | [5 empty] | 0.36647 |  | 0.72878
2014/03/17 01:23:14 | 85 | 20140020.064 | 871 |  | 0.819 | 2 | 2925 | 0 |  |  | 9.4621 | 57.652 | 57.213 | 13.3 | 10.2 | 243.0 | [5 empty] | 0.36647 |  | 0.72879
2014/03/17 01:23:14 | 85 | 20140020.064 | 871 |  | 0.819 | 3 | 2903 | 14.271 |  |  | 10.9389 | 66.65 | 66.133 | 13.4 | 10.2 | 230.1 | [5 empty] | 0.37168 |  | 0.73919
2014/03/17 01:23:14 | 85 | 20140020.064 | 871 |  | 0.819 | 4 |  |  | 2338 | -15.156 | 49.4456 | 59.234 |  |  |  |  | 58.308 | 0.9 | 1.1 | 3.2 | 1.0891 |  | 1.1607
2014/03/17 01:23:14 | 85 | 20140020.064 | 871 |  | 0.819 | 5 |  |  | 3259 | 0 | 53.8706 | 64.535 |  |  |  |  | 63.52 | 2.7 | 3.2 | 5.8 | 1.1057 |  | 1.1768
2014/03/17 01:31:44 | 86 | 20140020.065 | 796 |  | 0.841 | 1 | 2931 | 0.03 |  |  | 9.2575 | 57.92 | 57.479 | 13.5 | 10.4 | 237.9 | [5 empty] | 0.36648 |  | 0.72881
2014/03/17 01:31:44 | 86 | 20140020.065 | 796 |  | 0.841 | 2 | 2931 | 0 |  |  | 9.2366 | 57.79 | 57.35 | 13.4 | 10.2 | 244.2 | [5 empty] | 0.36647 |  | 0.72879
2014/03/17 01:31:44 | 86 | 20140020.065 | 796 |  | 0.841 | 3 | 2992 | 14.379 |  |  | 11.1377 | 69.684 | 69.143 | 13.5 | 10.3 | 232.2 | [5 empty] | 0.37172 |  | 0.73927
2014/03/17 01:31:44 | 86 | 20140020.065 | 796 |  | 0.841 | 4 |  |  | 2069 | -14.346 | 43.0031 | 52.9 |  |  |  |  | 52.073 | 0.9 | 1.1 | 3.2 | 1.09 |  | 1.1616
2014/03/17 01:31:44 | 86 | 20140020.065 | 796 |  | 0.841 | 5 |  |  | 3252 | 0 | 52.2677 | 64.297 |  |  |  |  | 63.285 | 2.6 | 3.1 | 5.6 | 1.1057 |  | 1.1768
2014/03/17 01:40:14 | 87 | 20140020.066 | 712 |  | 0.811 | 1 | 2925 | 0.058 |  |  | 9.5685 | 57.731 | 57.291 | 13.4 | 10.3 | 237.7 | [5 empty] | 0.36649 |  | 0.72882
2014/03/17 01:40:14 | 87 | 20140020.066 | 712 |  | 0.811 | 2 | 2926 | 0 |  |  | 9.5685 | 57.731 | 57.291 | 13.3 | 10.2 | 244.2 | [5 empty] | 0.36647 |  | 0.72878
2014/03/17 01:40:14 | 87 | 20140020.066 | 712 |  | 0.811 | 3 | 2945 | 14.485 |  |  | 11.3154 | 68.271 | 67.74 | 13.4 | 10.2 | 232.0 | [5 empty] | 0.37176 |  | 0.73933
2014/03/17 01:40:14 | 87 | 20140020.066 | 712 |  | 0.811 | 4 |  |  | 2160 | -14.982 | 46.3813 | 55.02 |  |  |  |  | 54.16 | 0.9 | 1.1 | 3.2 | 1.0893 |  | 1.1609
2014/03/17 01:40:14 | 87 | 20140020.066 | 712 |  | 0.811 | 5 |  |  | 3260 | 0 | 54.387 | 64.517 |  |  |  |  | 63.502 | 2.7 | 3.1 | 5.6 | 1.1057 |  | 1.1768
2014/03/17 01:48:44 | 88 | 20140020.067 | 621 |  | 0.749 | 1 | 2927 | 0.002 |  |  | 10.3658 | 57.76 | 57.321 | 13.5 | 10.3 | 238.1 | [5 empty] | 0.36647 |  | 0.7288
2014/03/17 01:48:44 | 88 | 20140020.067 | 621 |  | 0.749 | 2 | 2931 | 0 |  |  | 10.374 | 57.806 | 57.366 | 13.4 | 10.2 | 244.8 | [5 empty] | 0.36647 |  | 0.7288
2014/03/17 01:48:44 | 88 | 20140020.067 | 621 |  | 0.749 | 3 | 2867 | 14.757 |  |  | 11.9985 | 66.858 | 66.339 | 13.4 | 10.3 | 232.8 | [5 empty] | 0.37186 |  | 0.73956
2014/03/17 01:48:44 | 88 | 20140020.067 | 621 |  | 0.749 | 4 |  |  | 1948 | -14.959 | 45.3621 | 49.698 |  |  |  |  | 48.921 | 0.9 | 1.1 | 3.1 | 1.0893 |  | 1.1609
2014/03/17 01:48:44 | 88 | 20140020.067 | 621 |  | 0.749 | 5 |  |  | 3259 | 0 | 58.8681 | 64.494 |  |  |  |  | 63.48 | 2.5 | 3.0 | 5.5 | 1.1057 |  | 1.1768
2014/03/17 01:57:14 | 89 | 20140020.068 | 693 |  | 0.834 | 1 | 2928 | 0.009 |  |  | 9.3162 | 57.803 | 57.363 | 13.4 | 10.3 | 239.1 | [5 empty] | 0.36647 |  | 0.72878
2014/03/17 01:57:14 | 89 | 20140020.068 | 693 |  | 0.834 | 2 | 2928 | 0 |  |  | 9.3001 | 57.703 | 57.264 | 13.3 | 10.2 | 245.2 | [5 empty] | 0.36647 |  | 0.72877
2014/03/17 01:57:14 | 89 | 20140020.068 | 693 |  | 0.834 | 3 | 3129 | 14.829 |  |  | 11.6949 | 72.562 | 71.998 | 13.4 | 10.2 | 233.1 | [5 empty] | 0.37189 |  | 0.73958
2014/03/17 01:57:14 | 89 | 20140020.068 | 693 |  | 0.834 | 4 |  |  | 2259 | -15.123 | 47.4505 | 57.885 |  |  |  |  | 56.98 | 0.9 | 1.1 | 3.1 | 1.0891 |  | 1.1607
2014/03/17 01:57:14 | 89 | 20140020.068 | 693 |  | 0.834 | 5 |  |  | 3259 | 0 | 52.8792 | 64.508 |  |  |  |  | 63.493 | 2.7 | 3.2 | 5.7 | 1.1057 |  | 1.1768
2014/03/17 02:05:43 | 90 | 20140020.069 | 1095 |  | 0.786 | 1 | 2926 | 0.01 |  |  | 9.8741 | 57.738 | 57.299 | 13.5 | 10.4 | 239.2 | [5 empty] | 0.36648 |  | 0.72876
2014/03/17 02:05:43 | 90 | 20140020.069 | 1095 |  | 0.786 | 2 | 2925 | 0 |  |  | 9.8818 | 57.783 | 57.343 | 13.4 | 10.2 | 245.9 | [5 empty] | 0.36647 |  | 0.72876
2014/03/17 02:05:43 | 90 | 20140020.069 | 1095 |  | 0.786 | 3 | 3064 | 14.773 |  |  | 12.1314 | 70.938 | 70.387 | 13.4 | 10.2 | 233.6 | [5 empty] | 0.37187 |  | 0.73952
2014/03/17 02:05:43 | 90 | 20140020.069 | 1095 |  | 0.786 | 4 |  |  | 2067 | -15.148 | 45.6232 | 52.453 |  |  |  |  | 51.633 | 0.9 | 1.1 | 3.1 | 1.0891 |  | 1.1607
2014/03/17 02:05:43 | 90 | 20140020.069 | 1095 |  | 0.786 | 5 |  |  | 3257 | 0 | 56.0827 | 64.478 |  |  |  |  | 63.463 | 2.6 | 3.1 | 5.6 | 1.1057 |  | 1.1768
2014/03/17 02:14:13 | 91 | 20140020.070 | 1068 |  | 0.814 | 1 | 2925 | 0.054 |  |  | 9.5262 | 57.688 | 57.249 | 13.4 | 10.3 | 239.7 | [5 empty] | 0.36649 |  | 0.7288
2014/03/17 02:14:13 | 91 | 20140020.070 | 1068 |  | 0.814 | 2 | 2924 | 0 |  |  | 9.5257 | 57.685 | 57.246 | 13.3 | 10.2 | 246.4 | [5 empty] | 0.36647 |  | 0.72876
2014/03/17 02:14:13 | 91 | 20140020.070 | 1068 |  | 0.814 | 3 | 2905 | 14.739 |  |  | 11.0958 | 67.193 | 66.671 | 13.5 | 10.3 | 233.4 | [5 empty] | 0.37185 |  | 0.7395
2014/03/17 02:14:13 | 91 | 20140020.070 | 1068 |  | 0.814 | 4 |  |  | 2198 | -15.202 | 47.1242 | 56.108 |  |  |  |  | 55.231 | 0.9 | 1.1 | 3.2 | 1.089 |  | 1.1606
2014/03/17 02:14:13 | 91 | 20140020.070 | 1068 |  | 0.814 | 5 |  |  | 3253 | 0 | 54.0812 | 64.392 |  |  |  |  | 63.379 | 2.7 | 3.2 | 5.6 | 1.1057 |  | 1.1767
2014/03/17 02:22:43 | 92 | 20140020.071 | 1153 |  | 0.836 | 1 | 2919 | 0.011 |  |  | 9.2584 | 57.582 | 57.144 | 13.4 | 10.3 | 238.9 | [5 empty] | 0.36648 |  | 0.72878
2014/03/17 02:22:43 | 92 | 20140020.071 | 1153 |  | 0.836 | 2 | 2920 | 0 |  |  | 9.2485 | 57.52 | 57.083 | 13.3 | 10.2 | 245.6 | [5 empty] | 0.36647 |  | 0.72877
2014/03/17 02:22:43 | 92 | 20140020.071 | 1153 |  | 0.836 | 3 | 3137 | 12.912 |  |  | 11.7113 | 72.837 | 72.272 | 13.4 | 10.2 | 232.9 | [5 empty] | 0.37119 |  | 0.73818
2014/03/17 02:22:43 | 92 | 20140020.071 | 1153 |  | 0.836 | 4 |  |  | 2065 | -14.09 | 43.0499 | 52.643 |  |  |  |  | 51.819 | 0.9 | 1.1 | 3.2 | 1.0902 |  | 1.1618
2014/03/17 02:22:43 | 92 | 20140020.071 | 1153 |  | 0.836 | 5 |  |  | 3244 | 0 | 52.4976 | 64.196 |  |  |  |  | 63.186 | 2.6 | 3.1 | 5.6 | 1.1057 |  | 1.1768
2014/03/17 02:31:14 | 93 | 20140020.072 | 1131 |  | 0.827 | 1 | 2913 | 0.022 |  |  | 9.3348 | 57.432 | 56.995 | 13.4 | 10.3 | 238.6 | [5 empty] | 0.36648 |  | 0.72879
2014/03/17 02:31:14 | 93 | 20140020.072 | 1131 |  | 0.827 | 2 | 2912 | 0 |  |  | 9.33 | 57.402 | 56.965 | 13.3 | 10.2 | 244.9 | [5 empty] | 0.36647 |  | 0.72877
2014/03/17 02:31:14 | 93 | 20140020.072 | 1131 |  | 0.827 | 3 | 3051 | 13 |  |  | 11.5218 | 70.888 | 70.338 | 13.4 | 10.2 | 232.7 | [5 empty] | 0.37122 |  | 0.73825
2014/03/17 02:31:14 | 93 | 20140020.072 | 1131 |  | 0.827 | 4 |  |  | 2095 | -14.162 | 44.2307 | 53.504 |  |  |  |  | 52.667 | 0.9 | 1.1 | 3.1 | 1.0902 |  | 1.1618
2014/03/17 02:31:14 | 93 | 20140020.072 | 1131 |  | 0.827 | 5 |  |  | 3244 | 0 | 53.0146 | 64.13 |  |  |  |  | 63.121 | 2.6 | 3.1 | 5.5 | 1.1057 |  | 1.1768
2014/03/17 02:39:44 | 94 | 36-UWSIF-UT Glut 120140020.211 | 36-UWSIF-UT Glut 1 |  | 0.7732 | 1 | 2916 | 0.002 |  |  | 9.999 | 57.517 | 57.079 | 13.4 | 10.3 | 239.7 | [5 empty] | 0.36647 |  | 0.72879
2014/03/17 02:39:44 | 94 | 36-UWSIF-UT Glut 120140020.211 | 36-UWSIF-UT Glut 1 |  | 0.7732 | 2 | 2915 | 0 |  |  | 9.9967 | 57.503 | 57.065 | 13.3 | 10.2 | 246.4 | [5 empty] | 0.36647 |  | 0.72878
2014/03/17 02:39:44 | 94 | 36-UWSIF-UT Glut 120140020.211 | 36-UWSIF-UT Glut 1 |  | 0.7732 | 3 | 2269 | -4.22 |  |  | 9.1336 | 52.539 | 52.138 | 13.4 | 10.2 | 233.2 | [5 empty] | 0.36493 |  | 0.72571
2014/03/17 02:39:44 | 94 | 36-UWSIF-UT Glut 120140020.211 | 36-UWSIF-UT Glut 1 |  | 0.7732 | 4 |  |  | 1766 | -14.337 | 39.2406 | 44.38 |  |  |  |  | 43.686 | 0.9 | 1.1 | 3.2 | 1.09 |  | 1.1616
2014/03/17 02:39:44 | 94 | 36-UWSIF-UT Glut 120140020.211 | 36-UWSIF-UT Glut 1 |  | 0.7732 | 5 |  |  | 3239 | 0 | 56.6981 | 64.124 |  |  |  |  | 63.115 | 2.4 | 2.9 | 5.3 | 1.1057 |  | 1.1768
2014/03/17 02:48:14 | 95 | 36-UWSIF-UT Glut 120140020.212 | 36-UWSIF-UT Glut 1 |  | 0.7356 | 1 | 2911 | -0.03 |  |  | 10.4951 | 57.434 | 56.997 | 13.4 | 10.3 | 238.0 | [5 empty] | 0.36646 |  | 0.72876
2014/03/17 02:48:14 | 95 | 36-UWSIF-UT Glut 120140020.212 | 36-UWSIF-UT Glut 1 |  | 0.7356 | 2 | 2910 | 0 |  |  | 10.4859 | 57.384 | 56.947 | 13.3 | 10.1 | 244.7 | [5 empty] | 0.36647 |  | 0.72878
2014/03/17 02:48:14 | 95 | 36-UWSIF-UT Glut 120140020.212 | 36-UWSIF-UT Glut 1 |  | 0.7356 | 3 | 2155 | -4.248 |  |  | 9.1169 | 49.892 | 49.512 | 13.3 | 10.1 | 231.7 | [5 empty] | 0.36492 |  | 0.72569
2014/03/17 02:48:14 | 95 | 36-UWSIF-UT Glut 120140020.212 | 36-UWSIF-UT Glut 1 |  | 0.7356 | 4 |  |  | 1681 | -14.322 | 39.1323 | 42.105 |  |  |  |  | 41.447 | 0.9 | 1.1 | 3.1 | 1.09 |  | 1.1616
2014/03/17 02:48:14 | 95 | 36-UWSIF-UT Glut 120140020.212 | 36-UWSIF-UT Glut 1 |  | 0.7356 | 5 |  |  | 3236 | 0 | 59.517 | 64.039 |  |  |  |  | 63.031 | 2.3 | 2.8 | 5.2 | 1.1057 |  | 1.1768
2014/03/17 02:56:45 | 96 | 39-UWSIF-UW Glut 220140020.27 | 39-UWSIF-UW Glut 2 |  | 0.714 | 1 | 2909 | 0.033 |  |  | 10.7967 | 57.35 | 56.913 | 13.3 | 10.2 | 236.6 | [5 empty] | 0.36648 |  | 0.72881
2014/03/17 02:56:45 | 96 | 39-UWSIF-UW Glut 220140020.27 | 39-UWSIF-UW Glut 2 |  | 0.714 | 2 | 2909 | 0 |  |  | 10.7961 | 57.347 | 56.91 | 13.2 | 10.1 | 243.7 | [5 empty] | 0.36647 |  | 0.72879
2014/03/17 02:56:45 | 96 | 39-UWSIF-UW Glut 220140020.27 | 39-UWSIF-UW Glut 2 |  | 0.714 | 3 | 2269 | 28.3 |  |  | 9.9287 | 52.739 | 52.325 | 13.3 | 10.2 | 231.8 | [5 empty] | 0.3768 |  | 0.74941
2014/03/17 02:56:45 | 96 | 39-UWSIF-UW Glut 220140020.27 | 39-UWSIF-UW Glut 2 |  | 0.714 | 4 |  |  | 1761 | 37.545 | 42.5804 | 44.47 |  |  |  |  | 43.75 | 0.9 | 1.1 | 3.1 | 1.1467 |  | 1.2189
2014/03/17 02:56:45 | 96 | 39-UWSIF-UW Glut 220140020.27 | 39-UWSIF-UW Glut 2 |  | 0.714 | 5 |  |  | 3239 | 0 | 61.4013 | 64.126 |  |  |  |  | 63.117 | 2.4 | 2.9 | 5.3 | 1.1057 |  | 1.1771
2014/03/17 03:05:15 | 97 | 39-UWSIF-UW Glut 220140020.28 | 39-UWSIF-UW Glut 2 |  | 0.7284 | 1 | 2911 | 0.023 |  |  | 10.5975 | 57.427 | 56.99 | 13.3 | 10.2 | 237.5 | [5 empty] | 0.36648 |  | 0.72878
2014/03/17 03:05:15 | 97 | 39-UWSIF-UW Glut 220140020.28 | 39-UWSIF-UW Glut 2 |  | 0.7284 | 2 | 2913 | 0 |  |  | 10.6015 | 57.449 | 57.011 | 13.3 | 10.1 | 244.2 | [5 empty] | 0.36647 |  | 0.72877
2014/03/17 03:05:15 | 97 | 39-UWSIF-UW Glut 220140020.28 | 39-UWSIF-UW Glut 2 |  | 0.7284 | 3 | 2315 | 28.343 |  |  | 9.9442 | 53.887 | 53.464 | 13.3 | 10.2 | 232.5 | [5 empty] | 0.37682 |  | 0.74942
2014/03/17 03:05:15 | 97 | 39-UWSIF-UW Glut 220140020.28 | 39-UWSIF-UW Glut 2 |  | 0.7284 | 4 |  |  | 1799 | 37.613 | 42.6414 | 45.432 |  |  |  |  | 44.696 | 0.9 | 1.1 | 3.2 | 1.1468 |  | 1.219
2014/03/17 03:05:15 | 97 | 39-UWSIF-UW Glut 220140020.28 | 39-UWSIF-UW Glut 2 |  | 0.7284 | 5 |  |  | 3236 | 0 | 60.1707 | 64.108 |  |  |  |  | 63.099 | 2.4 | 3.0 | 5.4 | 1.1057 |  | 1.1771
2014/03/17 03:13:46 | 98 | 01-UWSIF-Liver 20140020.27 | 01-UWSIF-Liver |  | 0.777 | 1 | 2913 | 0.011 |  |  | 9.9365 | 57.438 | 57 | 13.3 | 10.3 | 237.7 | [5 empty] | 0.36648 |  | 0.72879
2014/03/17 03:13:46 | 98 | 01-UWSIF-Liver 20140020.27 | 01-UWSIF-Liver |  | 0.777 | 2 | 2912 | 0 |  |  | 9.9394 | 57.455 | 57.017 | 13.2 | 10.2 | 244.5 | [5 empty] | 0.36647 |  | 0.72878
2014/03/17 03:13:46 | 98 | 01-UWSIF-Liver 20140020.27 | 01-UWSIF-Liver |  | 0.777 | 3 | 3137 | 7.22 |  |  | 12.5195 | 72.369 | 71.81 | 13.3 | 10.2 | 232.5 | [5 empty] | 0.36911 |  | 0.73404
2014/03/17 03:13:46 | 98 | 01-UWSIF-Liver 20140020.27 | 01-UWSIF-Liver |  | 0.777 | 4 |  |  | 2169 | -3.883 | 48.2819 | 54.874 |  |  |  |  | 54.009 | 0.9 | 1.1 | 3.1 | 1.1014 |  | 1.1732
2014/03/17 03:13:46 | 98 | 01-UWSIF-Liver 20140020.27 | 01-UWSIF-Liver |  | 0.777 | 5 |  |  | 3250 | 0 | 56.5915 | 64.318 |  |  |  |  | 63.306 | 2.6 | 3.1 | 5.5 | 1.1057 |  | 1.1769
2014/03/17 03:22:16 | 99 | 01-UWSIF-Liver 20140020.28 | 01-UWSIF-Liver |  | 0.7235 | 1 | 2918 | 0.047 |  |  | 10.7057 | 57.623 | 57.185 | 13.4 | 10.3 | 240.0 | [5 empty] | 0.36649 |  | 0.7288
2014/03/17 03:22:16 | 99 | 01-UWSIF-Liver 20140020.28 | 01-UWSIF-Liver |  | 0.7235 | 2 | 2918 | 0 |  |  | 10.6991 | 57.588 | 57.15 | 13.3 | 10.2 | 247.1 | [5 empty] | 0.36647 |  | 0.72877
2014/03/17 03:22:16 | 99 | 01-UWSIF-Liver 20140020.28 | 01-UWSIF-Liver |  | 0.7235 | 3 | 2935 | 7.157 |  |  | 12.5797 | 67.71 | 67.188 | 13.4 | 10.2 | 234.8 | [5 empty] | 0.36908 |  | 0.73399
2014/03/17 03:22:16 | 99 | 01-UWSIF-Liver 20140020.28 | 01-UWSIF-Liver |  | 0.7235 | 4 |  |  | 2031 | -3.947 | 48.3954 | 51.216 |  |  |  |  | 50.409 | 0.9 | 1.2 | 3.2 | 1.1013 |  | 1.1731
2014/03/17 03:22:16 | 99 | 01-UWSIF-Liver 20140020.28 | 01-UWSIF-Liver |  | 0.7235 | 5 |  |  | 3252 | 0 | 60.7846 | 64.327 |  |  |  |  | 63.315 | 2.5 | 3.0 | 5.5 | 1.1057 |  | 1.1769
Job: 2014-0020 Run 1 Stockwell
Analyst: DEw
Sample Count:376
Helium Pressure: 1200
Oxygen pressure: 2400
Mass 28:17
Mass 29:14
Mass 30: 129
Peak Center:3.065
Instrument: ANNIE
Data analyst:
Comments:Sample 41 lost while crimping, picked up sample off the floor, but was unsure if was correct sample. Since # is drastically different from other samples, decided sample was lost.
Some samples had leftovers in sample well, highlighted which weight percents are off.

## Original 2
Time Code | Row | Identifier 1 | Identifier 2 | Comment | Amount | Peak Nr | Ampl 28 | d 15N/14N | Ampl 44 | d 13C/12C | Amt% | Area All | Area 28 | BGD 28 | BGD 29 | BGD 30 | Area 44 | BGD 44 | BGD 45 | BGD 46 | AT% 13C/12C | AT% 15N/14N | rR 45CO2/44CO2 | rR 29N2/28N2
2014/03/24 09:54:35 | 1 | 36-UWSIF-UT Glut 120140020.31 | 36-UWSIF-UT Glut 1 | Job # : 2014-0020 Run 2 Stockwell | 0.72 | 1 | 2631 | 0.013 |  |  | 11.28 | 51.896 | 51.503 | 9.6 | 7.0 | 92.9 | [5 empty] | 0.36648 |  | 0.73111
2014/03/24 09:54:35 | 1 | 36-UWSIF-UT Glut 120140020.31 | 36-UWSIF-UT Glut 1 | Job # : 2014-0020 Run 2 Stockwell | 0.72 | 2 | 2636 | 0 |  |  | 11.2941 | 51.961 | 51.568 | 9.7 | 7.0 | 100.0 | [5 empty] | 0.36647 |  | 0.73111
2014/03/24 09:54:35 | 1 | 36-UWSIF-UT Glut 120140020.31 | 36-UWSIF-UT Glut 1 | Job # : 2014-0020 Run 2 Stockwell | 0.72 | 3 | 1903 | -4.131 |  |  | 9.6909 | 44.585 | 44.247 | 9.7 | 7.1 | 93.5 | [5 empty] | 0.36496 |  | 0.72808
2014/03/24 09:54:35 | 1 | 36-UWSIF-UT Glut 120140020.31 | 36-UWSIF-UT Glut 1 | Job # : 2014-0020 Run 2 Stockwell | 0.72 | 4 |  |  | 1479 | -14.045 | 41.0464 | 37.445 |  |  |  |  | 36.858 | -0.1 | -0.0 | 1.6 | 1.0903 |  | 1.1642
2014/03/24 09:54:35 | 1 | 36-UWSIF-UT Glut 120140020.31 | 36-UWSIF-UT Glut 1 | Job # : 2014-0020 Run 2 Stockwell | 0.72 | 5 |  |  | 2884 | 0 | 62.4397 | 56.96 |  |  |  |  | 56.063 | 1.8 | 2.2 | 4.4 | 1.1057 |  | 1.1792
2014/03/24 10:03:03 | 2 | 36-UWSIF-UT Glut 120140020.32 | 36-UWSIF-UT Glut 1 | Analyst: DEW | 0.852 | 1 | 2638 | 0.016 |  |  |  | 52.063 | 51.669 | 10.1 | 7.6 | 103.7 | [5 empty] | 0.36648 |  | 0.73143
2014/03/24 10:03:03 | 2 | 36-UWSIF-UT Glut 120140020.32 | 36-UWSIF-UT Glut 1 | Analyst: DEW | 0.852 | 2 | 2639 | 0 |  |  |  | 52.11 | 51.716 | 10.1 | 7.5 | 109.1 | [5 empty] | 0.36647 |  | 0.73142
2014/03/24 10:03:03 | 2 | 36-UWSIF-UT Glut 120140020.32 | 36-UWSIF-UT Glut 1 | Analyst: DEW | 0.852 | 3 | 2241 | -4.148 |  |  | 9.52 | 52.133 | 51.738 | 10.1 | 7.5 | 100.9 | [5 empty] | 0.36496 |  | 0.72839
2014/03/24 10:03:03 | 2 | 36-UWSIF-UT Glut 120140020.32 | 36-UWSIF-UT Glut 1 | Analyst: DEW | 0.852 | 4 |  |  | 1746 | -14.133 | 40.81 | 44.554 |  |  |  |  | 43.856 | 0.4 | 0.5 | 2.4 | 1.0902 |  | 1.1639
2014/03/24 10:03:03 | 2 | 36-UWSIF-UT Glut 120140020.32 | 36-UWSIF-UT Glut 1 | Analyst: DEW | 0.852 | 5 |  |  | 2895 | 0 |  | 57.298 |  |  |  |  | 56.395 | 2.2 | 2.6 | 4.9 | 1.1057 |  | 1.179
2014/03/24 10:11:30 | 3 | 36-UWSIF-UT Glut 120140020.33 | 36-UWSIF-UT Glut 1 | Sample Count: 0547 | 0.418 | 1 | 2641 | 0.009 |  |  | 19.3779 | 52.062 | 51.668 | 10.4 | 7.8 | 107.0 | [5 empty] | 0.36647 |  | 0.73145
2014/03/24 10:11:30 | 3 | 36-UWSIF-UT Glut 120140020.33 | 36-UWSIF-UT Glut 1 | Sample Count: 0547 | 0.418 | 2 | 2642 | 0 |  |  | 19.4004 | 52.122 | 51.728 | 10.3 | 7.7 | 112.0 | [5 empty] | 0.36647 |  | 0.73144
2014/03/24 10:11:30 | 3 | 36-UWSIF-UT Glut 120140020.33 | 36-UWSIF-UT Glut 1 | Sample Count: 0547 | 0.418 | 3 | 1082 | -4.225 |  |  | 9.5111 | 25.553 | 25.36 | 10.3 | 7.7 | 103.5 | [5 empty] | 0.36493 |  | 0.72835
2014/03/24 10:11:30 | 3 | 36-UWSIF-UT Glut 120140020.33 | 36-UWSIF-UT Glut 1 | Sample Count: 0547 | 0.418 | 4 |  |  | 859 | -14.037 | 40.3277 | 21.6 |  |  |  |  | 21.262 | 0.5 | 0.7 | 2.5 | 1.0903 |  | 1.164
2014/03/24 10:11:30 | 3 | 36-UWSIF-UT Glut 120140020.33 | 36-UWSIF-UT Glut 1 | Sample Count: 0547 | 0.418 | 5 |  |  | 2883 | 0 | 106.4753 | 57.03 |  |  |  |  | 56.131 | 1.6 | 1.9 | 4.0 | 1.1057 |  | 1.179
2014/03/24 10:19:57 | 4 | 36-UWSIF-UT Glut 120140020.34 | 36-UWSIF-UT Glut 1 | Helium Pressure:1850 | 0.813 | 1 | 2635 | 0.015 |  |  | 9.9371 | 51.926 | 51.533 | 10.3 | 7.7 | 106.4 | [5 empty] | 0.36648 |  | 0.73141
2014/03/24 10:19:57 | 4 | 36-UWSIF-UT Glut 120140020.34 | 36-UWSIF-UT Glut 1 | Helium Pressure:1850 | 0.813 | 2 | 2632 | 0 |  |  | 9.9491 | 51.989 | 51.595 | 10.2 | 7.6 | 111.9 | [5 empty] | 0.36647 |  | 0.7314
2014/03/24 10:19:57 | 4 | 36-UWSIF-UT Glut 120140020.34 | 36-UWSIF-UT Glut 1 | Helium Pressure:1850 | 0.813 | 3 | 2140 | -4.151 |  |  | 9.5108 | 49.699 | 49.323 | 10.3 | 7.7 | 104.0 | [5 empty] | 0.36496 |  | 0.72836
2014/03/24 10:19:57 | 4 | 36-UWSIF-UT Glut 120140020.34 | 36-UWSIF-UT Glut 1 | Helium Pressure:1850 | 0.813 | 4 |  |  | 1669 | -14.155 | 40.693 | 42.392 |  |  |  |  | 41.729 | 0.5 | 0.6 | 2.4 | 1.0902 |  | 1.1636
2014/03/24 10:19:57 | 4 | 36-UWSIF-UT Glut 120140020.34 | 36-UWSIF-UT Glut 1 | Helium Pressure:1850 | 0.813 | 5 |  |  | 2877 | 0 | 54.6093 | 56.89 |  |  |  |  | 55.993 | 2.2 | 2.6 | 4.9 | 1.1057 |  | 1.1788
2014/03/24 10:28:25 | 5 | 36-UWSIF-UT Glut 120140020.35 | 36-UWSIF-UT Glut 1 | Oxygen Pressure:2400 | 1.067 | 1 | 2628 | 0.014 |  |  | 7.5691 | 51.909 | 51.516 | 10.4 | 7.8 | 108.9 | [5 empty] | 0.36648 |  | 0.73138
2014/03/24 10:28:25 | 5 | 36-UWSIF-UT Glut 120140020.35 | 36-UWSIF-UT Glut 1 | Oxygen Pressure:2400 | 1.067 | 2 | 2630 | 0 |  |  | 7.556 | 51.82 | 51.427 | 10.3 | 7.7 | 114.3 | [5 empty] | 0.36647 |  | 0.73137
2014/03/24 10:28:25 | 5 | 36-UWSIF-UT Glut 120140020.35 | 36-UWSIF-UT Glut 1 | Oxygen Pressure:2400 | 1.067 | 3 | 2835 | -4.055 |  |  | 9.5074 | 65.203 | 64.708 | 10.4 | 7.7 | 106.7 | [5 empty] | 0.36499 |  | 0.7284
2014/03/24 10:28:25 | 5 | 36-UWSIF-UT Glut 120140020.35 | 36-UWSIF-UT Glut 1 | Oxygen Pressure:2400 | 1.067 | 4 |  |  | 2182 | -14.247 | 40.7368 | 55.697 |  |  |  |  | 54.825 | 0.5 | 0.7 | 2.6 | 1.0901 |  | 1.1633
2014/03/24 10:28:25 | 5 | 36-UWSIF-UT Glut 120140020.35 | 36-UWSIF-UT Glut 1 | Oxygen Pressure:2400 | 1.067 | 5 |  |  | 2876 | 0 | 41.6031 | 56.881 |  |  |  |  | 55.985 | 2.6 | 3.1 | 5.5 | 1.1057 |  | 1.1786
2014/03/24 10:36:53 | 6 | 39-UWSIF-UW Glut 220140020.31 | 39-UWSIF-UW Glut 2 | Mass 28:10 | 0.7874 | 1 | 2628 | 0.014 |  |  | 10.2351 | 51.799 | 51.407 | 10.6 | 8.0 | 112.3 | [5 empty] | 0.36648 |  | 0.73139
2014/03/24 10:36:53 | 6 | 39-UWSIF-UW Glut 220140020.31 | 39-UWSIF-UW Glut 2 | Mass 28:10 | 0.7874 | 2 | 2624 | 0 |  |  | 10.2297 | 51.772 | 51.38 | 10.4 | 7.9 | 117.4 | [5 empty] | 0.36647 |  | 0.73138
2014/03/24 10:36:53 | 6 | 39-UWSIF-UW Glut 220140020.31 | 39-UWSIF-UW Glut 2 | Mass 28:10 | 0.7874 | 3 | 2235 | 28.531 |  |  | 10.2629 | 51.94 | 51.534 | 10.5 | 7.8 | 109.1 | [5 empty] | 0.37689 |  | 0.75225
2014/03/24 10:36:53 | 6 | 39-UWSIF-UW Glut 220140020.31 | 39-UWSIF-UW Glut 2 | Mass 28:10 | 0.7874 | 4 |  |  | 1739 | 37.875 | 43.9371 | 44.331 |  |  |  |  | 43.612 | 0.6 | 0.8 | 2.7 | 1.1471 |  | 1.2209
2014/03/24 10:36:53 | 6 | 39-UWSIF-UW Glut 220140020.31 | 39-UWSIF-UW Glut 2 | Mass 28:10 | 0.7874 | 5 |  |  | 2868 | 0 | 56.2572 | 56.761 |  |  |  |  | 55.867 | 2.3 | 2.9 | 5.0 | 1.1057 |  | 1.1789
2014/03/24 10:45:21 | 7 | 39-UWSIF-UW Glut 220140020.32 | 39-UWSIF-UW Glut 2 | Mass 29:7 | 0.7924 | 1 | 2621 | 0.004 |  |  | 10.145 | 51.67 | 51.278 | 10.6 | 8.0 | 113.0 | [5 empty] | 0.36647 |  | 0.73142
2014/03/24 10:45:21 | 7 | 39-UWSIF-UW Glut 220140020.32 | 39-UWSIF-UW Glut 2 | Mass 29:7 | 0.7924 | 2 | 2618 | 0 |  |  | 10.1493 | 51.692 | 51.3 | 10.4 | 7.8 | 118.4 | [5 empty] | 0.36647 |  | 0.73142
2014/03/24 10:45:21 | 7 | 39-UWSIF-UW Glut 220140020.32 | 39-UWSIF-UW Glut 2 | Mass 29:7 | 0.7924 | 3 | 2255 | 28.441 |  |  | 10.2695 | 52.304 | 51.895 | 10.5 | 7.8 | 110.5 | [5 empty] | 0.37685 |  | 0.75222
2014/03/24 10:45:21 | 7 | 39-UWSIF-UW Glut 220140020.32 | 39-UWSIF-UW Glut 2 | Mass 29:7 | 0.7924 | 4 |  |  | 1754 | 37.894 | 43.966 | 44.642 |  |  |  |  | 43.918 | 0.6 | 0.8 | 2.6 | 1.1471 |  | 1.2211
2014/03/24 10:45:21 | 7 | 39-UWSIF-UW Glut 220140020.32 | 39-UWSIF-UW Glut 2 | Mass 29:7 | 0.7924 | 5 |  |  | 2860 | 0 | 55.738 | 56.594 |  |  |  |  | 55.702 | 2.3 | 2.9 | 5.1 | 1.1057 |  | 1.179
2014/03/24 10:53:49 | 8 | 01-UWSIF-Liver 20140020.31 | 01-UWSIF-Liver | Mass 30:64 | 0.68 | 1 | 2609 | -0.013 |  |  | 11.7784 | 51.479 | 51.089 | 10.6 | 8.0 | 114.3 | [5 empty] | 0.36647 |  | 0.73143
2014/03/24 10:53:49 | 8 | 01-UWSIF-Liver 20140020.31 | 01-UWSIF-Liver | Mass 30:64 | 0.68 | 2 | 2609 | 0 |  |  | 11.7704 | 51.444 | 51.055 | 10.4 | 7.9 | 119.0 | [5 empty] | 0.36647 |  | 0.73144
2014/03/24 10:53:49 | 8 | 01-UWSIF-Liver 20140020.31 | 01-UWSIF-Liver | Mass 30:64 | 0.68 | 3 | 2420 | 7.275 |  |  | 12.8774 | 56.283 | 55.851 | 10.4 | 7.8 | 111.2 | [5 empty] | 0.36913 |  | 0.73676
2014/03/24 10:53:49 | 8 | 01-UWSIF-Liver 20140020.31 | 01-UWSIF-Liver | Mass 30:64 | 0.68 | 4 |  |  | 1695 | -3.752 | 49.653 | 43.265 |  |  |  |  | 42.583 | 0.6 | 0.8 | 2.6 | 1.1016 |  | 1.1751
2014/03/24 10:53:49 | 8 | 01-UWSIF-Liver 20140020.31 | 01-UWSIF-Liver | Mass 30:64 | 0.68 | 5 |  |  | 2849 | 0 | 64.657 | 56.338 |  |  |  |  | 55.45 | 2.3 | 2.8 | 5.1 | 1.1057 |  | 1.1788
2014/03/24 11:02:17 | 9 | 01-UWSIF-Liver 20140020.32 | 01-UWSIF-Liver | Peak Center:3.060 | 0.759 | 1 | 2607 | 0.024 |  |  | 10.544 | 51.438 | 51.048 | 10.6 | 8.0 | 115.9 | [5 empty] | 0.36648 |  | 0.73139
2014/03/24 11:02:17 | 9 | 01-UWSIF-Liver 20140020.32 | 01-UWSIF-Liver | Peak Center:3.060 | 0.759 | 2 | 2608 | 0 |  |  | 10.5443 | 51.439 | 51.05 | 10.5 | 7.9 | 121.0 | [5 empty] | 0.36647 |  | 0.73137
2014/03/24 11:02:17 | 9 | 01-UWSIF-Liver 20140020.32 | 01-UWSIF-Liver | Peak Center:3.060 | 0.759 | 3 | 2723 | 7.337 |  |  | 12.8629 | 62.75 | 62.269 | 10.5 | 7.9 | 112.9 | [5 empty] | 0.36915 |  | 0.73673
2014/03/24 11:02:17 | 9 | 01-UWSIF-Liver 20140020.32 | 01-UWSIF-Liver | Peak Center:3.060 | 0.759 | 4 |  |  | 1890 | -3.786 | 49.4913 | 48.134 |  |  |  |  | 47.375 | 0.6 | 0.8 | 2.6 | 1.1015 |  | 1.175
2014/03/24 11:02:17 | 9 | 01-UWSIF-Liver 20140020.32 | 01-UWSIF-Liver | Peak Center:3.060 | 0.759 | 5 |  |  | 2851 | 0 | 58.0567 | 56.464 |  |  |  |  | 55.574 | 2.4 | 2.9 | 5.3 | 1.1057 |  | 1.1788
2014/03/24 11:10:45 | 10 | 20140020.073 | 1261 | Instrument: ANNIE | 0.828 | 1 | 2608 | 0.017 |  |  | 9.6716 | 51.471 | 51.081 | 10.7 | 8.1 | 118.0 | [5 empty] | 0.36648 |  | 0.73138
2014/03/24 11:10:45 | 10 | 20140020.073 | 1261 | Instrument: ANNIE | 0.828 | 2 | 2612 | 0 |  |  | 9.6781 | 51.506 | 51.116 | 10.6 | 8.0 | 123.3 | [5 empty] | 0.36647 |  | 0.73137
2014/03/24 11:10:45 | 10 | 20140020.073 | 1261 | Instrument: ANNIE | 0.828 | 3 | 2628 | 13.81 |  |  | 11.4307 | 60.833 | 60.364 | 10.6 | 7.9 | 115.5 | [5 empty] | 0.37151 |  | 0.74147
2014/03/24 11:10:45 | 10 | 20140020.073 | 1261 | Instrument: ANNIE | 0.828 | 4 |  |  | 1878 | -15.037 | 45.4236 | 48.194 |  |  |  |  | 47.44 | 0.6 | 0.8 | 2.7 | 1.0892 |  | 1.1626
2014/03/24 11:10:45 | 10 | 20140020.073 | 1261 | Instrument: ANNIE | 0.828 | 5 |  |  | 2855 | 0 | 53.2421 | 56.489 |  |  |  |  | 55.599 | 2.5 | 2.9 | 5.3 | 1.1057 |  | 1.1787
2014/03/24 11:19:12 | 11 | 20140020.074 | 1235 | Data Analyst:DEW | 0.848 | 1 | 2606 | 0.026 |  |  | 9.4361 | 51.431 | 51.042 | 10.7 | 8.1 | 119.8 | [5 empty] | 0.36648 |  | 0.73134
2014/03/24 11:19:12 | 11 | 20140020.074 | 1235 | Data Analyst:DEW | 0.848 | 2 | 2606 | 0 |  |  | 9.4219 | 51.354 | 50.965 | 10.6 | 8.0 | 124.8 | [5 empty] | 0.36647 |  | 0.73132
2014/03/24 11:19:12 | 11 | 20140020.074 | 1235 | Data Analyst:DEW | 0.848 | 3 | 2915 | 14.783 |  |  | 12.3554 | 67.342 | 66.822 | 10.7 | 8.0 | 116.8 | [5 empty] | 0.37187 |  | 0.74214
2014/03/24 11:19:12 | 11 | 20140020.074 | 1235 | Data Analyst:DEW | 0.848 | 4 |  |  | 1930 | -14.782 | 45.4202 | 49.354 |  |  |  |  | 48.582 | 0.6 | 0.8 | 2.7 | 1.0895 |  | 1.1628
2014/03/24 11:19:12 | 11 | 20140020.074 | 1235 | Data Analyst:DEW | 0.848 | 5 |  |  | 2842 | 0 | 51.781 | 56.266 |  |  |  |  | 55.379 | 2.5 | 3.0 | 5.4 | 1.1057 |  | 1.1786
2014/03/24 11:27:40 | 12 | 20140020.075 | 1360 | Comments: | 0.756 | 1 | 2596 | -0.001 |  |  | 10.5376 | 51.204 | 50.816 | 10.7 | 8.1 | 121.2 | [5 empty] | 0.36647 |  | 0.73134
2014/03/24 11:27:40 | 12 | 20140020.075 | 1360 | Comments: | 0.756 | 2 | 2594 | 0 |  |  | 10.5354 | 51.193 | 50.805 | 10.6 | 8.0 | 126.1 | [5 empty] | 0.36647 |  | 0.73134
2014/03/24 11:27:40 | 12 | 20140020.075 | 1360 | Comments: | 0.756 | 3 | 2819 | 15.53 |  |  | 13.4617 | 65.412 | 64.907 | 10.6 | 8.0 | 118.0 | [5 empty] | 0.37214 |  | 0.7427
2014/03/24 11:27:40 | 12 | 20140020.075 | 1360 | Comments: | 0.756 | 4 |  |  | 1815 | -14.218 | 47.9222 | 46.423 |  |  |  |  | 45.697 | 0.6 | 0.8 | 2.7 | 1.0901 |  | 1.1635
2014/03/24 11:27:40 | 12 | 20140020.075 | 1360 | Comments: | 0.756 | 5 |  |  | 2850 | 0 | 58.2002 | 56.38 |  |  |  |  | 55.491 | 2.4 | 2.9 | 5.3 | 1.1057 |  | 1.1787
2014/03/24 11:36:08 | 13 | 20140020.076 | 1340 |  | 0.846 | 1 | 2607 | 0.018 |  |  | 9.4508 | 51.39 | 51 | 10.8 | 8.2 | 123.9 | [5 empty] | 0.36648 |  | 0.73137
2014/03/24 11:36:08 | 13 | 20140020.076 | 1340 |  | 0.846 | 2 | 2608 | 0 |  |  | 9.4547 | 51.411 | 51.021 | 10.7 | 8.1 | 128.8 | [5 empty] | 0.36647 |  | 0.73136
2014/03/24 11:36:08 | 13 | 20140020.076 | 1340 |  | 0.846 | 3 | 2925 | 15.158 |  |  | 12.4428 | 67.659 | 67.136 | 10.7 | 8.0 | 121.0 | [5 empty] | 0.37201 |  | 0.74245
2014/03/24 11:36:08 | 13 | 20140020.076 | 1340 |  | 0.846 | 4 |  |  | 1896 | -14.535 | 44.7823 | 48.546 |  |  |  |  | 47.786 | 0.6 | 0.9 | 2.8 | 1.0898 |  | 1.1632
2014/03/24 11:36:08 | 13 | 20140020.076 | 1340 |  | 0.846 | 5 |  |  | 2859 | 0 | 52.2168 | 56.605 |  |  |  |  | 55.713 | 2.5 | 3.0 | 5.4 | 1.1057 |  | 1.1787
2014/03/24 11:44:36 | 14 | 20140020.077 | 1580 |  | 0.781 | 1 | 2608 | 0.031 |  |  | 10.246 | 51.433 | 51.043 | 10.9 | 8.2 | 125.6 | [5 empty] | 0.36648 |  | 0.73141
2014/03/24 11:44:36 | 14 | 20140020.077 | 1580 |  | 0.781 | 2 | 2609 | 0 |  |  | 10.2627 | 51.517 | 51.127 | 10.7 | 8.1 | 130.7 | [5 empty] | 0.36647 |  | 0.73138
2014/03/24 11:44:36 | 14 | 20140020.077 | 1580 |  | 0.781 | 3 | 2363 | 13.89 |  |  | 10.9214 | 54.823 | 54.401 | 10.7 | 8.1 | 122.7 | [5 empty] | 0.37154 |  | 0.74154
2014/03/24 11:44:36 | 14 | 20140020.077 | 1580 |  | 0.781 | 4 |  |  | 1760 | -14.425 | 44.7579 | 44.792 |  |  |  |  | 44.091 | 0.7 | 0.9 | 2.8 | 1.0899 |  | 1.1633
2014/03/24 11:44:36 | 14 | 20140020.077 | 1580 |  | 0.781 | 5 |  |  | 2852 | 0 | 56.394 | 56.437 |  |  |  |  | 55.547 | 2.4 | 2.9 | 5.2 | 1.1057 |  | 1.1788
2014/03/24 11:53:04 | 15 | 20140020.078 | 1540 |  | 0.811 | 1 | 2608 | -0.007 |  |  | 9.8698 | 51.448 | 51.058 | 10.9 | 8.2 | 126.1 | [5 empty] | 0.36647 |  | 0.73138
2014/03/24 11:53:04 | 15 | 20140020.078 | 1540 |  | 0.811 | 2 | 2604 | 0 |  |  | 9.8546 | 51.368 | 50.979 | 10.8 | 8.1 | 131.2 | [5 empty] | 0.36647 |  | 0.73138
2014/03/24 11:53:04 | 15 | 20140020.078 | 1540 |  | 0.811 | 3 | 2508 | 14.075 |  |  | 11.0554 | 57.628 | 57.183 | 10.7 | 8.1 | 122.6 | [5 empty] | 0.37161 |  | 0.74168
2014/03/24 11:53:04 | 15 | 20140020.078 | 1540 |  | 0.811 | 4 |  |  | 1983 | -15.281 | 48.658 | 50.565 |  |  |  |  | 49.774 | 0.6 | 0.8 | 2.8 | 1.0889 |  | 1.1624
2014/03/24 11:53:04 | 15 | 20140020.078 | 1540 |  | 0.811 | 5 |  |  | 2852 | 0 | 54.3818 | 56.514 |  |  |  |  | 55.623 | 2.6 | 3.1 | 5.6 | 1.1057 |  | 1.1788
2014/03/24 12:01:31 | 16 | 20140020.079 | 1441 |  | 0.84 | 1 | 2607 | 0.004 |  |  | 9.5251 | 51.427 | 51.037 | 11.0 | 8.3 | 127.8 | [5 empty] | 0.36647 |  | 0.73141
2014/03/24 12:01:31 | 16 | 20140020.079 | 1441 |  | 0.84 | 2 | 2604 | 0 |  |  | 9.5212 | 51.406 | 51.016 | 10.8 | 8.2 | 132.8 | [5 empty] | 0.36647 |  | 0.7314
2014/03/24 12:01:31 | 16 | 20140020.079 | 1441 |  | 0.84 | 3 | 2867 | 14.653 |  |  | 12.1785 | 65.752 | 65.244 | 10.8 | 8.1 | 124.8 | [5 empty] | 0.37182 |  | 0.74212
2014/03/24 12:01:31 | 16 | 20140020.079 | 1441 |  | 0.84 | 4 |  |  | 2145 | -15.24 | 50.8891 | 54.775 |  |  |  |  | 53.918 | 0.7 | 0.9 | 2.8 | 1.089 |  | 1.1625
2014/03/24 12:01:31 | 16 | 20140020.079 | 1441 |  | 0.84 | 5 |  |  | 2854 | 0 | 52.4884 | 56.496 |  |  |  |  | 55.606 | 2.7 | 3.3 | 5.7 | 1.1057 |  | 1.1788
2014/03/24 12:09:59 | 17 | 20140020.080 | 1477 |  | 0.77 | 1 | 2603 | 0.049 |  |  | 10.3754 | 51.349 | 50.96 | 11.0 | 8.4 | 129.7 | [5 empty] | 0.36649 |  | 0.73138
2014/03/24 12:09:59 | 17 | 20140020.080 | 1477 |  | 0.77 | 2 | 2600 | 0 |  |  | 10.3699 | 51.322 | 50.933 | 10.8 | 8.2 | 134.3 | [5 empty] | 0.36647 |  | 0.73134
2014/03/24 12:09:59 | 17 | 20140020.080 | 1477 |  | 0.77 | 3 | 2765 | 14.889 |  |  | 12.8982 | 63.835 | 63.342 | 10.9 | 8.2 | 126.1 | [5 empty] | 0.37191 |  | 0.74223
2014/03/24 12:09:59 | 17 | 20140020.080 | 1477 |  | 0.77 | 4 |  |  | 1776 | -14.725 | 45.7998 | 45.189 |  |  |  |  | 44.482 | 0.7 | 0.9 | 2.8 | 1.0896 |  | 1.163
2014/03/24 12:09:59 | 17 | 20140020.080 | 1477 |  | 0.77 | 5 |  |  | 2839 | 0 | 56.9387 | 56.179 |  |  |  |  | 55.294 | 2.5 | 3.0 | 5.4 | 1.1057 |  | 1.1787
2014/03/24 12:18:27 | 18 | 20140020.081 | 1603 |  | 0.76 | 1 | 2591 | -0.005 |  |  | 10.4463 | 51.029 | 50.642 | 10.9 | 8.3 | 130.0 | [5 empty] | 0.36647 |  | 0.73136
2014/03/24 12:18:27 | 18 | 20140020.081 | 1603 |  | 0.76 | 2 | 2592 | 0 |  |  | 10.4633 | 51.112 | 50.724 | 10.8 | 8.1 | 134.9 | [5 empty] | 0.36647 |  | 0.73136
2014/03/24 12:18:27 | 18 | 20140020.081 | 1603 |  | 0.76 | 3 | 2675 | 14.675 |  |  | 12.6115 | 61.605 | 61.13 | 10.8 | 8.1 | 126.8 | [5 empty] | 0.37183 |  | 0.7421
2014/03/24 12:18:27 | 18 | 20140020.081 | 1603 |  | 0.76 | 4 |  |  | 1915 | -14.971 | 50.1513 | 48.84 |  |  |  |  | 48.075 | 0.7 | 0.9 | 2.8 | 1.0893 |  | 1.1627
2014/03/24 12:18:27 | 18 | 20140020.081 | 1603 |  | 0.76 | 5 |  |  | 2837 | 0 | 57.577 | 56.071 |  |  |  |  | 55.188 | 2.6 | 3.1 | 5.5 | 1.1057 |  | 1.1787
2014/03/24 12:26:55 | 19 | 20140020.082 | 1604 |  | 0.809 | 1 | 2593 | -0.011 |  |  | 9.8435 | 51.184 | 50.796 | 11.0 | 8.3 | 131.2 | [5 empty] | 0.36647 |  | 0.73129
2014/03/24 12:26:55 | 19 | 20140020.082 | 1604 |  | 0.809 | 2 | 2592 | 0 |  |  | 9.8321 | 51.125 | 50.738 | 10.8 | 8.2 | 136.2 | [5 empty] | 0.36647 |  | 0.73129
2014/03/24 12:26:55 | 19 | 20140020.082 | 1604 |  | 0.809 | 3 | 2384 | 13.758 |  |  | 10.0654 | 52.338 | 51.936 | 11.1 | 8.3 | 122.0 | [5 empty] | 0.3715 |  | 0.74135
2014/03/24 12:26:55 | 19 | 20140020.082 | 1604 |  | 0.809 | 4 |  |  | 1949 | -14.859 | 48.1132 | 49.876 |  |  |  |  | 49.095 | 0.6 | 0.7 | 2.7 | 1.0894 |  | 1.1629
2014/03/24 12:26:55 | 19 | 20140020.082 | 1604 |  | 0.809 | 5 |  |  | 2839 | 0 | 54.2166 | 56.203 |  |  |  |  | 55.317 | 6.1 | 7.1 | 10.5 | 1.1057 |  | 1.1787
2014/03/24 12:35:23 | 20 | 20140020.083 | 1472 |  | 0.794 | 1 | 2594 | -0.026 |  |  | 10.022 | 51.146 | 50.758 | 11.2 | 8.6 | 133.0 | [5 empty] | 0.36646 |  | 0.73135
2014/03/24 12:35:23 | 20 | 20140020.083 | 1472 |  | 0.794 | 2 | 2594 | 0 |  |  | 10.0299 | 51.186 | 50.798 | 11.0 | 8.3 | 137.9 | [5 empty] | 0.36647 |  | 0.73137
2014/03/24 12:35:23 | 20 | 20140020.083 | 1472 |  | 0.794 | 3 | 2593 | 13.158 |  |  | 11.7333 | 59.88 | 59.418 | 11.0 | 8.3 | 129.7 | [5 empty] | 0.37128 |  | 0.74099
2014/03/24 12:35:23 | 20 | 20140020.083 | 1472 |  | 0.794 | 4 |  |  | 1707 | -13.312 | 42.6235 | 43.366 |  |  |  |  | 42.686 | 0.8 | 0.9 | 2.9 | 1.0911 |  | 1.1646
2014/03/24 12:35:23 | 20 | 20140020.083 | 1472 |  | 0.794 | 5 |  |  | 2829 | 0 | 55.012 | 55.97 |  |  |  |  | 55.088 | 2.5 | 2.9 | 5.3 | 1.1057 |  | 1.1787
2014/03/24 12:43:51 | 21 | 36-UWSIF-UT Glut 120140020.36 | 36-UWSIF-UT Glut 1 |  | 0.788 | 1 | 2584 | -0.018 |  |  | 10.0745 | 51.025 | 50.639 | 11.0 | 8.3 | 133.2 | [5 empty] | 0.36647 |  | 0.73133
2014/03/24 12:43:51 | 21 | 36-UWSIF-UT Glut 120140020.36 | 36-UWSIF-UT Glut 1 |  | 0.788 | 2 | 2580 | 0 |  |  | 10.0509 | 50.906 | 50.52 | 10.8 | 8.2 | 137.9 | [5 empty] | 0.36647 |  | 0.73135
2014/03/24 12:43:51 | 21 | 36-UWSIF-UT Glut 120140020.36 | 36-UWSIF-UT Glut 1 |  | 0.788 | 3 | 2049 | -4.102 |  |  | 9.2844 | 47.024 | 46.667 | 10.8 | 8.1 | 129.3 | [5 empty] | 0.36497 |  | 0.72835
2014/03/24 12:43:51 | 21 | 36-UWSIF-UT Glut 120140020.36 | 36-UWSIF-UT Glut 1 |  | 0.788 | 4 |  |  | 1603 | -14.173 | 39.7634 | 40.15 |  |  |  |  | 39.521 | 0.7 | 0.8 | 2.8 | 1.0902 |  | 1.1636
2014/03/24 12:43:51 | 21 | 36-UWSIF-UT Glut 120140020.36 | 36-UWSIF-UT Glut 1 |  | 0.788 | 5 |  |  | 2817 | 0 | 55.2376 | 55.775 |  |  |  |  | 54.896 | 2.3 | 2.8 | 5.1 | 1.1057 |  | 1.1786
2014/03/24 12:52:20 | 22 | 36-UWSIF-UT Glut 120140020.37 | 36-UWSIF-UT Glut 1 |  | 0.814 | 1 | 2573 | -0.002 |  |  | 9.7097 | 50.801 | 50.416 | 10.9 | 8.3 | 132.6 | [5 empty] | 0.36647 |  | 0.73138
2014/03/24 12:52:20 | 22 | 36-UWSIF-UT Glut 120140020.37 | 36-UWSIF-UT Glut 1 |  | 0.814 | 2 | 2575 | 0 |  |  | 9.7092 | 50.798 | 50.413 | 10.8 | 8.1 | 137.6 | [5 empty] | 0.36647 |  | 0.73138
2014/03/24 12:52:20 | 22 | 36-UWSIF-UT Glut 120140020.37 | 36-UWSIF-UT Glut 1 |  | 0.814 | 3 | 2108 | -4.142 |  |  | 9.271 | 48.505 | 48.138 | 10.8 | 8.1 | 129.6 | [5 empty] | 0.36496 |  | 0.72835
2014/03/24 12:52:20 | 22 | 36-UWSIF-UT Glut 120140020.37 | 36-UWSIF-UT Glut 1 |  | 0.814 | 4 |  |  | 1653 | -14.234 | 39.8612 | 41.577 |  |  |  |  | 40.926 | 0.6 | 0.8 | 2.8 | 1.0901 |  | 1.1636
2014/03/24 12:52:20 | 22 | 36-UWSIF-UT Glut 120140020.37 | 36-UWSIF-UT Glut 1 |  | 0.814 | 5 |  |  | 2819 | 0 | 53.5035 | 55.806 |  |  |  |  | 54.927 | 2.4 | 2.8 | 5.2 | 1.1057 |  | 1.1788
2014/03/24 13:00:48 | 23 | 39-UWSIF-UW Glut 220140020.33 | 39-UWSIF-UW Glut 2 |  | 0.78 | 1 | 2576 | -0.034 |  |  | 10.1371 | 50.821 | 50.436 | 11.0 | 8.4 | 133.8 | [5 empty] | 0.36646 |  | 0.73131
2014/03/24 13:00:48 | 23 | 39-UWSIF-UW Glut 220140020.33 | 39-UWSIF-UW Glut 2 |  | 0.78 | 2 | 2579 | 0 |  |  | 10.1438 | 50.855 | 50.47 | 10.9 | 8.2 | 139.0 | [5 empty] | 0.36647 |  | 0.73134
2014/03/24 13:00:48 | 23 | 39-UWSIF-UW Glut 220140020.33 | 39-UWSIF-UW Glut 2 |  | 0.78 | 3 | 2215 | 28.427 |  |  | 10.1401 | 50.836 | 50.439 | 10.9 | 8.2 | 130.8 | [5 empty] | 0.37685 |  | 0.75213
2014/03/24 13:00:48 | 23 | 39-UWSIF-UW Glut 220140020.33 | 39-UWSIF-UW Glut 2 |  | 0.78 | 4 |  |  | 1729 | 37.674 | 43.6059 | 43.583 |  |  |  |  | 42.876 | 0.6 | 0.8 | 2.8 | 1.1468 |  | 1.221
2014/03/24 13:00:48 | 23 | 39-UWSIF-UW Glut 220140020.33 | 39-UWSIF-UW Glut 2 |  | 0.78 | 5 |  |  | 2824 | 0 | 55.9521 | 55.923 |  |  |  |  | 55.041 | 2.4 | 3.0 | 5.4 | 1.1057 |  | 1.179
2014/03/24 13:09:16 | 24 | 39-UWSIF-UW Glut 220140020.34 | 39-UWSIF-UW Glut 2 |  | 0.7282 | 1 | 2580 | 0.016 |  |  | 10.8673 | 50.864 | 50.479 | 11.0 | 8.4 | 134.8 | [5 empty] | 0.36648 |  | 0.73135
2014/03/24 13:09:16 | 24 | 39-UWSIF-UW Glut 220140020.34 | 39-UWSIF-UW Glut 2 |  | 0.7282 | 2 | 2582 | 0 |  |  | 10.8796 | 50.921 | 50.536 | 11.0 | 8.3 | 140.5 | [5 empty] | 0.36647 |  | 0.73134
2014/03/24 13:09:16 | 24 | 39-UWSIF-UW Glut 220140020.34 | 39-UWSIF-UW Glut 2 |  | 0.7282 | 3 | 2044 | 28.389 |  |  | 10.0958 | 47.253 | 46.883 | 10.9 | 8.2 | 132.4 | [5 empty] | 0.37684 |  | 0.7521
2014/03/24 13:09:16 | 24 | 39-UWSIF-UW Glut 220140020.34 | 39-UWSIF-UW Glut 2 |  | 0.7282 | 4 |  |  | 1610 | 37.899 | 43.5575 | 40.644 |  |  |  |  | 39.984 | 0.7 | 0.9 | 2.8 | 1.1471 |  | 1.2213
2014/03/24 13:09:16 | 24 | 39-UWSIF-UW Glut 220140020.34 | 39-UWSIF-UW Glut 2 |  | 0.7282 | 5 |  |  | 2850 | 0 | 60.43 | 56.387 |  |  |  |  | 55.499 | 2.4 | 2.9 | 5.2 | 1.1057 |  | 1.1791
2014/03/24 13:17:44 | 25 | 01-UWSIF-Liver 20140020.33 | 01-UWSIF-Liver |  | 0.7543 | 1 | 2611 | 0.003 |  |  | 10.6398 | 51.584 | 51.193 | 11.2 | 8.5 | 138.8 | [5 empty] | 0.36647 |  | 0.73132
2014/03/24 13:17:44 | 25 | 01-UWSIF-Liver 20140020.33 | 01-UWSIF-Liver |  | 0.7543 | 2 | 2610 | 0 |  |  | 10.6217 | 51.496 | 51.106 | 11.1 | 8.4 | 143.9 | [5 empty] | 0.36647 |  | 0.73131
2014/03/24 13:17:44 | 25 | 01-UWSIF-Liver 20140020.33 | 01-UWSIF-Liver |  | 0.7543 | 3 | 2720 | 7.303 |  |  | 12.854 | 62.319 | 61.841 | 11.1 | 8.3 | 135.1 | [5 empty] | 0.36914 |  | 0.73665
2014/03/24 13:17:44 | 25 | 01-UWSIF-Liver 20140020.33 | 01-UWSIF-Liver |  | 0.7543 | 4 |  |  | 1893 | -3.763 | 49.4953 | 47.839 |  |  |  |  | 47.085 | 0.7 | 0.9 | 2.8 | 1.1015 |  | 1.1753
2014/03/24 13:17:44 | 25 | 01-UWSIF-Liver 20140020.33 | 01-UWSIF-Liver |  | 0.7543 | 5 |  |  | 2845 | 0 | 58.2774 | 56.328 |  |  |  |  | 55.44 | 2.6 | 3.1 | 5.6 | 1.1057 |  | 1.179
2014/03/24 13:26:12 | 26 | 01-UWSIF-Liver 20140020.34 | 01-UWSIF-Liver |  | 0.731 | 1 | 2601 | -0.008 |  |  | 10.9138 | 51.278 | 50.889 | 11.2 | 8.6 | 139.9 | [5 empty] | 0.36647 |  | 0.73132
2014/03/24 13:26:12 | 26 | 01-UWSIF-Liver 20140020.34 | 01-UWSIF-Liver |  | 0.731 | 2 | 2598 | 0 |  |  | 10.9129 | 51.274 | 50.885 | 11.1 | 8.4 | 144.8 | [5 empty] | 0.36647 |  | 0.73133
2014/03/24 13:26:12 | 26 | 01-UWSIF-Liver 20140020.34 | 01-UWSIF-Liver |  | 0.731 | 3 | 2630 | 7.241 |  |  | 12.842 | 60.338 | 59.875 | 11.2 | 8.4 | 136.6 | [5 empty] | 0.36912 |  | 0.73662
2014/03/24 13:26:12 | 26 | 01-UWSIF-Liver 20140020.34 | 01-UWSIF-Liver |  | 0.731 | 4 |  |  | 1834 | -3.71 | 49.4498 | 46.319 |  |  |  |  | 45.588 | 0.7 | 0.9 | 2.9 | 1.1016 |  | 1.1753
2014/03/24 13:26:12 | 26 | 01-UWSIF-Liver 20140020.34 | 01-UWSIF-Liver |  | 0.731 | 5 |  |  | 2831 | 0 | 59.8836 | 56.092 |  |  |  |  | 55.208 | 2.6 | 3.1 | 5.5 | 1.1057 |  | 1.1788
Job # : 2014-0020 Run 2 Stockwell
Analyst: DEW
Sample Count: 0547
Helium Pressure:1850
Oxygen Pressure:2400
Mass 28:10
Mass 29:7
Mass 30:64
Peak Center:3.060
Instrument: ANNIE
Data Analyst:DEW
Comments:

## Analysis Information
Principal of operation
●The Finnigan DeltaPlus XP is run in continuous flow mode and connected to either a Carlo Erba 1110 or a Costech 4010 elemental analyzer via a Finnigan ConFlo III interface.
●A sample contained within a tin capsule is dropped into a combustion reactor held at 1020°C. For nitrogen and carbon analysis, the combustion reactor contains chromium oxide for oxidation and silvered cobaltous/cobaltic oxide for removal of sulfur. When the tin capsule is exposed to a gas flow temporarily enriched with ultra-high purity oxygen (25-30 mL min-1), flash combustion occurs which raises the temperature of the sample to >1700°C. The encapsulated sample, depending on its composition, combusts generating one or more of these gases: N2, NxOx, CO2, and H2O. The reduction reactor contains reduced copper wires for the reduction of nitrogen oxides to N2 and the removal of excess O2. An adsorption trap containing magnesium perchlorate removes the H2O. The remaining N2 and CO2 gases travel through a Porapak Q chromatographic column (80-100 mL min-1 at 50˚C) as the carrier gas and then moves to the ConFlo III open-split interface.
●The gas then moves through the ConFlo III to the Finnigan DeltaPlus XP mass spectrometer. Reference gases are pulsed by the ConFlo III into the mass spectrometer for proper mass balance.
QA/QC
Quality assurance of carbon and nitrogen isotope composition of solids is based on the standard uncertainty of the known value of the secondary laboratory reference material calculated on multiple analyses. For carbon isotope composition, if the standard uncertainty is greater than 0.15‰, the unknowns are re-analyzed (until the 2-sigma expanded standard uncertainty of the result is better than 0.3‰). The carbon isotopic composition is reported in per mil relative to VPDB scale such that USGS40 glutamic acid, USGS41 glutamic acid, and NIST 8542 sucrose ANU, respectively are -26.24‰, +37.76‰, and -10.05‰. For nitrogen isotope ratio composition, if the standard uncertainty is greater than 0.2‰, the unknowns are re-analyzed (until the 2-sigma expanded standard uncertainty of the result is better than 0.4‰). The nitrogen isotopic composition is reported in per mil relative to nitrogen in air on a scale such that USGS40 glutamic acid, USGS41 glutamic acid, and NIST8549 potassium nitrate, respectively are -4.52‰, +47.57‰, and 4.36‰.
Normalization
Whenever isotopic analyses are performed, reference materials must be included with the unknowns. In general two ‘bracketing’ reference materials should be used for correcting purposes. These reference materials ideally would be similar in chemical complexity to the unknowns but have very different isotopic compositions. Using two such reference materials, you can generate two linear equations, one for each referenece material. Since these linear equations will have the same slope and intercept, combining the two equations will generate a single linear equation that can be used to correct the unknowns.
The linear equation is derived as follows:
Slope (m): m=[δRM 1-known – δRM 2-known] / [δRM 1-measured – δRM 2-measured]
Intercept (b): b= δRM 1-known – [m* δRM 1-measured]
Normalization equation:
δSA-COR = m* δSA-measured + b
Where:
δRM 1-known = the accepted delta value of reference material 1
δRM 2-known = the accepted delta value of reference material 2
δRM 1-measured = the measured delta value of reference material 1
δRM 2-measured = the measured delta value of reference material 2
δSA-measured = measured delta value of sample
δSA-COR = corrected delta value of sample
13C and 15N analyses of solids

## Contact
For questions about the analysis, please contact:
|  | Chandelle Macdonald | Master Technician
| email: | cmacdon1@uwyo.edu
| Phone: | (307) 766-6373
Contact information for UW Stable Isotope Facility:
| Personnel: | Dr. David G. Williams | Faculty Director
dgw@uwyo.edu
|  | Craig Cook | Facility Director
ccook21@uwyo.edu
|  | Chandelle Macdonald | Master Technician
cmacdon1@uwyo.edu
| Address: | UWYO Stable Isotope Facility
University of Wyoming
Berry Biodiversity Center Rm 214
Laramie, WY 82071
| email: | uwyosif@uwyo.edu
| Phone: | (307) 766-6373
| Fax: | (307) 766-6403